## Instructions
This spreadsheet is provided by Morehouse Instrument Company. It is to be used as a guide to help calculate CMC for Force or Torque Measurements. Most of this workbook is setup for labs following the ASTM E74 standard. With slight modification the uncertainty sheet could be used for instruments calibrated in accordance with another test method or standard. Only enter information in light grey boxes and choose a drop down for the dark grey boxes. All information entered must converted to like units. Disclaimer: This template is for simplified calculation of CMC compliments of E=mc3 Solutions and Morehouse Instrument Company who assumes no responsibility or liability for any unintentional errors. It remains the users responsibility to validate the calculations and confirm suitability for practical purpose. Dilip Shah of E=mc3 Solutions can be reached at emc3solu@aol.com and Henry Zumbrun of Morehouse can be reached at hzumbrun@mhforce.com.
Uncertainty Worksheet - Section 1: Data Entry
This sheet is to be used as the master form to calculate CMC for a range of up to 12 points. Its primary purpose is to aid in figuring CMC by including repeatability and resolution of an Unit Under Test (UUT).
Cell | Description | Notes
B4 | Laboratory
B5 | Technician Initials
B6 | Date
B7 | Applied Maximum
B8 | Standards Used (Ref and UUT)
B10 & C10 | Resolution UUT | This is the resolution of the unit under test (UUT) that you are using for the Repeatability Study (what you are testing). CMC requires repeatability studies. We recommend using a stable device with high resolution. An example of how to calculate resolution is in the tab labeled "Calc Resolution Ex".
A13 | Standards Drop Down | Choose one of the three options from the drop down.
B13 & C13 | ASTM E74 LLF | This is your ASTM A74 LLF, which is found on your ASTM E74 calibration report. LLF is a statistical estimate of the error in forces computed from the calibration equation of a force–measuring instrument when the instrument is calibrated in accordance with the ASTM E74 standard. If you are not using ASTM E74, leave these cells blank.
B14 & C14 | Resolution of Reference | The resolution of the reference standard is found on your calibration report.
B15 | Temperature Spec per degree C % | This is found on the load cell specification sheet. Temperature Effect on Sensitivity, % RDG/100 F
B18 | Max Temperature Variation per degree C of Environment | This is the max temperature variation of load cells, which is found on the specification sheet under Temperature - Effect on Sensitivity, % reading/degrees. During a typical calibration in a tightly controlled environment the temperature varies by no more than 1 degree C. Environmental factors - This would be the difference in temperature from 23 degree C. This is 0.0015 % per degree C from 23. If you are operating at 20 degrees C, you would multiply 0.0015 % by 3. Example provided is from a Morehouse load cell. Specifications can vary.
B20 | Morehouse CMC (Ref Lab) | Optional: If entering Engineering Units, leave blank and complete the Ref Laboratory Uncertainty Per Point (P29 - 40). This is the reference lab CMC. It is the CMC statement for the range calibrated found on the certificate of calibration.
B22 | Non ASTM or ISO 376 (Tolerance, Non-Linearity, SEB) | If you chose ASTM E74 or ISO 376 in A13, then leave blank. If you chose Other in A13, then use this field to enter Tolerance with nonlinearity or SEB if making ascending and descending measurements. When not using ASTM E74, SEB output can be used, or the square root of the sum of the squares for non-linearity worst case and hysteresis worst case.
B23 | Miscellaneous Error | This can be measurement error, creep, side load sensitivity or other known error sources. For example, if using a Morehouse UCM and Morehouse Cell the press will transfer the force applied to the load cell at an angle no more than 1/16th inch measured off. Example provided is from a Morehouse load cell. centerline of the load cell. This number is side load sensitivity 0.05 % * 0.0625 = 0.003 %.
B25 | Conv Repeatability Data To Eng. Units Drop Down | If data is in engineering units LBF, KGF, or N, then choose no. If data is in DIV, mV/V or other non-engineering unit, then choose yes.
N7-18 & O7-18 | Ref Standard Stability | You will need to have two calibrations, performed at different times to characterize the Reference Standard Stability. You should use a per point analysis between the previous calibration and your current calibration. If you do not have previous data, we suggest a conservative number. For Morehouse Precision load cells, we believe 0.05 % of applied force could be used from 10 % of Capacity through 100 %. Below 10 LBF, we believe 0.15 % of applied could be used.
N22, O22, P22 | ISO 376 Uncertainty Coefficients | If you chose ISO 376 in A13, then fill in the uncertainty coefficients from your calibration report or calculated with methods provided in ISO 376. If no C2 provided, then leave as zero.
B29-40 | Applied | After you complete the Ref Standard Stability (N7-18), the same numbers will automatically be entered here.
C29-40, D29-40, E29-40, F29-40 | Repeatability of UUT (Run1-Run4) | Do a repeatability study using a test artifact. We strongly recommend using an Ultra Precision Load Cell that is very stable and repeatable as not doing so will have a negative impact on the CMC. We suggest purchasing an artifact that can also be used for SPC testing as this will strengthen the measurement assurance process. The repeatability study on this worksheet is set up for 3 or 4 test runs, each with up to 12 test points. We suggest exercising the artifact and performing 4 runs of tests and recording the artifacts output at each test point. The worksheet provides for up to 12 different test points. When testing repeatability, moving or rotating the artifact is not required. Using the pooled std deviation may not be in line with current ISO and ILAC standards. It is recommended to test each individual test point against the predicted response from the ASTM E74 curve fit to ensure compliance with ISO and ILAC standards. If the pooled standard deviation is larger than the difference between the predicted response for each test point, then this method should not be argued. Morehouse has conferred with several organizations, including NIST on this subject. Until a document is drafted for continuous force reading devices, it is our belief that each lab should do their own analysis and determine the appropriate method for their use. The example is how Morehouse derives our CMC's and this was deemed acceptable. ASTM E74 Appendix Section X1.4.1.2 includes the pooled standard deviation method in the Appendix for calculating Ur (Uncertainty due to force-measuring instrument responses).
S29-40 | % or Eng. Drop Down | Choose % from the drop down if the CMC of the reference lab is % of applied. Choose Eng. Units if the reference lab reported CMC in Engineering Units.
P29-40 | Ref Laboratory Uncertainty Per Point | Optional: If the reference lab reported CMC in Engineering Units, then fill in the values per point.
Uncertainty Worksheet - Section2: Data Analysis
Cell | Description | Notes
D51-60 | Distribution | It may be necessary to select the appropriate distribution to change the pre-selected values. Once everything is selected, the uncertainty per test point will be determined on each test point.
R & R Between Techs Worksheet
Use this Repeatability and Reproducibility Worksheet for R & R between technicians, between machines, or between different instruments. It is setup to be used as R & R between technicians, if using multiple technicians. This form will be populated with information found in the Data Entry Sheet. Per ISO 5725 the general term for variability between repeated measurements is precision. Two conditions of precision, termed repeatability and reproducibility conditions, have been found necessary and, for many practical cases, useful for describing the variability of a measurement method, under repeatability conditions By default repeatability between technicians is set to none. It is recommended that this contributions be included, but most accreditation bodies do not require it. There are several other methods for determining reproducility and this is just one method. The ASTM LLF is the really an uncertainty calculation that quantifies reproducibility of the equipment and the deviation from the interpolation equation per section 8.4 of ASTM E74. ISO 376 has rotational tests to capture reproducibility. Neither capture reproducibility of the process with different operators. The ANOVA analysis seems to complete this and capture that source of uncertainty. This worksheet was created by Dilip Shah of E=Mc3 Solutions and edited by Henry Zumbrun of Morehouse Instrument Company.
Cell | Description | Notes
D7 and D8 | Repeatability and Reproducibility Between Techs | Here one can decide to include or exclude the uncertainty contribution for Repeatability or Repeatability and Reproducibility
E5 | Y or N Converts to Students T | If you want to report Uncertainties at k = 2, or other coverage factors to comply with Ilac P14 or have all points in the range agree to the appropriate coverage factor you with to report, set this value to Y. It will give Type A's 1000 degrees of freedom, and to do this; requires the correction for the population standard deviation for infinite degrees of freedom for Type A data.
E25-J44 | R & R Study between multiple techs or machines | Repeatability and Reproducibility between technicians: This should be performed whenever there is a change in personnel or the first time a budget is established. Data can be entered for up to 6 technicians. The technicians should measure the same test point up to 20 times each and the reduced data should be entered under the Technician cells in light yellow. We recommend a minimum of 5 points per technician. When we run these tests we typically perform 10 points per technician.
F21 | Select the Confidence Internal | There is a drop down where one can select the Confidence Interval needed. If you wanted to set to k = 2, 95.45 % should be selected.
J50 | Convert to Eng Units | If the readings are in mV/V or not exact Engineering Units, select "Yes" to convert to the appropriate units. When R & R studies are performed this must be changed to yes and the force point used for R & R must be entered. This is not needed if the force applied is 10,000 lbf and the meter reads 10,000 lbf. It needs to be used if the force applied is 10,000 lbf and the output is read in mV/V.
CMC Summary
When all data has been entered properly, the CMC summary tab will will graph and calculate the uncertainty per test point. This worksheet calculates the measurement uncertainty between test points. Visit www.mhforce.com for additional information on calculating Measurement Uncertainty for Force or Torque Measurements. Information on the training offered can be found at http://www.mhforce.com/Training/TrainingCourses
Cell | Description | Notes
K8-19 | Exclude a Point | Points can be excluded. The slope will not work if this is done, though points can be removed for a better coefficient fit
I9-19 | Force Value | Interpolates in between points if needed
I42 | Expanded UNC (CMC) | ISO/IEC 17025 Accredited Calibrations. The International Laboratory Accreditation Cooperation (ILAC) P-14 section 6.4 requires contributions such as the resolution of the Unit Under Test (UUT) be included in the uncertainty per point value reported on the certificate of calibration. Euramet CG4 v2.0 Uncertainty of Force Measurements also requires the resolution of both the standard and unit under test be reported as standard contributors. The per point analysis consists of the UUT resolution and the CMC of our standard (which includes the resolution of the reference standard if primary standards were not used). When a full uncertainty analysis per ILAC P-14 is performed, as required by ISO/IEC 17025, the CMC (Calibration and Measurement Capability) is derived and reported to the user. There is a tab in this worksheet to help with calculating the CMC per point. Morehouse can provide guidance on using this workbook. It is recommended a professional consultant be contacted if you are having problems figuring out CMC or what to report. We offer training with a focus on measurement uncertainty with an independent consultant (Dilip Shah of E=mc3 solutions) two to three times per year. Information on the training offered can be found at http://www.mhforce.com/customer-education
There are two sheets for additional calculations:
En for PT: This workseet can be used for Intra Laboratory Comparisons. Establishing a PT plan is a requirement for accreditation. Morehouse can calibrate reference standard force or torque cells, which can be used with the En For PT sheet to satisfy Proficiency Testing Requirements.
Calc Resolution Ex: This worksheet shows how to calculate resolution of a force-measuring device.
Revision Log
Date | Changes
11.14.2018 | Changed some references to the new ISO/IEC 17025 standard and fixed an error with R & R between techs as degrees of freedom was not calculating properly.
9.4.2020 | Created instructions, definitions, and new sheet for Calc Resolution Ex. Added color for drop down boxes.
10.03.2021 | Added the ability to choose a coverage factor and then set normal data to have infinite degrees of freedom. CMC exclude ability More Compliance to ILAC P-14
9.16.2022 | Automatically Calculated Best Fit Line - R & R between technicians scales maximum value throughout the range on a per point basis, some additional cleanup

## Definitions
Definitions
ASTM E74 – Standard Practices for Calibration and Veriﬁcation for Force-Measuring Instruments: ASTM E74 is a practice that specifies procedures for the calibration of force-measuring instruments.
Best existing force measuring instrument (ILAC P14): The term “best existing force measuring instrument” is understood as a force measuring instrument to be calibrated that is commercially or otherwise available for customers, even if it has a special performance (stability) or has a long history of calibration. For force calibrations this is often a very stable force transducer (load cell) and indicator with enough resolution to observe differences in repeatability conditions.
Calibration and Measurement Capability (ILAC-P14): A CMC is a Calibration and Measurement Capability available to customers under normal conditions:
a) as described in the laboratory’s scope of accreditation granted by a signatory to the ILAC Arrangement; or
b) as published in the BIPM key comparison database (KCDB) of the CIPM MRA.
The scope of accreditation of an accredited calibration laboratory shall include the Calibration and Measurement Capability (CMC) expressed in terms of:
a) measurand or reference material;
b) calibration/measurement method/procedure and/or type of instrument/material to be calibrated/measured;
c) measurement range and additional parameters where applicable, e.g., frequency of applied voltage;
d) uncertainty of measurement.
Commercial Calibration or Quality Compliance Test: This is usually a 5-6-point calibration done by the manufacturer to verify the force measuring instrument is within the Manufacturer’s specification in regard to non-linearity, Static Error Band (SEB), and Hysteresis. It does not give the expected performance of the force measuring instrument as it has not been subjected to testing repeatability and reproducibility during the calibration. ASTM E74 and ISO 376 are standards that give better guidance on what should be tested to determine expected performance of the force measuring instrument.
Environmental Factors: Environmental conditions, such as temperature, influences the force transducer output. The most common specification is temperature effect found on the force measuring instruments specification sheet. It is important to note that any deviation in environmental conditions from the temperature the force measuring instrument was calibrated at needs to be accounted for in the measurement uncertainty in measurements using the force transducer by the user.
Force Units: A force unit can be any unit representing a force. Common force units are N, kgf, lbf. The SI unit for force is N (Newton).
Hysteresis: the phenomenon in which the value of a physical property lags behind changes in the effect causing it, as for instance when magnetic induction lags behind the magnetizing force. For Force measurements Hysteresis is often defined as the algebraic difference between output at a given load descending from the maximum load and output at the same load ascending from the minimum load. Normally expressed in units of % Full Scale. It is most commonly calculated between 40 - 60 % of full scale.
lower limit factor (llf): ASTM specific term where the ASTM E74 standard uses a method of least squares to fit a polynomial function to the data points. The standard deviation of all the deviations from the predicted values by the fit function versus the observed values is found by taking the square root of the sum of all the squared deviations divided by the number of samples minus the degree of polynomial fit used minus one. This number is then multiplied by a coverage factor (k) of 2.4 and then multiplied by the average ratio of force to deflection from the calibration data. The llf is a statistical estimate of the error in forces computed from the calibration equation of a force–measuring instrument when the instrument is calibrated in accordance with this practice.
ISO 376 - Calibration of force proving instruments used for the verification of uniaxial testing machines: ISO 376 is an International Standard that specifies a method for the calibration of force-proving instruments used for the static verification of uniaxial testing machines (e.g. tension/compression testing machines) and describes a procedure for the classification of these instruments.
Metrological traceability (JCGM 200:2012, 2.41): property of a measurement result whereby the result can be related to a reference through a documented unbroken chain of calibrations, each contributing to the measurement uncertainty.
Non-Linearity: The quality of a function that expresses a relationship that is not one of direct proportion. For Force measurements, Non-Linearity is defined as the algebraic difference between the output at a specific load and the corresponding point on the straight line drawn between the outputs at minimum load and maximum load. Normally expressed in units of % of Full Scale. It is most commonly calculated between 40 - 60 % of full scale.
Non-Repeatability (per force transducer specification and not JCGM 200:2012): The maximum difference between output readings for repeated loadings under identical loading and environmental conditions. Normally expressed in units as a % of Rated Output (RO).
Other Force Measurement Errors: Most force measuring instruments are susceptible to misalignment error, errors from not exercising the force measuring instrument to full capacity, and errors from improper adapter use. In almost all cases, there will be additional errors if the end user fails to have the force measuring instrument calibrated with the same adapters being used in their application. Other errors may include temperature change under no load conditions. Errors from loading equipment not being level, square and rigid can have significant contributions.
Primary Standard: Per ASTM E74 a deadweight force applied directly without intervening mechanisms such as levers, hydraulic multipliers, or the like whose mass has been determined by comparison with reference standards traceable to the International System of Units (SI) of mass.
Note:
Weights used for force measurement require the correction for the effects of Local Gravity, Air Buoyancy and must be adjusted to within 0.005 % of nominal force value. The uncertainty budget for primary standards also needs to consider possible force-generating mechanisms other than gravity and air buoyancy, including magnetic, electrostatic and aerodynamic effects.
Rated Output or RO: The output corresponding to capacity, equal to the algebraic difference between the signal at “(minimal load + capacity)” and the signal at minimum load.
Reference Standard(s) Calibration Uncertainty: This is usually the measurement uncertainty in the calibration of reference standard(s) used to calibrate the force measuring instrument. It is the uncertainty contribution in the calibration of reference standard(s) used to calibrate the force measuring instrument.
Repeatability condition of measurement, repeatability condition (JCGM 200:2012, 2.20): condition of measurement, out of a set of conditions that includes the same measurement procedure, same operators, same measuring system, same operating conditions and same location, and replicate measurements on the same or similar objects over a short period of time
Measurement repeatability, Repeatability (JCGM 200:2012, VIM 2.21): measurement precision under a set of repeatability conditions of measurement.
Example: a laboratory performing calibrations from 10 N through 10,000 N. The ranges calibrated may be 10 N - 100 N, 100 N - 1,000 N, and 1 000 N – 10 000 N. Recommended practice would be to take test points at 10, 20, 30, 40, 50, 60, 70, 80, 90, 100, 200, 300, 400, 500, 600, 700, 800, 900, 1 000, 2 000, 3 000, 4 000, 5 000, 6 000, 7 000, 8 000, 9 000, and
10 000 N.
Note 1:
For this application zero should never be considered as a first test point. A force measuring instrument should not be used to calibrate other force measuring instruments outside the range it was calibrated over.
Note 2:
A force measuring instrument calibrated from 10 % through 100 % of its range may not be capable of calibrating force measuring instruments outside of this range.
Resolution (JCGM 200:2012, VIM 4.14): Smallest change in a quantity being measured that causes a perceptible change in the corresponding indication.
Resolution of a Displaying Device (JCGM 200:2012, VIM 4.15): smallest difference between displayed indications that can be meaningfully distinguished
Secondary force standard (ASTM E74): an instrument or mechanism, the calibration of which has been established by comparison with primary force standards.
Static Error Band (SEB): The band of maximum deviations of the ascending and descending calibration points from a best-fit line through zero output. it includes the effects of nonlinearity, hysteresis, and non-return to minimum load. normally expressed in units of % of FS.
VIM definition of Reproducibility 2.24 (3.7, Note 2) reproducibility condition of measurement reproducibility condition of measurement, out of a set of conditions that includes different locations, operators, measuring systems, and replicate measurements on the same or similar objects NOTE 1 The different measuring systems may use different measurement procedures. NOTE 2 A specification should give the conditions changed and unchanged, to the extent practical.

## Uncertainty Worksheet
Morehouse Measurement Uncertainty Calibration and Measurement Capability Worksheet
Start on this sheet and fill in only the light grey boxes. Choose a drop down option for the dark grey boxes.
Section 1: Data Entry
Laboratory | Morehouse | [11 empty] | Ref Standard Stability |  |  |  |  | Temperature |  | This is used when the points tested for stability
Technician Initials |  |  |  | Enter Information into these color cells | [8 empty] | FORCE | Change From | Interpolation | Actual |  | Effect |  | do not match the points calibrated
Date: |  |  |  |  | Initial Setup and after new calibration | [7 empty] | APPLIED | Previous % | Value | LBF |  | 1.5e-05 |  | FORCE | SLOPE | INTERCEPT | STABILITY | Interp | 0 | 0
Applied Maximum |  |  |  |  | Initial Setup | [6 empty] | 1 | [5 empty] | 0 |  | 0 |  |  |  |  | 0
Standards Used (Ref and UUT) |  |  |  | Uncertainty Worksheet and R & R Between Techs Need to Filled Out for CMC Summary to be correct | [7 empty] | 2 | [5 empty] | 0 |  | 0
Force or Torque/Subrange | [11 empty] | 3 | [5 empty] | 0 |  | 0
Resolution UUT |  | LBF | [9 empty] | 4 | [5 empty] | 0 |  | 0 | [6 empty] | ASTM E74 LLF | * This is your ASTM E74 LLF Found on Your ASTM E74 Report. It will be converted to a pooled std dev
[12 empty] | 5 | [5 empty] | 0 |  | 0 | [6 empty] | ISO 376 Uncertainty | Leave this section blank and fill in Uncertainty Coefficients In Cell N22 From Your Calibration Report or ISO 376 Annex C
Reference Standard Information | [11 empty] | 6 | [5 empty] | 0 |  | 0 | [6 empty] | Other (Leave Blank) | If non ASTM E74 or ISO 376 use this field & use nonlinearity or SEB if making ascending and descending measurements
ASTM E74 LLF |  | LBF | * This is your ASTM E74 LLF Found on Your ASTM E74 Report. It will be converted to a pooled std dev | [8 empty] | 7 | [5 empty] | 0 |  | 0
Resolution of Reference Standard |  | LBF | [9 empty] | 8 | [5 empty] | 0 |  | 0
Temperature Spec per degree C % | 0.0015% | [10 empty] | 9 | [5 empty] | 0 |  | 0
[12 empty] | 10 | [5 empty] | 0 |  | 0
Max Temperature Variation | [11 empty] | 11 | [5 empty] | 0 |  | 0
per degree C of Environment | 1 |  | * This is set for a Morehouse load cell based on the specification sheet | [8 empty] | 12 | [5 empty] | 0 |  | 0
Morehouse CMC (Ref Lab) | 0.0016% |  | Optional: If entering Engineering Units, leave blank and complete the Ref Laboratory Uncertainty Per Point (P29 - 40) | [9 empty] | ISO 376 Uncertainty Coefficients
[13 empty] | C0 | C1 | C2
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | % | Leave Blank if Not Used | [15 empty] | YES |  | %
Miscellaneous Error | 0.003 | % | Leave Blank if Not Used (Select % or Eng. Units if a constant) | [9 empty] | Expanded Uncertainty = C0 + (C1 * F) + (C2 * F)^2 | [5 empty] | NO |  | Eng. Units
[13 empty] | Where F = Force Applied, C0 = Intercept, C1 = Slope | [7 empty] | ISO
Conv Repeatability Data To Eng. Units | YES
| Repeatability of UUT | [11 empty] | Ref Laboratory Uncertainty Per Point |  |  |  |  | MUST SELECT |  |  | ISO 376 | UNCERTAINTY
| Applied | Run1 | Run2 | Run3 | Run4 | Average | Resolution | STD DEV | CONVERTED |  |  |  | Force | % | Eng. Units | Conv % | Force | % or Eng. |  |  | Force | LBF
1 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 1 | 0
2 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 2 | 0
3 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 3 | 0
4 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 4 | 0
5 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 5 | 0
6 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 6 | 0
7 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 7 | 0
8 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 8 | 0
9 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 9 | 0
10 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 10 | 0
11 | [13 empty] | 0.0016% |  | 1.6e-05 |  | % |  | 11 | 0
12 | [13 empty] | 0.0016% |  | 1.6e-05 | 0 | % |  | 12 | 0
Avg Std Dev of Runs
Section 2: Data Analysis
Measurement Uncertainty Budget Worksheet Data Analysis (Choose a drop down option for the dark grey boxes.) | [9 empty] | Expanded (68.26% k=1) | 1
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 |  |  |  | Miscellaneous Error |  | %
Parameter | FORCE | Range | 0 | Sub-Range | N/A |  |  |  |  | Expanded (95.45% k=2) | 2 |  |  |  | Force | % | Reading
Technician | 0 | Standards Used | 0 | [6 empty] | Expanded (99% k=2.58) | 2.578 |  |  | 1 | 0 | 0
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 |  |  | 2 | 0 | 0
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0 |  |  | 3 | 0 | 0
Reproducibility & Rep ( R & R sheet) | #VALUE! | A | None | 0 | -1 |  |  |  |  | Normal | 1 |  |  | 4 | 0 | 0
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 |  |  | 5 | 0 | 0
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 |  |  | 6 | 0 | 0
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495 |  |  | 7 | 0 | 0
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0 |  |  |  |  | 8 | 0 | 0
Stability of Ref Standard | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0 |  |  |  |  | 9 | 0 | 0
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 |  |  |  |  | 10 | 0 | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) | 0 | B | Rectangular | 1.7321 |  | 0 | 0 | [6 empty] | 11 | 0 | 0
Miscellaneous Error | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0 |  |  |  |  | 12 | 0 | 0
Morehouse CMC (Ref Lab) | 0 | B | Expanded (95.45% k=2) | 2 | 200 | 0 | 0 |  | 0
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
|  |  | Effective Degrees of Freedom |  |  |  | Formula | [6 empty] | The grey column Ref CMC is what
NOTE: ONLY ENTER INFO IN GREY BOXES IN SECTION 1 UNLESS CHANGING DISTRIBUTION |  |  | Coverage Factor (k) = | [10 empty] | populates individual sheets
|  |  | Expanded Uncertainty (U) = |  |  |  | #VALUE!
Slope Regression Worksheet | [13 empty] | Applied | Ref CMC | 0
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Error | Calculated | Residuals
1 | 0 | [13 empty] | 1.6e-05
2 | [14 empty] | 1.6e-05
3 | [14 empty] | 1.6e-05
4 | [14 empty] | 1.6e-05
5 | [14 empty] | 1.6e-05
6 | [14 empty] | 1.6e-05
7 | [14 empty] | 1.6e-05
8 | [14 empty] | 1.6e-05
9 | [14 empty] | 1.6e-05
10 | [14 empty] | 1.6e-05
11 | [14 empty] | 1.6e-05
12
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
Slope |  | #DIV/0! | [12 empty] | Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | %
Intercept |  | #DIV/0! | [12 empty] | Force | LBF | Reading
STEYX (k=1) |  | #DIV/0! | [11 empty] | 1 |  |  | 0%
[14 empty] | 2 |  |  | 0%
[14 empty] | 3 |  |  | 0%
[14 empty] | 4 |  |  | 0%
[14 empty] | 5 |  |  | 0%
[14 empty] | 6 |  |  | 0%
[14 empty] | 7 |  |  | 0%
[14 empty] | 8 |  |  | 0%
[14 empty] | 9 |  |  | 0%
[14 empty] | 10 |  |  | 0%
[14 empty] | 11 |  |  | 0%
[14 empty] | 12 |  |  | 0%
Note on A1: This workbook is not protected and was created by Henry Zumbrun (www.mhforce.com) and Dilip Shah of E=Mc3 solutions. This sheet is intended to help labs figure out their CMC's for force or torque measurement. Morehouse and E=Mc3 offer training course twice a year. More information can be found at https://mhforce.com/Training/TrainingCourses?tab=Training%20Courses
Note on X4: This part calculates interpolated stability values. Your should enter the values from the calibration certificate. This sheet is set up to find interpolated values, if different calibration points are used for repeatability
Note on B10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is the resolution we used for our repeatability study load cell
Note on A13: Select between three standards - ASTM E74 fill in LLF - ISO 376 Leave Blank (Fill in N22, O22, P22) - Other Leave Blank (Fill in A22)
Note on B23: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Used in UCM with less than 1/16 of an inch misalignment. 0.0625 * 0.0005 (load cell spec sheet) Take 1/2 of the specification from the LC spec sheet
Note on N27: Use % if the CMC of th reference lab is % of applied. If the reference lab reported CMC in Engineering Units, you will need to fill in the values per point and select Eng. Units in the cell on the right
Note on Z27: ISO 376 equation is often a combined uncertainty. This formula reduces that value to a standard uncertainty.
Note on B29: Ref Standard Stability (top right) should be populated first. The same number that is used for reference standard stability should match these numbers
Note on I41: Many times, when a laboratory needs to look at the "average" of the standard deviations across the many points of data, they will look at the POOLED STANDARD DEVIATION. The way to calculate POOLED STANDARD DEVIATION is to take the AVERAGE of the SUM of the SQUARE of the STANDARD DEVIATIONS (i.e. VARIANCE) and DIVIDE it by the Number (COUNT) of the data points. Many labs will erroneously calculate the AVERAGE of the STANDARD DEVIATIONS which will be incorrect. and not an acceptable statistical practice.
Note on A45: This Worksheet is the intellectual Property of E= mc3 Solutions and Morehouse Instrument Company. Permission is granted to the E=mc3 Solution Clients and Morehouse Instrument Company Clients. All others should contact at emc3solu@aol.com (call 330-328-4400) for purchasing a license.
Note on A51: Henry_000: CMC summary uses a per point analysis found on sheets 1-12 that includes both Repeatability and Reproducibility separated. This sheet gives an overall uncertainty snapshot and uses both repeatability and reproducibility from R & R Between Techs)
Note on F53: This number is the number of calibration points minus 1. It is the Sample Size minus 1 and is Type A Data since it used the standard deviation of a number of test points.
Note on G64: This workbook was developed by Dilip Shah of E=mc3 solutions and Henry Zumbrun of Morehouse Instrument Company. mhforce.com

## R & R Between Techs
Repeatability and Reproducibility Worksheet for Measurement Uncertainty Budget Worksheet | [9 empty] | Normal (68.26%, k=1) | 1 | Y | 68.27% | Normal (68.26%, k=1) | 1
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | N | 95% | None | 0
Parameter | FORCE | Range | 0 | Sub-Range | N/A |  |  |  |  | Expanded (95.45% k=2) | 2 |  | 95.45%
Technician | 0 | Standards Used |  | 0 | [5 empty] | Expanded (98.36% k=2.4) | 2.4 |  | 99%
Date | 0 | For Student's t correction input "Y" ->> |  |  | Specific Divisor |  | 3.4 |  |  | Expanded (99% k=2.58) | 2.578 |  | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | Expanded (99.73% k=3) | 3
Repeatability Between Techs |  | A | None | 0 | 0 |  |  |  |  | Specific Divisor | 3.4
Reproducibility Between Techs |  | A | Normal (68.26%, k=1) | 1 | -1 |  |  |  |  | None | 0
|  |  | None | 0 | [5 empty] | U-Shaped (sqrt 2) | 1.4142
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular (sqrt 3) | 1.7321
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Triangular (sqrt 6) | 2.4495
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Resolution (sqrt 12) | 3.4641
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) | 0 | B | Rectangular | 1.7321 |  | 0 | 0
Miscellaneous Error | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Morehouse CMC (Ref Lab) | 0 | B | Expanded (95.45% k=2) | 2 | 200 | 0 | 0 |  | 0
|  |  | Combined Uncertainty (uc) = |  |  | 0 | 0 | 0% | 0
|  |  | Effective Degrees of Freedom |  |  |  | Effective Degrees of Freedom Formula
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
Expanded Uncertainty (U) =
Repeatability and Reproducibility Worksheet
|  |  |  | Technician 1 | Technician 2 | Technician 3 | Technician 4 | Technician 5 | Technician 6
1
2
3
4
5
6
7
8
9
|  |  | 10 | [11 empty] | YES
|  |  | 11 | [11 empty] | NO
12
Distribution | Divisor | 1/Divisor | 13
Rectangular | 1.7321 | 0.57735 | 14
Triangular | 2.4495 | 0.40825 | 15
U - Shaped | 1.4142 | 0.70711 | 16
Resolution of instrument | 3.4641 | 0.28868 | 17
Coverage Factor (k) | Confidence Level |  | 18
1 | 68.27 |  | 19
1.645 | 90 |  | 20
1.96 | 95 |  | Std. Dev.
2 | 95.45 |  | Average
2.576 | 99 |  | Variance
3 | 99.73 |  | Repeatability |  |  |  | #DIV/0! | #VALUE! | LBF
|  |  | Reproducibility |  |  |  |  | #VALUE! | LBF
|  |  | Std. Dev. Of the Mean |  | #DIV/0! | Convert to Eng Unit (Use Values Above) |  |  | YES | [6 empty] | Repeatability | #VALUE!
|  |  | dfNumerator |  |  | Force Value corresponding to R & R Output |  |  | 0 | [6 empty] | Reproducibility | #VALUE!
dfDenominator
This sheet is to be used to test the repeatability and reproducibility between technicians
| Excel Tab | Applied Force | Repeatability Between Techs | Reproducibility | [5 empty] | Repeatability | Reproducibility | Multiplier
| Sheet 1 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 2 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 3 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 4 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 5 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 6 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 7 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 8 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 9 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 10 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 11 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
| Sheet 12 | 0 | #VALUE! | #VALUE! | [7 empty] | #DIV/0!
Note on C5: Dilip Shah: Type A data in the first two rows will corrected to infinite degrees of freedom if "Y" is chosen.
Note on F5: Dilip Shah: Allows you to choose ONE Custom Divisor for the Distribution column
Note on A38: Dilip Shah: Use for specifications, or where no other information is available. It is the most conservative distribution.
Note on A39: Dilip Shah: Use for parameters that have a central tendency. Examples: Frequency measurements.
Note on A40: Dilip Shah: Use for Cyclical Type of data. Example: Temperature control
Note on A41: Dilip Shah: e.g. 0.0001 to be entered in Magnitude Column
Note on F48: Hank527: If you are not working in Engineering Units these values will need to be converted to like units. If these are in mV/V you will need to convert to force units. Formula to the right should do this for you

## CMC Summary
|  | Morehouse Measurement Uncertainty Calibration and Measurement Capability Worksheet | [13 empty] | Y
Measurement Uncertainty Budget Summary | [15 empty] | N
Laboratory | Morehouse
Parameter | FORCE | Range | 0 | Sub-Range | N/A
Technician |  | Standards Used | 0
Date | 0
| Applied | Expanded Uncertainty | Expanded Uncertainty % |  | Slope | Intercept |  | Fitted Curve | Difference | Enter Force | Estimated Expanded | Exclude Y/N | Sheet # | Applied | Expanded UNC (CMC)
1 | 0 |  | #VALUE! | [6 empty] | Value Below |  | N | 1 | 30000
2 |  |  | #VALUE! | [8 empty] | N | 2 | 50000
3 |  |  | #VALUE! | [8 empty] | N | 3 | 100000
4 | [11 empty] | Y | 4 | 200000
5 | [11 empty] | Y | 5 | 300000
6 | [11 empty] | Y | 6 | 400000
7 | [11 empty] | Y | 7 | 500000
8 | [11 empty] | Y | 8 | 600000
9 | [11 empty] | Y | 9 | 700000
10 | [11 empty] | Y | 10 | 800000
11 | [11 empty] | Y | 11 | 900000
12 | [11 empty] | Y | 12 | 1000000
[8 empty] | Max Difference | 0
Note: Force value should be entered between the segmented ranges above to calculate MU per point
Note: This is a summary sheet for all test points
|  |  |  | Unc Order of Fit | 1 | MAY NOT GIVE BEST FIT WITH ORDERS HIGHER THAN 1
Uncertainty Per Point Fit Coefficients
|  |  |  | a3= | 0
|  |  |  | a2= | 0
|  |  |  | a1= | #VALUE!
|  |  |  | a0= | 0
[17 empty] | 1st Try | Difference | Adjusted Intercept | Difference | Attempt 1 | Difference | Attempt 2 | Difference
|  | Applied | CMC | Best Fit | Difference |  |  | CMC Summary | [8 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | #N/A | #VALUE! | #VALUE!
|  | 0 |  | #VALUE! | #VALUE! |  |  | Sheet # | Applied | % | Constant Corrected | Orig % |  |  |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  |  | #VALUE! | #VALUE! |  |  | 1 | 0 | #VALUE! | #VALUE! | #VALUE! |  |  |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  |  | #VALUE! | #VALUE! |  |  | 2 |  | #VALUE! | #VALUE! | #VALUE!
[8 empty] | 3 |  | #VALUE! | #VALUE! | #VALUE!
4
5
6
7
8
9
10
11
12
| Non Corrected Formula |  | Corrected to Formula
Coefficients | A1 | #VALUE! | A1 | #VALUE! | A1 | #N/A | A1 | #VALUE! | A1 | #VALUE!
| A0 |  | A0 | #VALUE! | A0 |  | A0 |  | A0
CMC = ((APPLIED FORCE X SLOPE A1) + ADJUSTED INTERCEPT A0)
|  |  | A1 | #VALUE! |  |  |  | Desired % |  |  | 0.01%
|  |  | A0 | 0 |  |  |  | First Force Point that Matches Desired % |  |  | #VALUE!
Note: The corrected formula is used to make sure no numbers are underreported
|  |  | Max | #VALUE!
𝑈=𝑎_3 𝐹^3+𝑎_2 𝐹^2+𝑎_1 𝐹+𝑎_0
Note on A1: This workbook is not protected and was created by Henry Zumbrun (www.mhforce.com) and Dilip Shah of E=Mc3 solutions.
Note on Q7: ISO/IEC 17025 Accredited Calibrations. The International Laboratory Accreditation Cooperation (ILAC) P-14 section 6.4 requires contributions such as the resolution of the Unit Under Test (UUT) be included in the uncertainty per point value reported on the certificate of calibration. Euramet CG4 v2.0 Uncertainty of Force Measurements also requires the resolution of both the standard and unit under test be reported as standard contributors. The per point analysis consists of the UUT resolution and the CMC of our standard (which includes the resolution of the reference standard if primary standards were not used). When a full uncertainty analysis per ILAC P-14 is performed, as required by ISO/IEC 17025, the CMC (Calibration and Measurement Capability) is derived and reported to the user. There is a tab in this worksheet to help with calculating the CMC per point. Morehouse can provide guidance on using this workbook. It is recommended a professional consultant be contacted if you are having problems figuring out CMC or what to report. We offer training with a focus on measurement uncertainty with an independent consultant (Dilip Shah of E=mc3 solutions) two to three times per year. Information on the training offered can be found at http://www.mhforce.com/customer-education
Note on E26: Future Release may do higher order fit. Currently limited to straight line
Note on I42: ISO/IEC 17025 Accredited Calibrations. The International Laboratory Accreditation Cooperation (ILAC) P-14 section 6.4 requires contributions such as the resolution of the Unit Under Test (UUT) be included in the uncertainty per point value reported on the certificate of calibration. Euramet CG4 v2.0 Uncertainty of Force Measurements also requires the resolution of both the standard and unit under test be reported as standard contributors. The per point analysis consists of the UUT resolution and the CMC of our standard (which includes the resolution of the reference standard if primary standards were not used). When a full uncertainty analysis per ILAC P-14 is performed, as required by ISO/IEC 17025, the CMC (Calibration and Measurement Capability) is derived and reported to the user. There is a tab in this worksheet to help with calculating the CMC per point. Morehouse can provide guidance on using this workbook. It is recommended a professional consultant be contacted if you are having problems figuring out CMC or what to report. We offer training with a focus on measurement uncertainty with an independent consultant (Dilip Shah of E=mc3 solutions) two to three times per year. Information on the training offered can be found at http://www.mhforce.com/customer-education

## En For PT
PROFICIENCY TEST SHEET
Date | 0 | ILC Sheet for figuring En Ratios Use this Sheet for PT tests
Morehouse
|  | Reference |  | Test Lab |  |  | Reference | Test Lab
| Lab or Machine | M-4644 |  | M-8407 |  |  | M-4644 | M-8407
Reduced Data
| Force Applied | REF | REF % | LAB | LAB % |  | REF | LAB
| LBF | VALUE | UNC PER POINT | VALUE | UNC PER POINT |  | VALUE | VALUE | Force Applied | Difference | Difference | Lab Uncertainty | Ref Uncertainty | En | Pass/Fail
Test Point # | Beg 0 | 0 |  | 0 |  | Beg - Beg 0 | 0 | 0 | LBF | Counts | LBF | LBF | LBF
1 | 500 | -0.20547 | 0.0008% | -0.20546 | 0.0008% |  | -0.20546 | -0.20546 | 500 | 3.3333e-06 | 0.0081276 | 0.016192 | 0.016192 | 0.35493 | PASS
2 | 2500 | -1.0208 | 0.0008% | -1.0208 | 0.0008% |  | -1.0208 | -1.0208 | 2500 | 1.6667e-05 | 0.040638 | 0.042405 | 0.042405 | 0.67764 | PASS
3 | 5000 | -2.0557 | 0.0008% | -2.0557 | 0.0008% |  | -2.0556 | -2.0557 | 5000 | 3e-05 | 0.073148 | 0.081229 | 0.081229 | 0.63676 | PASS
4
5
6
7
8
9
10
11
12
| End 0 | -4e-05 |  | 0 |  | Beg - End 0 | -4e-05 | 0
[6 empty] | End 0 | 0 | 0
| # of test point | 3
| Resolution | -2438.2713
| Readability | 1e-05
| Resolution | LBF |  | 0.024383
|  | K=2 | k=1
| Lab Exp Unc | 0.0016% | 0.0008%
Morehouse
| Ref Exp Unc | 0.0016% | 0.0008%
| Notes: | Must have at least 3 test points for this to work. Maximum number of pts is set to 12.
This sheet is to be used for direct Intra Laboratory Comparisons
It is recommended to have a force measuring instrument calibrated by Primary Standards for X ref
This sheet would be used once you have calculated your CMC's. Morehouse can calibrate an artifact sent to us, or supply (rent) an artifact for you to compare the results you get against our results.
This would be the final test to prove you are making measurements within your claimed measurement capability. Contact us at info@mhforce.com
Note:
Note on C2: Hank527: This sheet was created by Henry Zumbrun and is property of Morehouse Instrument Company www.mhforce.com Use at your own risk
Note on B50: Does the Accreditation Body (AB) require Proficiency Testing (PT)? The En for PT sheet is designed to help meet this requirement. We strongly recommend using an Ultra Precision Load Cell that is very stable and repeatable when used with a higher end meter. Not doing so will have a negative impact on the CMC. This load cell can double as an artifact used for SPC testing and will strengthen the measurement assurance process. To meet PT requirements, have this artifact calibrated by dead weight primary standards. Enter the calibration data for this artifact in the Ref Value Cells and enter the calibration data for the same artifact when calibrated by comparison with a secondary standard (reference load cell). Enter the Lab Unc (CMC) against the Ref Unc (CMC) and the sheet will calculate the En ratio and indicate a "pass" or a "fail". Per ISO 17025 Section 7.7 Ensuring the validity of the results 7.7.2 The laboratory shall monitor its performance by comparison with results of other laboratories, where available and appropriate. This monitoring shall be planned and reviewed and shall include, but not be limited to, either or both of the following: a) participation in proficiency testing; NOTE ISO/IEC 17043 contains additional information on proficiency tests and proficiency testing providers. Proficiency testing providers that meet the requirements of ISO/IEC 17043 are considered to be competent) participation in interlaboratory comparisons other than proficiency testing. 7.7.3 Data from monitoring activities shall be analyzed, used to control and, if applicable, improve the laboratory's activities. If the results of the analysis of data from monitoring activities are found to be outside pre-defined criteria, appropriate action shall be taken

## Calc Resoution Ex
How to Calculate Resolution of a Force-Measuring Device
Step 1 | Enter Force Applied | 5000
Step 2 | Enter Output of Instrument Corresponding to Force Applied | 1.9969
| Resolution | 2503.881
Step 3 | Multiply by readability | 0.00012
Step 4 | Count By | 1
| Resolution per count | 0.30047
Examples
| Example 1: | Device Reads in mV/V & Counts by 1
| Force Applied | Measured Output of UUT | Readability of UUT | Count by
| 10000 | 4.1124 | 1e-05 | 1
| Resolution = | 0.024317
| Example 2: | Device Reads in lbf & Counts by 2
| Force Applied | Measured Output of UUT | Readability of UUT | Count by
| 10000 | 9995.9 | 0.1 | 2
| Resolution = | 0.20008
If we have multiple force points, we should then take the average resolution is shown in example 2
| Example 3: | Multiple Points mV/V
| Force Applied | Measured Output of UUT | Readability of UUT | Count by | Resolution
| 2000 | 0.82199 | 1e-05 | 1 | 0.024331
| 4000 | 1.6432 | 1e-05 | 1 | 0.024343
| 6000 | 2.4674 | 1e-05 | 1 | 0.024317
| 8000 | 3.2863 | 1e-05 | 1 | 0.024343
| 10000 | 4.1124 | 1e-05 | 1 | 0.024317
|  |  | Average Resolution |  | 0.02433

## 1
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error |  |  |  | #N/A
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2 | 200
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | 0 | [7 empty] | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 1

## 2
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
#DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 2

## 3
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 3

## 4
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 4

## 5
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 5

## 6
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 6

## 7
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 7

## 8
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 8

## 9
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 9

## 10
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 10

## 11
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs | #VALUE! | A | None | 0 | 0
Reproducibility Between Techs | #VALUE! | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 11

## 12
Measurement Uncertainty Budget Worksheet | [9 empty] | Expanded (68.26% k=1) | 1 | 68.27%
Laboratory | Morehouse | [8 empty] | Expanded (95% k=1.96) | 1.96 | 95%
Parameter | FORCE | Range | 0 | Sub-Range | [5 empty] | Expanded (95.45% k=2) | 2 | 95.45%
Technician | HZ | Standards Used | [7 empty] | Expanded (99% k=2.58) | 2.578 | 99%
Date | 0 | [8 empty] | Expanded (99.73% k=3) | 3 | 99.73%
Uncertainty Contributor | Magnitude | Type | Distribution | Divisor | df | Std. Uncert | Variance (Std. Uncert^2) | % Contribution | u^4/df | None | 0
Repeatability Between Techs |  | A | None | 0 | 0
Reproducibility Between Techs |  | A | Normal (68.26%, k=1) | 1 | -1
[10 empty] | Normal | 1
Repeatability | #DIV/0! | A | Normal | 1 | -1 |  |  |  |  | Rectangular | 1.7321 | #VALUE! | 1000
ASTM E74 LLF | 0 | A | Normal | 1 | 32 | 0 | 0 |  | 0 | Resolution | 3.4641 | #DIV/0! | -1
Resolution of UUT | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0 | Triangular | 2.4495
Environmental Conditions | 0 | B | Rectangular | 1.7321 | 200 | 0 | 0 |  | 0
Stability of Ref Standard |  | B | Rectangular | 1.7321 | 200
Ref Standard Resolution | 0 | B | Resolution | 3.4641 | 200 | 0 | 0 |  | 0
Non ASTM or ISO 376 (Tolerance,Non-Linearity,SEB) |  | B | Rectangular | 1.7321 | 200
Miscellaneous Error
Morehouse CMC (Ref Lab) | #VALUE! | B | Expanded (95.45% k=2) | 2
|  |  | Combined Uncertainty (uc)= |  |  | 0 | 0 | 0% | 0
Effective Degrees of Freedom
|  |  | Coverage Factor (k), Confidence Interval = |  | 95.45%
|  |  | Expanded Uncertainty (U) K = |  |  |  | #VALUE!
Slope Regression Worksheet
| Applied | Run 1 | Run 2 | Run 3 | Run 4 | Average | Std. Dev. | Ref CMC | LBF
1 | [7 empty] | 0.0016% | #VALUE!
Repeatability (Of Error) |  |  | Average Standard Deviation of Runs |  |  | #DIV/0!
This sheet is to be used as the master form to calculate CMC for PT 12