## Sample Size Selector
Resolution/Repeatbility/CMC Risk Analyzer | [12 empty] | Make Repeatability Better |  |  | Yes |  |  | Regenerate | Y | 1
User Inputs (Light Yellow Background Cells with Red Text) |  |  | Make Repeatability Better (Tolerance/#) |  |  |  | 100 |  |  | Raw Data | Regenerate Random Raw Data |  |  |  |  | No |  |  | Raw Data | N | 2
Nominal Value | 100 | Tolerance (%) |  | 0.1% | + No of Counts |  | 0 | [13 empty] | 3
CMC (k=2) | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | [8 empty] | 0.1 |  |  |  |  | 4
k = | 2 | 2 | 2 | 2 | 2 | 2 | 2 | [8 empty] | 0.001 |  |  |  |  | 5
Sample Size (Maximum 30) | 19 | For Student's t correction input "Y" |  |  | Y | df | 18 | [13 empty] | 6
Starting Resolution and increments | 0.001 | 0.01 | 0.02 | 0.05 | 0.1 | 0.2 | 0.5 |  |  |  | Graph Resolution and Uncertainty | Expanded Uncertainty (95% C.I.) | 0.0027788 | 0.0059976 | 0.011671 | 0.028948 | 0.057829 | 0.11563 | 0.28904 |  | 7
Measured Value (Average of Sample Size) | 99.9997 | 100 | 100 | 100 | 100 | 100 | 100 |  |  |  |  | Starting Resolution and increments | 0.001 | 0.01 | 0.02 | 0.05 | 0.1 | 0.2 | 0.5 |  | 8
Tolerance due to the No. of Counts | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  |  | TUR | 36.0093 | 16.6914 | 8.5783 | 3.4588 | 1.7314 | 0.86594 | 0.3464 |  | 9
Tolerance | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | [13 empty] | 10
Tolerance in Units (including Counts if any) | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | [13 empty] | 11
Lower specification Limit | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | [13 empty] | 12
Upper Specification Limit | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | [13 empty] | 13
CMC (k=1) | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | [13 empty] | 14
Repeatability Uncertainty | 0.0010976 | 0 | 0 | 0 | 0 | 0 | 0 | [13 empty] | 15
Resolution Uncertainty | 0.00028868 | 0.0028868 | 0.0057735 | 0.014434 | 0.028868 | 0.057735 | 0.14434 | [13 empty] | 16
df - CMC (k=1) | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | [13 empty] | 17
df - Repeatability Uncertainty | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | [13 empty] | 18
df - Resolution Uncertainty | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | [13 empty] | 19
Combined Variance (RSS) | 1.928e-06 | 8.9733e-06 | 3.3973e-05 | 0.00020897 | 0.00083397 | 0.003334 | 0.020834 | [13 empty] | 20
Combined Uncertainty (k=1) | 0.0013885 | 0.0029956 | 0.0058287 | 0.014456 | 0.028879 | 0.057741 | 0.14434 | [13 empty] | 21
Effective Degrees of Freedom | 1990 | 1152 | 1038 | 1006 | 1001 | 1000 | 1000 | [13 empty] | 22
k95% = | 2.0013 | 2.0022 | 2.0024 | 2.0025 | 2.0025 | 2.0025 | 2.0025 | [13 empty] | 23
Expanded Uncertainty (95% C.I.) | 0.0027788 | 0.0059976 | 0.011671 | 0.028948 | 0.057829 | 0.11563 | 0.28904 | [13 empty] | 24
TUR | 36.0093 | 16.6914 | 8.5783 | 3.4588 | 1.7314 | 0.86594 | 0.3464 | [13 empty] | 25
u^4/df - CMC (k=1) | 4.096e-16 | 4.096e-16 | 4.096e-16 | 4.096e-16 | 4.096e-16 | 4.096e-16 | 4.096e-16 | [13 empty] | 26
u^4/df - Repeatability Uncertainty | 1.4512e-15 | 0 | 0 | 0 | 0 | 0 | 0 | [13 empty] | 27
u^4/df - Resolution Uncertainty | 6.9444e-18 | 6.9444e-14 | 1.1111e-12 | 4.3403e-11 | 6.9444e-10 | 1.1111e-08 | 4.3403e-07 | [13 empty] | 28
CMC Percent Contribution | 33.19% | 7.132% | 1.884% | 0.3063% | 0.07674% | 0.0192% | 0.003072% | [13 empty] | 29
Repeatability Percent Contribution | 62.48% | 0% | 0% | 0% | 0% | 0% | 0% | [13 empty] | 30
Resolution Percent Contibution | 4.322% | 92.87% | 98.12% | 99.69% | 99.92% | 99.98% | 100%
Total Contribution | 100% | 100% | 100% | 100% | 100% | 100% | 100%
Nominal Value | 100 | 100 | 100 | 100 | 100 | 100 | 100
Lower Specification Limit | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9
Upper Specification Limit | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1
Measured Value | 99.9997 | 100 | 100 | 100 | 100 | 100 | 100
Combined Uncert. (k=1) | 0.0013885 | 0.0029956 | 0.0058287 | 0.014456 | 0.028879 | 0.057741 | 0.14434
Total Risk | 0% | 1.217e-242% | 2.806e-64% | 4.594e-10% | 0.05346% | 8.329% | 48.84%
Upper Limit Risk | 0% | 1.217e-242% | 2.806e-64% | 2.297e-10% | 0.02673% | 4.165% | 24.42%
Lower Limit Risk | 0% | 0% | 0% | 2.297e-10% | 0.02673% | 4.165% | 24.42%
TUR | 36.0093 | 16.6914 | 8.5783 | 3.4588 | 1.7314 | 0.86594 | 0.3464
CPK | 23.9431 | 11.1276 | 5.7189 | 2.3059 | 1.1543 | 0.57729 | 0.23094
Randomly Generated RAW Data
Resolution | 0.001 | 0.01 | 0.02 | 0.05 | 0.1 | 0.2 | 0.5
100.0014 | 100.001 | 100 | 100 | 100 | 100 | 100 | 100
99.9997 | 100 | 100 | 100 | 100 | 100 | 100 | 100
99.9995 | 99.999 | 100 | 100 | 100 | 100 | 100 | 100
99.9996 | 100 | 100 | 100 | 100 | 100 | 100 | 100
99.9989 | 99.999 | 100 | 100 | 100 | 100 | 100 | 100
100.0006 | 100.001 | 100 | 100 | 100 | 100 | 100 | 100
99.9991 | 99.999 | 100 | 100 | 100 | 100 | 100 | 100
99.9995 | 99.999 | 100 | 100 | 100 | 100 | 100 | 100
99.9995 | 99.999 | 100 | 100 | 100 | 100 | 100 | 100
99.9994 | 99.999 | 100 | 100 | 100 | 100 | 100 | 100
99.9996 | 100 | 100 | 100 | 100 | 100 | 100 | 100
100.0003 | 100 | 100 | 100 | 100 | 100 | 100 | 100
100.0007 | 100.001 | 100 | 100 | 100 | 100 | 100 | 100
99.9991 | 99.999 | 100 | 100 | 100 | 100 | 100 | 100
99.9972 | 99.997 | 100 | 100 | 100 | 100 | 100 | 100
100.0002 | 100 | 100 | 100 | 100 | 100 | 100 | 100
100.0015 | 100.002 | 100 | 100 | 100 | 100 | 100 | 100
100.0003 | 100 | 100 | 100 | 100 | 100 | 100 | 100
100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
100.0006 | 100.001 | 100 | 100 | 100 | 100 | 100 | 100
99.9997 | 100 | 100 | 100 | 100 | 100 | 100 | 100
100.0009 | 100.001 | 100 | 100 | 100 | 100 | 100 | 100
100.0014 | 100.001 | 100 | 100 | 100 | 100 | 100 | 100
100.0001 | 100 | 100 | 100 | 100 | 100 | 100 | 100
99.9988 | 99.999 | 100 | 100 | 100 | 100 | 100 | 100
100.0015 | 100.001 | 100 | 100 | 100 | 100 | 100 | 100
100.0022 | 100.002 | 100 | 100 | 100 | 100 | 100 | 100
100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
99.9989 | 99.999 | 100 | 100 | 100 | 100 | 100 | 100
99.9987 | 99.999 | 100 | 100 | 100 | 100 | 100 | 100 | The two colums below are just for validating can be romoved after checking.
Standard Deviation of Data Set | 0.0010807 | 0 | 0 | 0 | 0 | 0 | 0
Standard Deviation Calculation
df | 0.001 | 0.01 | 0.02 | 0.05 | 0.1 | 0.2 | 0.5 | N | Y
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 0.00070711 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00070711 | 0.0012988
2 | 0.001 | 0 | 0 | 0 | 0 | 0 | 0 | 0.001 | 0.0018367
3 | 0.0008165 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0008165 | 0.0014997
4 | 0.00083666 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00083666 | 0.0015367
5 | 0.00089443 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00089443 | 0.0016428
6 | 0.00089974 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00089974 | 0.0016526
7 | 0.00088641 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00088641 | 0.0016281
8 | 0.00086603 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00086603 | 0.0015906
9 | 0.00084327 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00084327 | 0.0015489
10 | 0.00080904 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00080904 | 0.001486
11 | 0.0007785 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0007785 | 0.0014299
12 | 0.00083205 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00083205 | 0.0015282
13 | 0.00082542 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00082542 | 0.0015161
14 | 0.0010601 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010601 | 0.0019471
15 | 0.0010308 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010308 | 0.0018932
16 | 0.00116 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00116 | 0.0021306
17 | 0.0011275 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0011275 | 0.0020709
18 | 0.0010976 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010976 | 0.0020159
19 | 0.001105 | 0 | 0 | 0 | 0 | 0 | 0 | 0.001105 | 0.0020296
20 | 0.0010779 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010779 | 0.0019798
21 | 0.0010821 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010821 | 0.0019876
22 | 0.0010835 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010835 | 0.00199
23 | 0.0010598 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010598 | 0.0019466
24 | 0.0010536 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010536 | 0.0019351
25 | 0.0010554 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010554 | 0.0019385
26 | 0.0011094 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0011094 | 0.0020377
27 | 0.0010887 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010887 | 0.0019996
28 | 0.0010851 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010851 | 0.0019929
29 | 0.0010807 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010807 | 0.0019849
Average Calculation for Nominal Value
Sample | 0.001 | 0.01 | 0.02 | 0.05 | 0.1 | 0.2 | 0.5
1 | 100.001 | 100 | 100 | 100 | 100 | 100 | 100
2 | 100.0005 | 100 | 100 | 100 | 100 | 100 | 100
3 | 100 | 100 | 100 | 100 | 100 | 100 | 100
4 | 100 | 100 | 100 | 100 | 100 | 100 | 100
5 | 99.9998 | 100 | 100 | 100 | 100 | 100 | 100
6 | 100 | 100 | 100 | 100 | 100 | 100 | 100
7 | 99.9999 | 100 | 100 | 100 | 100 | 100 | 100
8 | 99.9998 | 100 | 100 | 100 | 100 | 100 | 100
9 | 99.9997 | 100 | 100 | 100 | 100 | 100 | 100
10 | 99.9996 | 100 | 100 | 100 | 100 | 100 | 100
11 | 99.9996 | 100 | 100 | 100 | 100 | 100 | 100
12 | 99.9997 | 100 | 100 | 100 | 100 | 100 | 100
13 | 99.9998 | 100 | 100 | 100 | 100 | 100 | 100
14 | 99.9997 | 100 | 100 | 100 | 100 | 100 | 100
15 | 99.9995 | 100 | 100 | 100 | 100 | 100 | 100
16 | 99.9996 | 100 | 100 | 100 | 100 | 100 | 100
17 | 99.9997 | 100 | 100 | 100 | 100 | 100 | 100
18 | 99.9997 | 100 | 100 | 100 | 100 | 100 | 100
19 | 99.9997 | 100 | 100 | 100 | 100 | 100 | 100
20 | 99.9998 | 100 | 100 | 100 | 100 | 100 | 100
21 | 99.9998 | 100 | 100 | 100 | 100 | 100 | 100
22 | 99.9999 | 100 | 100 | 100 | 100 | 100 | 100
23 | 99.9999 | 100 | 100 | 100 | 100 | 100 | 100
24 | 99.9999 | 100 | 100 | 100 | 100 | 100 | 100
25 | 99.9999 | 100 | 100 | 100 | 100 | 100 | 100
26 | 99.9999 | 100 | 100 | 100 | 100 | 100 | 100
27 | 100 | 100 | 100 | 100 | 100 | 100 | 100
28 | 100 | 100 | 100 | 100 | 100 | 100 | 100
29 | 100 | 100 | 100 | 100 | 100 | 100 | 100
30 | 99.9999 | 100 | 100 | 100 | 100 | 100 | 100
Note on I4: henry zumbrun: By default the repeatability is randomized and set to the tolerance. Enter a number to divide the tolerance by to make repeatability better
Note on M4: henry zumbrun: This will regenerate the randomized raw data. A macro would have been better though it creates issues for some
Note on I5: Dilip Shah: Resolution is to multiply by the Number of Counts to add to te tolerance. This may increase the tolerance and TUR. Used when the Tolerance might be a % ± # of counts