## Data Entry
Data Entry Page
Company |  |  |  |  | Uncertainty Calculation | [8 empty] | Compression
Company Name | Calibrations R Us |  | Enter Information and Data From Morehouse Certificate |  | uc | U | U | U | U |  |  |  | lbf | Tension
Date | 2022-04-20 |  |  |  | u | k = 2 | Rounded | As a % | Response Units |  |  |  | kgf
Instrument Type | Load Cell |  |  |  | 0.066536 | 0.13307 | 0.13 | 0.0026% | 1.04e-05 |  |  |  | N
Instrument Serial Number | U-Sample |  |  |  | 0.11748 | 0.23496 | 0.23 | 0.0023% | 1.84e-05
Meter Serial Number | MY025154 |  |  |  | 0.17154 | 0.34309 | 0.34 | 0.002267% | 2.7201e-05
Force Units | N |  |  |  | 0.2265 | 0.453 | 0.45 | 0.00225% | 3.6001e-05
Mode Type | Tension |  |  |  | 0.28183 | 0.56366 | 0.56 | 0.00224% | 4.4801e-05
[5 empty] | 0.33735 | 0.67469 | 0.67 | 0.002233% | 5.3601e-05
Uncertainty |  |  |  |  | 0.39297 | 0.78594 | 0.79 | 0.002257% | 6.3202e-05
Ref Exp Unc in a % or numerical form? |  | Percentage |  |  | 0.44867 | 0.89733 | 0.9 | 0.00225% | 7.2002e-05
Ref Exp if % Unc k = 2 | 0.002% |  |  |  | 0.50441 | 1.0088 | 1.01 | 0.002244% | 8.0802e-05
mV/V % Unc k = 2 | 0.001% |  |  |  | 0.56018 | 1.1204 | 1.12 | 0.00224% | 8.9602e-05
Readability | 1e-05
ASTM LLF (From Cert)
Round SIG FIG | 2
Enter Force Applied and Reduced Run Data From Certificate
Tension Uncertainty Analysis at the Time of Calibration
|  | Fitted Curve or | Force Applied | Force Applied | Force Applied | mV/V |  | Resolution |  |  | ASTM llf
| Force Applied | Average or |  |  | uw |  | uv |  |  | ures |  | ur
Test Point # | N | Measured Output | k = 2 | k = 2 | k = 1 | k = 1 | k = 1 | UUT Res | Readability | Res/3.464 | k = 2.4 | k = 1
1 | 5000 | 0.40001 | 0.002% |  | 0.05 | 0.0005% | 0.025 | 12499.6875 | 1e-05 | 0.036085 | 0 | 0
2 | 10000 | 0.80002 | 0.002% |  | 0.1 | 0.0005% | 0.05 | 12499.6875 | 1e-05 | 0.036085 | 0 | 0
3 | 15000 | 1.2 | 0.002% |  | 0.15 | 0.0005% | 0.075 | 12499.6875 | 1e-05 | 0.036085 | 0 | 0
4 | 20000 | 1.6 | 0.002% |  | 0.2 | 0.0005% | 0.1 | 12499.6875 | 1e-05 | 0.036085 | 0 | 0
5 | 25000 | 2 | 0.002% |  | 0.25 | 0.0005% | 0.125 | 12499.6875 | 1e-05 | 0.036085 | 0 | 0
6 | 30000 | 2.4001 | 0.002% |  | 0.3 | 0.0005% | 0.15 | 12499.6875 | 1e-05 | 0.036085 | 0 | 0
7 | 35000 | 2.8001 | 0.002% |  | 0.35 | 0.0005% | 0.175 | 12499.6875 | 1e-05 | 0.036085 | 0 | 0
8 | 40000 | 3.2001 | 0.002% |  | 0.4 | 0.0005% | 0.2 | 12499.6875 | 1e-05 | 0.036085 | 0 | 0
9 | 45000 | 3.6001 | 0.002% |  | 0.45 | 0.0005% | 0.225 | 12499.6875 | 1e-05 | 0.036085 | 0 | 0
10 | 50000 | 4.0001 | 0.002% |  | 0.5 | 0.0005% | 0.25 | 12499.6875 | 1e-05 | 0.036085 | 0 | 0
11
12
13
INTERCEPT | A0 | 0 | B0 | -3.638e-12
SLOPE | A1 | 8.0002e-05 | B1 | 12499.6875
Note on C11: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: If Compression and Tension are required, then perform each test seperately