## Specific Vs Global
Risk Calculator | [16 empty] | Au Value Prior to Rounding
Upper Tolerance TU | 1500.2 |  |  | Conformance Probability Table |  | Global Risk - JCGM 106 Resistor Example | [5 empty] | Global Risk |  |  |  |  | 1500.18
Lower Tolerance TL | 1499.8 |  |  | Conformance Probability , Pc | Guard Band Multiplier, r Two Sided | w = r * U95 | r = TU-AU/2 um | AL=TL-r*2um | Au=TU-r*2um | Rc | Rp | TU | 1500.2
Nominal Value | 1500 | [10 empty] | TL | 1499.8
Measurement Unc um | 0.04 |  |  | 0.0668 | -0.75003 | -0.060002 | -0.75003 | 1499.74 | 1500.26 | 5.817% | 0.2418% | Au | 1500.18 | Rc | 0.9881%
Measured Value xm | 1500.16 |  |  | 0.159 | -0.49929 | -0.039943 | -0.49929 | 1499.7601 | 1500.2399 | 4.487% | 0.7128% | Al | 1499.82 | Rp | 6.907%
Tolerance T | 0.4 |  |  | 0.3085 | -0.25005 | -0.020004 | -0.25005 | 1499.78 | 1500.22 | 3.115% | 1.757% | y0 | 1500
|  |  |  | 0.5 | 0 | 0 | 0 | 1499.8 | 1500.2 | 1.895% | 3.724% | u0 | 0.12
|  |  |  | 0.6914 | 0.24991 | 0.019993 | 0.24991 | 1499.82 | 1500.18 | 0.9884% | 6.906% | um | 0.04
Probability of Conformance (pc) | 84.13% |  |  | 0.8 | 0.42081 | 0.033665 | 0.42081 | 1499.8337 | 1500.1663 | 0.574% | 9.872% | Cm | 2.5
Probability of NonConformance (1 - pc) | 15.87% |  |  | 0.85 | 0.51822 | 0.041457 | 0.51822 | 1499.8415 | 1500.1585 | 0.4054% | 11.86% | Prod U | (tl-yo)/uo | -1.6667 | < Cons U |  | Lookup Table
|  |  |  | 0.9 | 0.64078 | 0.051262 | 0.64078 | 1499.8513 | 1500.1487 | 0.2512% | 14.66% | Prod L | (tu-yo)/uo | 1.6667 | < Cons L |  | 0.24991
Setting the Guard Band Upper and Lower pc |  |  |  | 0.95 | 0.82243 | 0.065794 | 0.82243 | 1499.8658 | 1500.1342 | 0.1133% | 19.43%
Select Desired Conformance Probability | 0.6914 |  |  | 0.9545 | 0.84507 | 0.067606 | 0.84507 | 1499.8676 | 1500.1324 | 0.1018% | 20.08%
Guard Band Upper Gu | 1500.18 |  |  | 0.97725 | 1 | 0.08 | 1 | 1499.88 | 1500.12 | 0.04688% | 24.77%
Guard Band Lower GL | 1499.82 |  |  | 0.98 | 1.0269 | 0.08215 | 1.0269 | 1499.8821 | 1500.1179 | 0.04065% | 25.64%
|  |  |  | 0.99 | 1.1632
Specific Risk (Bench-Level) |  |  |  | 0.999 | 1.5451
Conditional Probability False Accept | 15.87% |  |  |  |  | Global Risk - Working Application
Conditional Probability False Reject | 84.13% |  |  |  |  | w = r * U95 | r = TU-AU/2 um | AL=TL-r*2um | Au=TU-r*2um | Rc | Rp
Global Risk | [5 empty] | -0.060002 | -0.75003 | 1499.74 | 1500.26 | 5.817% | 0.2418%
Au=TU-r*2um | 1500.18 |  |  |  |  | -0.039943 | -0.49929 | 1499.7601 | 1500.2399 | 4.487% | 0.7128%
AL=TL+r*2um | 1499.82 |  |  |  |  | -0.020004 | -0.25005 | 1499.78 | 1500.22 | 3.115% | 1.757%
r = TU-AU/2 um | 0.25 |  |  |  |  | 0 | 0 | 1499.8 | 1500.2 | 1.895% | 3.724%
u0 | 0.12 |  |  |  |  | 0.019993 | 0.24991 | 1499.82 | 1500.18 | 0.9884% | 6.906%
Cm | 2.5 |  |  |  |  | 0.033665 | 0.42081 | 1499.8337 | 1500.1663 | 0.574% | 9.872%
Cp | 0.55556 |  |  |  |  | 0.041457 | 0.51822 | 1499.8415 | 1500.1585 | 0.4054% | 11.86%
[6 empty] | 0.051262 | 0.64078 | 1499.8513 | 1500.1487 | 0.2512% | 14.66%
Global Risk | [5 empty] | 0.065794 | 0.82243 | 1499.8658 | 1500.1342 | 0.1133% | 19.43%
Consumer Risk (RC) | 0.9881% |  |  |  |  | 0.067606 | 0.84507 | 1499.8676 | 1500.1324 | 0.1018% | 20.08%
Producer Risk (RP) | 6.907% |  |  |  |  | 0.08 | 1 | 1499.88 | 1500.12 | 0.04688% | 24.77%
[6 empty] | 0.08215 | 1.0269 | 1499.8821 | 1500.1179 | 0.04065% | 25.64%
Difference Between Global and Specific Risk
Consumer (PFA) | 14.88%
Producer (PFR) | 77.23% | Conditional Probability Calculator
Note: be sure to use same units of measurement for all inputs
|  | Upper Tolerance | 1500.2
|  | Lower Tolerance | 1499.8
|  | Measured Value | 1500.16
|  | Expanded Uncertainty | 0.08
|  | Tu-y/u | 1
|  | Phi | 0.84134
|  | Tl-y/u | -9
|  | Phi | 0
|  | Probability of False Accept | 15.87%
[17 empty] | 0 |  |  |  | As-Found
[17 empty] | 0 | Nominal Value | LSL | USL | MV
[18 empty] | 1500 | 1499.8 | 1500.2 | 1500.16 | 0
[18 empty] | 1500 | 1499.8 | 1500.2 | 1500.16 | 9.9736
[19 empty] | LAL | UAL
[16 empty] | Increment | 0.00032 |  | 1499.82 | 1500.18 |  | 0
[19 empty] | 1499.82 | 1500.18 |  | 9.9736
Norm. Dist.
[16 empty] | x | y
[16 empty] | 1500 | 0.0033458
[16 empty] | 1500.0003 | 0.0034544
[16 empty] | 1500.0006 | 0.0035664
[16 empty] | 1500.001 | 0.0036818
[16 empty] | 1500.0013 | 0.0038007 | NORMAL DISTRIBUTION FUNCTION
[16 empty] | 1500.0016 | 0.0039231
[16 empty] | 1500.0019 | 0.0040493 | x = | x-axis value
[16 empty] | 1500.0022 | 0.0041792 | u = | Centerline of distribution
[16 empty] | 1500.0026 | 0.004313 | σ = | Std. Dev.
[16 empty] | 1500.0029 | 0.0044509
[16 empty] | 1500.0032 | 0.0045928
[16 empty] | 1500.0035 | 0.004739
[16 empty] | 1500.0038 | 0.0048895
[16 empty] | 1500.0042 | 0.0050444
[16 empty] | 1500.0045 | 0.005204
[16 empty] | 1500.0048 | 0.0053682
[16 empty] | 1500.0051 | 0.0055373
[16 empty] | 1500.0054 | 0.0057113
[16 empty] | 1500.0058 | 0.0058904
[16 empty] | 1500.0061 | 0.0060747
[16 empty] | 1500.0064 | 0.0062645
[16 empty] | 1500.0067 | 0.0064597
[16 empty] | 1500.007 | 0.0066606
[16 empty] | 1500.0074 | 0.0068673
[16 empty] | 1500.0077 | 0.0070799
[16 empty] | 1500.008 | 0.0072987
[16 empty] | 1500.0083 | 0.0075237
[16 empty] | 1500.0086 | 0.0077552
[16 empty] | 1500.009 | 0.0079933
[16 empty] | 1500.0093 | 0.0082382
[16 empty] | 1500.0096 | 0.00849
[16 empty] | 1500.0099 | 0.008749
[16 empty] | 1500.0102 | 0.0090153
[16 empty] | 1500.0106 | 0.0092891
[16 empty] | 1500.0109 | 0.0095706
[16 empty] | 1500.0112 | 0.0098601
[16 empty] | 1500.0115 | 0.010158
[16 empty] | 1500.0118 | 0.010463
[16 empty] | 1500.0122 | 0.010778
[16 empty] | 1500.0125 | 0.011101
[16 empty] | 1500.0128 | 0.011433
[16 empty] | 1500.0131 | 0.011774
[16 empty] | 1500.0134 | 0.012125
[16 empty] | 1500.0138 | 0.012485
[16 empty] | 1500.0141 | 0.012855
[16 empty] | 1500.0144 | 0.013235
[16 empty] | 1500.0147 | 0.013626
[16 empty] | 1500.015 | 0.014027
[16 empty] | 1500.0154 | 0.014439
[16 empty] | 1500.0157 | 0.014863
[16 empty] | 1500.016 | 0.015298
[16 empty] | 1500.0163 | 0.015744
[16 empty] | 1500.0166 | 0.016202
[16 empty] | 1500.017 | 0.016673
[16 empty] | 1500.0173 | 0.017157
[16 empty] | 1500.0176 | 0.017653
[16 empty] | 1500.0179 | 0.018162
[16 empty] | 1500.0182 | 0.018685
[16 empty] | 1500.0186 | 0.019222
[16 empty] | 1500.0189 | 0.019773
[16 empty] | 1500.0192 | 0.020338
[16 empty] | 1500.0195 | 0.020918
[16 empty] | 1500.0198 | 0.021514
[16 empty] | 1500.0202 | 0.022124
[16 empty] | 1500.0205 | 0.022751
[16 empty] | 1500.0208 | 0.023394
[16 empty] | 1500.0211 | 0.024054
[16 empty] | 1500.0214 | 0.024731
[16 empty] | 1500.0218 | 0.025425
[16 empty] | 1500.0221 | 0.026137
[16 empty] | 1500.0224 | 0.026867
[16 empty] | 1500.0227 | 0.027616
[16 empty] | 1500.023 | 0.028383
[16 empty] | 1500.0234 | 0.029171
[16 empty] | 1500.0237 | 0.029978
[16 empty] | 1500.024 | 0.030805
[16 empty] | 1500.0243 | 0.031654
[16 empty] | 1500.0246 | 0.032524
[16 empty] | 1500.025 | 0.033415
[16 empty] | 1500.0253 | 0.034329
[16 empty] | 1500.0256 | 0.035265
[16 empty] | 1500.0259 | 0.036225
[16 empty] | 1500.0262 | 0.037208
[16 empty] | 1500.0266 | 0.038216
[16 empty] | 1500.0269 | 0.039248
[16 empty] | 1500.0272 | 0.040306
[16 empty] | 1500.0275 | 0.041389
[16 empty] | 1500.0278 | 0.042499
[16 empty] | 1500.0282 | 0.043636
[16 empty] | 1500.0285 | 0.044801
[16 empty] | 1500.0288 | 0.045993
[16 empty] | 1500.0291 | 0.047214
[16 empty] | 1500.0294 | 0.048465
[16 empty] | 1500.0298 | 0.049746
[16 empty] | 1500.0301 | 0.051057
[16 empty] | 1500.0304 | 0.052399
[16 empty] | 1500.0307 | 0.053773
[16 empty] | 1500.031 | 0.05518
[16 empty] | 1500.0314 | 0.05662
[16 empty] | 1500.0317 | 0.058094
[16 empty] | 1500.032 | 0.059602
[16 empty] | 1500.0323 | 0.061146
[16 empty] | 1500.0326 | 0.062725
[16 empty] | 1500.033 | 0.064341
[16 empty] | 1500.0333 | 0.065995
[16 empty] | 1500.0336 | 0.067687
[16 empty] | 1500.0339 | 0.069418
[16 empty] | 1500.0342 | 0.071188
[16 empty] | 1500.0346 | 0.072999
[16 empty] | 1500.0349 | 0.074851
[16 empty] | 1500.0352 | 0.076745
[16 empty] | 1500.0355 | 0.078682
[16 empty] | 1500.0358 | 0.080663
[16 empty] | 1500.0362 | 0.082689
[16 empty] | 1500.0365 | 0.08476
[16 empty] | 1500.0368 | 0.086877
[16 empty] | 1500.0371 | 0.089041
[16 empty] | 1500.0374 | 0.091254
[16 empty] | 1500.0378 | 0.093515
[16 empty] | 1500.0381 | 0.095827
[16 empty] | 1500.0384 | 0.098189
[16 empty] | 1500.0387 | 0.1006
[16 empty] | 1500.039 | 0.10307
[16 empty] | 1500.0394 | 0.10559
[16 empty] | 1500.0397 | 0.10817
[16 empty] | 1500.04 | 0.1108
[16 empty] | 1500.0403 | 0.11348
[16 empty] | 1500.0406 | 0.11623
[16 empty] | 1500.041 | 0.11903
[16 empty] | 1500.0413 | 0.1219
[16 empty] | 1500.0416 | 0.12482
[16 empty] | 1500.0419 | 0.12781
[16 empty] | 1500.0422 | 0.13086
[16 empty] | 1500.0426 | 0.13397
[16 empty] | 1500.0429 | 0.13715
[16 empty] | 1500.0432 | 0.1404
[16 empty] | 1500.0435 | 0.14371
[16 empty] | 1500.0438 | 0.1471
[16 empty] | 1500.0442 | 0.15055
[16 empty] | 1500.0445 | 0.15407
[16 empty] | 1500.0448 | 0.15767
[16 empty] | 1500.0451 | 0.16134
[16 empty] | 1500.0454 | 0.16508
[16 empty] | 1500.0458 | 0.1689
[16 empty] | 1500.0461 | 0.1728
[16 empty] | 1500.0464 | 0.17678
[16 empty] | 1500.0467 | 0.18083
[16 empty] | 1500.047 | 0.18497
[16 empty] | 1500.0474 | 0.18919
[16 empty] | 1500.0477 | 0.1935
[16 empty] | 1500.048 | 0.19789
[16 empty] | 1500.0483 | 0.20236
[16 empty] | 1500.0486 | 0.20693
[16 empty] | 1500.049 | 0.21158
[16 empty] | 1500.0493 | 0.21632
[16 empty] | 1500.0496 | 0.22116
[16 empty] | 1500.0499 | 0.22609
[16 empty] | 1500.0502 | 0.23112
[16 empty] | 1500.0506 | 0.23624
[16 empty] | 1500.0509 | 0.24146
[16 empty] | 1500.0512 | 0.24678
[16 empty] | 1500.0515 | 0.2522
[16 empty] | 1500.0518 | 0.25772
[16 empty] | 1500.0522 | 0.26335
[16 empty] | 1500.0525 | 0.26908
[16 empty] | 1500.0528 | 0.27492
[16 empty] | 1500.0531 | 0.28087
[16 empty] | 1500.0534 | 0.28693
[16 empty] | 1500.0538 | 0.2931
[16 empty] | 1500.0541 | 0.29939
[16 empty] | 1500.0544 | 0.30579
[16 empty] | 1500.0547 | 0.31231
[16 empty] | 1500.055 | 0.31894
[16 empty] | 1500.0554 | 0.3257
[16 empty] | 1500.0557 | 0.33257
[16 empty] | 1500.056 | 0.33957
[16 empty] | 1500.0563 | 0.3467
[16 empty] | 1500.0566 | 0.35395
[16 empty] | 1500.057 | 0.36133
[16 empty] | 1500.0573 | 0.36885
[16 empty] | 1500.0576 | 0.37649
[16 empty] | 1500.0579 | 0.38427
[16 empty] | 1500.0582 | 0.39218
[16 empty] | 1500.0586 | 0.40023
[16 empty] | 1500.0589 | 0.40842
[16 empty] | 1500.0592 | 0.41675
[16 empty] | 1500.0595 | 0.42523
[16 empty] | 1500.0598 | 0.43384
[16 empty] | 1500.0602 | 0.44261
[16 empty] | 1500.0605 | 0.45152
[16 empty] | 1500.0608 | 0.46058
[16 empty] | 1500.0611 | 0.4698
[16 empty] | 1500.0614 | 0.47916
[16 empty] | 1500.0618 | 0.48869
[16 empty] | 1500.0621 | 0.49837
[16 empty] | 1500.0624 | 0.50821
[16 empty] | 1500.0627 | 0.51821
[16 empty] | 1500.063 | 0.52837
[16 empty] | 1500.0634 | 0.5387
[16 empty] | 1500.0637 | 0.5492
[16 empty] | 1500.064 | 0.55986
[16 empty] | 1500.0643 | 0.5707
[16 empty] | 1500.0646 | 0.58171
[16 empty] | 1500.065 | 0.59289
[16 empty] | 1500.0653 | 0.60425
[16 empty] | 1500.0656 | 0.61578
[16 empty] | 1500.0659 | 0.6275
[16 empty] | 1500.0662 | 0.6394
[16 empty] | 1500.0666 | 0.65148
[16 empty] | 1500.0669 | 0.66375
[16 empty] | 1500.0672 | 0.6762
[16 empty] | 1500.0675 | 0.68885
[16 empty] | 1500.0678 | 0.70168
[16 empty] | 1500.0682 | 0.71472
[16 empty] | 1500.0685 | 0.72794
[16 empty] | 1500.0688 | 0.74136
[16 empty] | 1500.0691 | 0.75499
[16 empty] | 1500.0694 | 0.76881
[16 empty] | 1500.0698 | 0.78284
[16 empty] | 1500.0701 | 0.79707
[16 empty] | 1500.0704 | 0.81151
[16 empty] | 1500.0707 | 0.82615
[16 empty] | 1500.071 | 0.84101
[16 empty] | 1500.0714 | 0.85608
[16 empty] | 1500.0717 | 0.87136
[16 empty] | 1500.072 | 0.88686
[16 empty] | 1500.0723 | 0.90258
[16 empty] | 1500.0726 | 0.91852
[16 empty] | 1500.073 | 0.93468
[16 empty] | 1500.0733 | 0.95106
[16 empty] | 1500.0736 | 0.96767
[16 empty] | 1500.0739 | 0.98451
[16 empty] | 1500.0742 | 1.0016
[16 empty] | 1500.0746 | 1.0189
[16 empty] | 1500.0749 | 1.0364
[16 empty] | 1500.0752 | 1.0542
[16 empty] | 1500.0755 | 1.0721
[16 empty] | 1500.0758 | 1.0904
[16 empty] | 1500.0762 | 1.1089
[16 empty] | 1500.0765 | 1.1276
[16 empty] | 1500.0768 | 1.1465
[16 empty] | 1500.0771 | 1.1657
[16 empty] | 1500.0774 | 1.1852
[16 empty] | 1500.0778 | 1.2049
[16 empty] | 1500.0781 | 1.2248
[16 empty] | 1500.0784 | 1.245
[16 empty] | 1500.0787 | 1.2654
[16 empty] | 1500.079 | 1.2861
[16 empty] | 1500.0794 | 1.3071
[16 empty] | 1500.0797 | 1.3283
[16 empty] | 1500.08 | 1.3498
[16 empty] | 1500.0803 | 1.3715
[16 empty] | 1500.0806 | 1.3935
[16 empty] | 1500.081 | 1.4157
[16 empty] | 1500.0813 | 1.4382
[16 empty] | 1500.0816 | 1.461
[16 empty] | 1500.0819 | 1.4841
[16 empty] | 1500.0822 | 1.5074
[16 empty] | 1500.0826 | 1.531
[16 empty] | 1500.0829 | 1.5548
[16 empty] | 1500.0832 | 1.5789
[16 empty] | 1500.0835 | 1.6033
[16 empty] | 1500.0838 | 1.628
[16 empty] | 1500.0842 | 1.6529
[16 empty] | 1500.0845 | 1.6781
[16 empty] | 1500.0848 | 1.7036
[16 empty] | 1500.0851 | 1.7293
[16 empty] | 1500.0854 | 1.7554
[16 empty] | 1500.0858 | 1.7817
[16 empty] | 1500.0861 | 1.8083
[16 empty] | 1500.0864 | 1.8352
[16 empty] | 1500.0867 | 1.8623
[16 empty] | 1500.087 | 1.8898
[16 empty] | 1500.0874 | 1.9175
[16 empty] | 1500.0877 | 1.9455
[16 empty] | 1500.088 | 1.9738
[16 empty] | 1500.0883 | 2.0023
[16 empty] | 1500.0886 | 2.0312
[16 empty] | 1500.089 | 2.0603
[16 empty] | 1500.0893 | 2.0897
[16 empty] | 1500.0896 | 2.1194
[16 empty] | 1500.0899 | 2.1494
[16 empty] | 1500.0902 | 2.1797
[16 empty] | 1500.0906 | 2.2102
[16 empty] | 1500.0909 | 2.2411
[16 empty] | 1500.0912 | 2.2722
[16 empty] | 1500.0915 | 2.3036
[16 empty] | 1500.0918 | 2.3353
[16 empty] | 1500.0922 | 2.3673
[16 empty] | 1500.0925 | 2.3995
[16 empty] | 1500.0928 | 2.4321
[16 empty] | 1500.0931 | 2.4649
[16 empty] | 1500.0934 | 2.498
[16 empty] | 1500.0938 | 2.5314
[16 empty] | 1500.0941 | 2.5651
[16 empty] | 1500.0944 | 2.599
[16 empty] | 1500.0947 | 2.6333
[16 empty] | 1500.095 | 2.6678
[16 empty] | 1500.0954 | 2.7026
[16 empty] | 1500.0957 | 2.7377
[16 empty] | 1500.096 | 2.773
[16 empty] | 1500.0963 | 2.8087
[16 empty] | 1500.0966 | 2.8446
[16 empty] | 1500.097 | 2.8807
[16 empty] | 1500.0973 | 2.9172
[16 empty] | 1500.0976 | 2.9539
[16 empty] | 1500.0979 | 2.9909
[16 empty] | 1500.0982 | 3.0282
[16 empty] | 1500.0986 | 3.0657
[16 empty] | 1500.0989 | 3.1035
[16 empty] | 1500.0992 | 3.1416
[16 empty] | 1500.0995 | 3.1799
[16 empty] | 1500.0998 | 3.2185
[16 empty] | 1500.1002 | 3.2574
[16 empty] | 1500.1005 | 3.2965
[16 empty] | 1500.1008 | 3.3359
[16 empty] | 1500.1011 | 3.3755
[16 empty] | 1500.1014 | 3.4154
[16 empty] | 1500.1018 | 3.4555
[16 empty] | 1500.1021 | 3.4959
[16 empty] | 1500.1024 | 3.5365
[16 empty] | 1500.1027 | 3.5774
[16 empty] | 1500.103 | 3.6185
[16 empty] | 1500.1034 | 3.6598
[16 empty] | 1500.1037 | 3.7014
[16 empty] | 1500.104 | 3.7432
[16 empty] | 1500.1043 | 3.7852
[16 empty] | 1500.1046 | 3.8275
[16 empty] | 1500.105 | 3.87
[16 empty] | 1500.1053 | 3.9127
[16 empty] | 1500.1056 | 3.9556
[16 empty] | 1500.1059 | 3.9988
[16 empty] | 1500.1062 | 4.0421
[16 empty] | 1500.1066 | 4.0857
[16 empty] | 1500.1069 | 4.1295
[16 empty] | 1500.1072 | 4.1734
[16 empty] | 1500.1075 | 4.2176
[16 empty] | 1500.1078 | 4.262
[16 empty] | 1500.1082 | 4.3065
[16 empty] | 1500.1085 | 4.3513
[16 empty] | 1500.1088 | 4.3962
[16 empty] | 1500.1091 | 4.4413
[16 empty] | 1500.1094 | 4.4866
[16 empty] | 1500.1098 | 4.532
[16 empty] | 1500.1101 | 4.5776
[16 empty] | 1500.1104 | 4.6234
[16 empty] | 1500.1107 | 4.6694
[16 empty] | 1500.111 | 4.7155
[16 empty] | 1500.1114 | 4.7617
[16 empty] | 1500.1117 | 4.8081
[16 empty] | 1500.112 | 4.8547
[16 empty] | 1500.1123 | 4.9013
[16 empty] | 1500.1126 | 4.9481
[16 empty] | 1500.113 | 4.9951
[16 empty] | 1500.1133 | 5.0421
[16 empty] | 1500.1136 | 5.0893
[16 empty] | 1500.1139 | 5.1366
[16 empty] | 1500.1142 | 5.184
[16 empty] | 1500.1146 | 5.2315
[16 empty] | 1500.1149 | 5.279
[16 empty] | 1500.1152 | 5.3267
[16 empty] | 1500.1155 | 5.3745
[16 empty] | 1500.1158 | 5.4223
[16 empty] | 1500.1162 | 5.4703
[16 empty] | 1500.1165 | 5.5183
[16 empty] | 1500.1168 | 5.5663
[16 empty] | 1500.1171 | 5.6145
[16 empty] | 1500.1174 | 5.6626
[16 empty] | 1500.1178 | 5.7109
[16 empty] | 1500.1181 | 5.7591
[16 empty] | 1500.1184 | 5.8074
[16 empty] | 1500.1187 | 5.8558
[16 empty] | 1500.119 | 5.9041
[16 empty] | 1500.1194 | 5.9525
[16 empty] | 1500.1197 | 6.0009
[16 empty] | 1500.12 | 6.0493
[16 empty] | 1500.1203 | 6.0977
[16 empty] | 1500.1206 | 6.146
[16 empty] | 1500.121 | 6.1944
[16 empty] | 1500.1213 | 6.2428
[16 empty] | 1500.1216 | 6.2911
[16 empty] | 1500.1219 | 6.3394
[16 empty] | 1500.1222 | 6.3877
[16 empty] | 1500.1226 | 6.4359
[16 empty] | 1500.1229 | 6.484
[16 empty] | 1500.1232 | 6.5322
[16 empty] | 1500.1235 | 6.5802
[16 empty] | 1500.1238 | 6.6282
[16 empty] | 1500.1242 | 6.6761
[16 empty] | 1500.1245 | 6.7239
[16 empty] | 1500.1248 | 6.7716
[16 empty] | 1500.1251 | 6.8192
[16 empty] | 1500.1254 | 6.8667
[16 empty] | 1500.1258 | 6.9141
[16 empty] | 1500.1261 | 6.9614
[16 empty] | 1500.1264 | 7.0086
[16 empty] | 1500.1267 | 7.0556
[16 empty] | 1500.127 | 7.1025
[16 empty] | 1500.1274 | 7.1493
[16 empty] | 1500.1277 | 7.1959
[16 empty] | 1500.128 | 7.2423
[16 empty] | 1500.1283 | 7.2886
[16 empty] | 1500.1286 | 7.3346
[16 empty] | 1500.129 | 7.3806
[16 empty] | 1500.1293 | 7.4263
[16 empty] | 1500.1296 | 7.4718
[16 empty] | 1500.1299 | 7.5171
[16 empty] | 1500.1302 | 7.5623
[16 empty] | 1500.1306 | 7.6072
[16 empty] | 1500.1309 | 7.6518
[16 empty] | 1500.1312 | 7.6963
[16 empty] | 1500.1315 | 7.7405
[16 empty] | 1500.1318 | 7.7845
[16 empty] | 1500.1322 | 7.8282
[16 empty] | 1500.1325 | 7.8716
[16 empty] | 1500.1328 | 7.9148
[16 empty] | 1500.1331 | 7.9577
[16 empty] | 1500.1334 | 8.0004
[16 empty] | 1500.1338 | 8.0427
[16 empty] | 1500.1341 | 8.0848
[16 empty] | 1500.1344 | 8.1266
[16 empty] | 1500.1347 | 8.168
[16 empty] | 1500.135 | 8.2091
[16 empty] | 1500.1354 | 8.25
[16 empty] | 1500.1357 | 8.2905
[16 empty] | 1500.136 | 8.3306
[16 empty] | 1500.1363 | 8.3704
[16 empty] | 1500.1366 | 8.4099
[16 empty] | 1500.137 | 8.449
[16 empty] | 1500.1373 | 8.4878
[16 empty] | 1500.1376 | 8.5261
[16 empty] | 1500.1379 | 8.5642
[16 empty] | 1500.1382 | 8.6018
[16 empty] | 1500.1386 | 8.639
[16 empty] | 1500.1389 | 8.6759
[16 empty] | 1500.1392 | 8.7123
[16 empty] | 1500.1395 | 8.7484
[16 empty] | 1500.1398 | 8.784
[16 empty] | 1500.1402 | 8.8192
[16 empty] | 1500.1405 | 8.854
[16 empty] | 1500.1408 | 8.8883
[16 empty] | 1500.1411 | 8.9222
[16 empty] | 1500.1414 | 8.9557
[16 empty] | 1500.1418 | 8.9887
[16 empty] | 1500.1421 | 9.0213
[16 empty] | 1500.1424 | 9.0534
[16 empty] | 1500.1427 | 9.085
[16 empty] | 1500.143 | 9.1162
[16 empty] | 1500.1434 | 9.1468
[16 empty] | 1500.1437 | 9.177
[16 empty] | 1500.144 | 9.2068
[16 empty] | 1500.1443 | 9.236
[16 empty] | 1500.1446 | 9.2647
[16 empty] | 1500.145 | 9.2929
[16 empty] | 1500.1453 | 9.3206
[16 empty] | 1500.1456 | 9.3478
[16 empty] | 1500.1459 | 9.3744
[16 empty] | 1500.1462 | 9.4006
[16 empty] | 1500.1466 | 9.4262
[16 empty] | 1500.1469 | 9.4512
[16 empty] | 1500.1472 | 9.4758
[16 empty] | 1500.1475 | 9.4997
[16 empty] | 1500.1478 | 9.5232
[16 empty] | 1500.1482 | 9.5461
[16 empty] | 1500.1485 | 9.5684
[16 empty] | 1500.1488 | 9.5902
[16 empty] | 1500.1491 | 9.6114
[16 empty] | 1500.1494 | 9.632
[16 empty] | 1500.1498 | 9.652
[16 empty] | 1500.1501 | 9.6715
[16 empty] | 1500.1504 | 9.6904
[16 empty] | 1500.1507 | 9.7087
[16 empty] | 1500.151 | 9.7265
[16 empty] | 1500.1514 | 9.7436
[16 empty] | 1500.1517 | 9.7601
[16 empty] | 1500.152 | 9.7761
[16 empty] | 1500.1523 | 9.7914
[16 empty] | 1500.1526 | 9.8061
[16 empty] | 1500.153 | 9.8203
[16 empty] | 1500.1533 | 9.8338
[16 empty] | 1500.1536 | 9.8467
[16 empty] | 1500.1539 | 9.859
[16 empty] | 1500.1542 | 9.8707
[16 empty] | 1500.1546 | 9.8817
[16 empty] | 1500.1549 | 9.8922
[16 empty] | 1500.1552 | 9.902
[16 empty] | 1500.1555 | 9.9112
[16 empty] | 1500.1558 | 9.9198
[16 empty] | 1500.1562 | 9.9277
[16 empty] | 1500.1565 | 9.935
[16 empty] | 1500.1568 | 9.9417
[16 empty] | 1500.1571 | 9.9477
[16 empty] | 1500.1574 | 9.9532
[16 empty] | 1500.1578 | 9.9579
[16 empty] | 1500.1581 | 9.9621
[16 empty] | 1500.1584 | 9.9656
[16 empty] | 1500.1587 | 9.9685
[16 empty] | 1500.159 | 9.9707
[16 empty] | 1500.1594 | 9.9723
[16 empty] | 1500.1597 | 9.9732
[16 empty] | 1500.16 | 9.9736
[16 empty] | 1500.1603 | 9.9732
[16 empty] | 1500.1606 | 9.9723
[16 empty] | 1500.161 | 9.9707
[16 empty] | 1500.1613 | 9.9685
[16 empty] | 1500.1616 | 9.9656
[16 empty] | 1500.1619 | 9.9621
[16 empty] | 1500.1622 | 9.9579
[16 empty] | 1500.1626 | 9.9532
[16 empty] | 1500.1629 | 9.9477
[16 empty] | 1500.1632 | 9.9417
[16 empty] | 1500.1635 | 9.935
[16 empty] | 1500.1638 | 9.9277
[16 empty] | 1500.1642 | 9.9198
[16 empty] | 1500.1645 | 9.9112
[16 empty] | 1500.1648 | 9.902
[16 empty] | 1500.1651 | 9.8922
[16 empty] | 1500.1654 | 9.8817
[16 empty] | 1500.1658 | 9.8707
[16 empty] | 1500.1661 | 9.859
[16 empty] | 1500.1664 | 9.8467
[16 empty] | 1500.1667 | 9.8338
[16 empty] | 1500.167 | 9.8203
[16 empty] | 1500.1674 | 9.8061
[16 empty] | 1500.1677 | 9.7914
[16 empty] | 1500.168 | 9.7761
[16 empty] | 1500.1683 | 9.7601
[16 empty] | 1500.1686 | 9.7436
[16 empty] | 1500.169 | 9.7265
[16 empty] | 1500.1693 | 9.7087
[16 empty] | 1500.1696 | 9.6904
[16 empty] | 1500.1699 | 9.6715
[16 empty] | 1500.1702 | 9.652
[16 empty] | 1500.1706 | 9.632
[16 empty] | 1500.1709 | 9.6114
[16 empty] | 1500.1712 | 9.5902
[16 empty] | 1500.1715 | 9.5684
[16 empty] | 1500.1718 | 9.5461
[16 empty] | 1500.1722 | 9.5232
[16 empty] | 1500.1725 | 9.4997
[16 empty] | 1500.1728 | 9.4758
[16 empty] | 1500.1731 | 9.4512
[16 empty] | 1500.1734 | 9.4262
[16 empty] | 1500.1738 | 9.4006
[16 empty] | 1500.1741 | 9.3744
[16 empty] | 1500.1744 | 9.3478
[16 empty] | 1500.1747 | 9.3206
[16 empty] | 1500.175 | 9.2929
[16 empty] | 1500.1754 | 9.2647
[16 empty] | 1500.1757 | 9.236
[16 empty] | 1500.176 | 9.2068
[16 empty] | 1500.1763 | 9.177
[16 empty] | 1500.1766 | 9.1468
[16 empty] | 1500.177 | 9.1162
[16 empty] | 1500.1773 | 9.085
[16 empty] | 1500.1776 | 9.0534
[16 empty] | 1500.1779 | 9.0213
[16 empty] | 1500.1782 | 8.9887
[16 empty] | 1500.1786 | 8.9557
[16 empty] | 1500.1789 | 8.9222
[16 empty] | 1500.1792 | 8.8883
[16 empty] | 1500.1795 | 8.854
[16 empty] | 1500.1798 | 8.8192
[16 empty] | 1500.1802 | 8.784
[16 empty] | 1500.1805 | 8.7484
[16 empty] | 1500.1808 | 8.7123
[16 empty] | 1500.1811 | 8.6759
[16 empty] | 1500.1814 | 8.639
[16 empty] | 1500.1818 | 8.6018
[16 empty] | 1500.1821 | 8.5642
[16 empty] | 1500.1824 | 8.5261
[16 empty] | 1500.1827 | 8.4878
[16 empty] | 1500.183 | 8.449
[16 empty] | 1500.1834 | 8.4099
[16 empty] | 1500.1837 | 8.3704
[16 empty] | 1500.184 | 8.3306
[16 empty] | 1500.1843 | 8.2905
[16 empty] | 1500.1846 | 8.25
[16 empty] | 1500.185 | 8.2091
[16 empty] | 1500.1853 | 8.168
[16 empty] | 1500.1856 | 8.1266
[16 empty] | 1500.1859 | 8.0848
[16 empty] | 1500.1862 | 8.0427
[16 empty] | 1500.1866 | 8.0004
[16 empty] | 1500.1869 | 7.9577
[16 empty] | 1500.1872 | 7.9148
[16 empty] | 1500.1875 | 7.8716
[16 empty] | 1500.1878 | 7.8282
[16 empty] | 1500.1882 | 7.7845
[16 empty] | 1500.1885 | 7.7405
[16 empty] | 1500.1888 | 7.6963
[16 empty] | 1500.1891 | 7.6518
[16 empty] | 1500.1894 | 7.6072
[16 empty] | 1500.1898 | 7.5623
[16 empty] | 1500.1901 | 7.5171
[16 empty] | 1500.1904 | 7.4718
[16 empty] | 1500.1907 | 7.4263
[16 empty] | 1500.191 | 7.3806
[16 empty] | 1500.1914 | 7.3346
[16 empty] | 1500.1917 | 7.2886
[16 empty] | 1500.192 | 7.2423
[16 empty] | 1500.1923 | 7.1959
[16 empty] | 1500.1926 | 7.1493
[16 empty] | 1500.193 | 7.1025
[16 empty] | 1500.1933 | 7.0556
[16 empty] | 1500.1936 | 7.0086
[16 empty] | 1500.1939 | 6.9614
[16 empty] | 1500.1942 | 6.9141
[16 empty] | 1500.1946 | 6.8667
[16 empty] | 1500.1949 | 6.8192
[16 empty] | 1500.1952 | 6.7716
[16 empty] | 1500.1955 | 6.7239
[16 empty] | 1500.1958 | 6.6761
[16 empty] | 1500.1962 | 6.6282
[16 empty] | 1500.1965 | 6.5802
[16 empty] | 1500.1968 | 6.5322
[16 empty] | 1500.1971 | 6.484
[16 empty] | 1500.1974 | 6.4359
[16 empty] | 1500.1978 | 6.3877
[16 empty] | 1500.1981 | 6.3394
[16 empty] | 1500.1984 | 6.2911
[16 empty] | 1500.1987 | 6.2428
[16 empty] | 1500.199 | 6.1944
[16 empty] | 1500.1994 | 6.146
[16 empty] | 1500.1997 | 6.0977
[16 empty] | 1500.2 | 6.0493
[16 empty] | 1500.2003 | 6.0009
[16 empty] | 1500.2006 | 5.9525
[16 empty] | 1500.201 | 5.9041
[16 empty] | 1500.2013 | 5.8558
[16 empty] | 1500.2016 | 5.8074
[16 empty] | 1500.2019 | 5.7591
[16 empty] | 1500.2022 | 5.7109
[16 empty] | 1500.2026 | 5.6626
[16 empty] | 1500.2029 | 5.6145
[16 empty] | 1500.2032 | 5.5663
[16 empty] | 1500.2035 | 5.5183
[16 empty] | 1500.2038 | 5.4703
[16 empty] | 1500.2042 | 5.4223
[16 empty] | 1500.2045 | 5.3745
[16 empty] | 1500.2048 | 5.3267
[16 empty] | 1500.2051 | 5.279
[16 empty] | 1500.2054 | 5.2315
[16 empty] | 1500.2058 | 5.184
[16 empty] | 1500.2061 | 5.1366
[16 empty] | 1500.2064 | 5.0893
[16 empty] | 1500.2067 | 5.0421
[16 empty] | 1500.207 | 4.9951
[16 empty] | 1500.2074 | 4.9481
[16 empty] | 1500.2077 | 4.9013
[16 empty] | 1500.208 | 4.8547
[16 empty] | 1500.2083 | 4.8081
[16 empty] | 1500.2086 | 4.7617
[16 empty] | 1500.209 | 4.7155
[16 empty] | 1500.2093 | 4.6694
[16 empty] | 1500.2096 | 4.6234
[16 empty] | 1500.2099 | 4.5776
[16 empty] | 1500.2102 | 4.532
[16 empty] | 1500.2106 | 4.4866
[16 empty] | 1500.2109 | 4.4413
[16 empty] | 1500.2112 | 4.3962
[16 empty] | 1500.2115 | 4.3513
[16 empty] | 1500.2118 | 4.3065
[16 empty] | 1500.2122 | 4.262
[16 empty] | 1500.2125 | 4.2176
[16 empty] | 1500.2128 | 4.1734
[16 empty] | 1500.2131 | 4.1295
[16 empty] | 1500.2134 | 4.0857
[16 empty] | 1500.2138 | 4.0421
[16 empty] | 1500.2141 | 3.9988
[16 empty] | 1500.2144 | 3.9556
[16 empty] | 1500.2147 | 3.9127
[16 empty] | 1500.215 | 3.87
[16 empty] | 1500.2154 | 3.8275
[16 empty] | 1500.2157 | 3.7852
[16 empty] | 1500.216 | 3.7432
[16 empty] | 1500.2163 | 3.7014
[16 empty] | 1500.2166 | 3.6598
[16 empty] | 1500.217 | 3.6185
[16 empty] | 1500.2173 | 3.5774
[16 empty] | 1500.2176 | 3.5365
[16 empty] | 1500.2179 | 3.4959
[16 empty] | 1500.2182 | 3.4555
[16 empty] | 1500.2186 | 3.4154
[16 empty] | 1500.2189 | 3.3755
[16 empty] | 1500.2192 | 3.3359
[16 empty] | 1500.2195 | 3.2965
[16 empty] | 1500.2198 | 3.2574
[16 empty] | 1500.2202 | 3.2185
[16 empty] | 1500.2205 | 3.1799
[16 empty] | 1500.2208 | 3.1416
[16 empty] | 1500.2211 | 3.1035
[16 empty] | 1500.2214 | 3.0657
[16 empty] | 1500.2218 | 3.0282
[16 empty] | 1500.2221 | 2.9909
[16 empty] | 1500.2224 | 2.9539
[16 empty] | 1500.2227 | 2.9172
[16 empty] | 1500.223 | 2.8807
[16 empty] | 1500.2234 | 2.8446
[16 empty] | 1500.2237 | 2.8087
[16 empty] | 1500.224 | 2.773
[16 empty] | 1500.2243 | 2.7377
[16 empty] | 1500.2246 | 2.7026
[16 empty] | 1500.225 | 2.6678
[16 empty] | 1500.2253 | 2.6333
[16 empty] | 1500.2256 | 2.599
[16 empty] | 1500.2259 | 2.5651
[16 empty] | 1500.2262 | 2.5314
[16 empty] | 1500.2266 | 2.498
[16 empty] | 1500.2269 | 2.4649
[16 empty] | 1500.2272 | 2.4321
[16 empty] | 1500.2275 | 2.3995
[16 empty] | 1500.2278 | 2.3673
[16 empty] | 1500.2282 | 2.3353
[16 empty] | 1500.2285 | 2.3036
[16 empty] | 1500.2288 | 2.2722
[16 empty] | 1500.2291 | 2.2411
[16 empty] | 1500.2294 | 2.2102
[16 empty] | 1500.2298 | 2.1797
[16 empty] | 1500.2301 | 2.1494
[16 empty] | 1500.2304 | 2.1194
[16 empty] | 1500.2307 | 2.0897
[16 empty] | 1500.231 | 2.0603
[16 empty] | 1500.2314 | 2.0312
[16 empty] | 1500.2317 | 2.0023
[16 empty] | 1500.232 | 1.9738
[16 empty] | 1500.2323 | 1.9455
[16 empty] | 1500.2326 | 1.9175
[16 empty] | 1500.233 | 1.8898
[16 empty] | 1500.2333 | 1.8623
[16 empty] | 1500.2336 | 1.8352
[16 empty] | 1500.2339 | 1.8083
[16 empty] | 1500.2342 | 1.7817
[16 empty] | 1500.2346 | 1.7554
[16 empty] | 1500.2349 | 1.7293
[16 empty] | 1500.2352 | 1.7036
[16 empty] | 1500.2355 | 1.6781
[16 empty] | 1500.2358 | 1.6529
[16 empty] | 1500.2362 | 1.628
[16 empty] | 1500.2365 | 1.6033
[16 empty] | 1500.2368 | 1.5789
[16 empty] | 1500.2371 | 1.5548
[16 empty] | 1500.2374 | 1.531
[16 empty] | 1500.2378 | 1.5074
[16 empty] | 1500.2381 | 1.4841
[16 empty] | 1500.2384 | 1.461
[16 empty] | 1500.2387 | 1.4382
[16 empty] | 1500.239 | 1.4157
[16 empty] | 1500.2394 | 1.3935
[16 empty] | 1500.2397 | 1.3715
[16 empty] | 1500.24 | 1.3498
[16 empty] | 1500.2403 | 1.3283
[16 empty] | 1500.2406 | 1.3071
[16 empty] | 1500.241 | 1.2861
[16 empty] | 1500.2413 | 1.2654
[16 empty] | 1500.2416 | 1.245
[16 empty] | 1500.2419 | 1.2248
[16 empty] | 1500.2422 | 1.2049
[16 empty] | 1500.2426 | 1.1852
[16 empty] | 1500.2429 | 1.1657
[16 empty] | 1500.2432 | 1.1465
[16 empty] | 1500.2435 | 1.1276
[16 empty] | 1500.2438 | 1.1089
[16 empty] | 1500.2442 | 1.0904
[16 empty] | 1500.2445 | 1.0721
[16 empty] | 1500.2448 | 1.0542
[16 empty] | 1500.2451 | 1.0364
[16 empty] | 1500.2454 | 1.0189
[16 empty] | 1500.2458 | 1.0016
[16 empty] | 1500.2461 | 0.98451
[16 empty] | 1500.2464 | 0.96767
[16 empty] | 1500.2467 | 0.95106
[16 empty] | 1500.247 | 0.93468
[16 empty] | 1500.2474 | 0.91852
[16 empty] | 1500.2477 | 0.90258
[16 empty] | 1500.248 | 0.88686
[16 empty] | 1500.2483 | 0.87136
[16 empty] | 1500.2486 | 0.85608
[16 empty] | 1500.249 | 0.84101
[16 empty] | 1500.2493 | 0.82615
[16 empty] | 1500.2496 | 0.81151
[16 empty] | 1500.2499 | 0.79707
[16 empty] | 1500.2502 | 0.78284
[16 empty] | 1500.2506 | 0.76881
[16 empty] | 1500.2509 | 0.75499
[16 empty] | 1500.2512 | 0.74136
[16 empty] | 1500.2515 | 0.72794
[16 empty] | 1500.2518 | 0.71472
[16 empty] | 1500.2522 | 0.70168
[16 empty] | 1500.2525 | 0.68885
[16 empty] | 1500.2528 | 0.6762
[16 empty] | 1500.2531 | 0.66375
[16 empty] | 1500.2534 | 0.65148
[16 empty] | 1500.2538 | 0.6394
[16 empty] | 1500.2541 | 0.6275
[16 empty] | 1500.2544 | 0.61578
[16 empty] | 1500.2547 | 0.60425
[16 empty] | 1500.255 | 0.59289
[16 empty] | 1500.2554 | 0.58171
[16 empty] | 1500.2557 | 0.5707
[16 empty] | 1500.256 | 0.55986
[16 empty] | 1500.2563 | 0.5492
[16 empty] | 1500.2566 | 0.5387
[16 empty] | 1500.257 | 0.52837
[16 empty] | 1500.2573 | 0.51821
[16 empty] | 1500.2576 | 0.50821
[16 empty] | 1500.2579 | 0.49837
[16 empty] | 1500.2582 | 0.48869
[16 empty] | 1500.2586 | 0.47916
[16 empty] | 1500.2589 | 0.4698
[16 empty] | 1500.2592 | 0.46058
[16 empty] | 1500.2595 | 0.45152
[16 empty] | 1500.2598 | 0.44261
[16 empty] | 1500.2602 | 0.43384
[16 empty] | 1500.2605 | 0.42523
[16 empty] | 1500.2608 | 0.41675
[16 empty] | 1500.2611 | 0.40842
[16 empty] | 1500.2614 | 0.40023
[16 empty] | 1500.2618 | 0.39218
[16 empty] | 1500.2621 | 0.38427
[16 empty] | 1500.2624 | 0.37649
[16 empty] | 1500.2627 | 0.36885
[16 empty] | 1500.263 | 0.36133
[16 empty] | 1500.2634 | 0.35395
[16 empty] | 1500.2637 | 0.3467
[16 empty] | 1500.264 | 0.33957
[16 empty] | 1500.2643 | 0.33257
[16 empty] | 1500.2646 | 0.3257
[16 empty] | 1500.265 | 0.31894
[16 empty] | 1500.2653 | 0.31231
[16 empty] | 1500.2656 | 0.30579
[16 empty] | 1500.2659 | 0.29939
[16 empty] | 1500.2662 | 0.2931
[16 empty] | 1500.2666 | 0.28693
[16 empty] | 1500.2669 | 0.28087
[16 empty] | 1500.2672 | 0.27492
[16 empty] | 1500.2675 | 0.26908
[16 empty] | 1500.2678 | 0.26335
[16 empty] | 1500.2682 | 0.25772
[16 empty] | 1500.2685 | 0.2522
[16 empty] | 1500.2688 | 0.24678
[16 empty] | 1500.2691 | 0.24146
[16 empty] | 1500.2694 | 0.23624
[16 empty] | 1500.2698 | 0.23112
[16 empty] | 1500.2701 | 0.22609
[16 empty] | 1500.2704 | 0.22116
[16 empty] | 1500.2707 | 0.21632
[16 empty] | 1500.271 | 0.21158
[16 empty] | 1500.2714 | 0.20693
[16 empty] | 1500.2717 | 0.20236
[16 empty] | 1500.272 | 0.19789
[16 empty] | 1500.2723 | 0.1935
[16 empty] | 1500.2726 | 0.18919
[16 empty] | 1500.273 | 0.18497
[16 empty] | 1500.2733 | 0.18083
[16 empty] | 1500.2736 | 0.17678
[16 empty] | 1500.2739 | 0.1728
[16 empty] | 1500.2742 | 0.1689
[16 empty] | 1500.2746 | 0.16508
[16 empty] | 1500.2749 | 0.16134
[16 empty] | 1500.2752 | 0.15767
[16 empty] | 1500.2755 | 0.15407
[16 empty] | 1500.2758 | 0.15055
[16 empty] | 1500.2762 | 0.1471
[16 empty] | 1500.2765 | 0.14371
[16 empty] | 1500.2768 | 0.1404
[16 empty] | 1500.2771 | 0.13715
[16 empty] | 1500.2774 | 0.13397
[16 empty] | 1500.2778 | 0.13086
[16 empty] | 1500.2781 | 0.12781
[16 empty] | 1500.2784 | 0.12482
[16 empty] | 1500.2787 | 0.1219
[16 empty] | 1500.279 | 0.11903
[16 empty] | 1500.2794 | 0.11623
[16 empty] | 1500.2797 | 0.11348
[16 empty] | 1500.28 | 0.1108
[16 empty] | 1500.2803 | 0.10817
[16 empty] | 1500.2806 | 0.10559
[16 empty] | 1500.281 | 0.10307
[16 empty] | 1500.2813 | 0.1006
[16 empty] | 1500.2816 | 0.098189
[16 empty] | 1500.2819 | 0.095827
[16 empty] | 1500.2822 | 0.093515
[16 empty] | 1500.2826 | 0.091254
[16 empty] | 1500.2829 | 0.089041
[16 empty] | 1500.2832 | 0.086877
[16 empty] | 1500.2835 | 0.08476
[16 empty] | 1500.2838 | 0.082689
[16 empty] | 1500.2842 | 0.080663
[16 empty] | 1500.2845 | 0.078682
[16 empty] | 1500.2848 | 0.076745
[16 empty] | 1500.2851 | 0.074851
[16 empty] | 1500.2854 | 0.072999
[16 empty] | 1500.2858 | 0.071188
[16 empty] | 1500.2861 | 0.069418
[16 empty] | 1500.2864 | 0.067687
[16 empty] | 1500.2867 | 0.065995
[16 empty] | 1500.287 | 0.064341
[16 empty] | 1500.2874 | 0.062725
[16 empty] | 1500.2877 | 0.061146
[16 empty] | 1500.288 | 0.059602
[16 empty] | 1500.2883 | 0.058094
[16 empty] | 1500.2886 | 0.05662
[16 empty] | 1500.289 | 0.05518
[16 empty] | 1500.2893 | 0.053773
[16 empty] | 1500.2896 | 0.052399
[16 empty] | 1500.2899 | 0.051057
[16 empty] | 1500.2902 | 0.049746
[16 empty] | 1500.2906 | 0.048465
[16 empty] | 1500.2909 | 0.047214
[16 empty] | 1500.2912 | 0.045993
[16 empty] | 1500.2915 | 0.044801
[16 empty] | 1500.2918 | 0.043636
[16 empty] | 1500.2922 | 0.042499
[16 empty] | 1500.2925 | 0.041389
[16 empty] | 1500.2928 | 0.040306
[16 empty] | 1500.2931 | 0.039248
[16 empty] | 1500.2934 | 0.038216
[16 empty] | 1500.2938 | 0.037208
[16 empty] | 1500.2941 | 0.036225
[16 empty] | 1500.2944 | 0.035265
[16 empty] | 1500.2947 | 0.034329
[16 empty] | 1500.295 | 0.033415
[16 empty] | 1500.2954 | 0.032524
[16 empty] | 1500.2957 | 0.031654
[16 empty] | 1500.296 | 0.030805
[16 empty] | 1500.2963 | 0.029978
[16 empty] | 1500.2966 | 0.029171
[16 empty] | 1500.297 | 0.028383
[16 empty] | 1500.2973 | 0.027616
[16 empty] | 1500.2976 | 0.026867
[16 empty] | 1500.2979 | 0.026137
[16 empty] | 1500.2982 | 0.025425
[16 empty] | 1500.2986 | 0.024731
[16 empty] | 1500.2989 | 0.024054
[16 empty] | 1500.2992 | 0.023394
[16 empty] | 1500.2995 | 0.022751
[16 empty] | 1500.2998 | 0.022124
[16 empty] | 1500.3002 | 0.021514
[16 empty] | 1500.3005 | 0.020918
[16 empty] | 1500.3008 | 0.020338
[16 empty] | 1500.3011 | 0.019773
[16 empty] | 1500.3014 | 0.019222
[16 empty] | 1500.3018 | 0.018685
[16 empty] | 1500.3021 | 0.018162
[16 empty] | 1500.3024 | 0.017653
[16 empty] | 1500.3027 | 0.017157
[16 empty] | 1500.303 | 0.016673
[16 empty] | 1500.3034 | 0.016202
[16 empty] | 1500.3037 | 0.015744
[16 empty] | 1500.304 | 0.015298
[16 empty] | 1500.3043 | 0.014863
[16 empty] | 1500.3046 | 0.014439
[16 empty] | 1500.305 | 0.014027
[16 empty] | 1500.3053 | 0.013626
[16 empty] | 1500.3056 | 0.013235
[16 empty] | 1500.3059 | 0.012855
[16 empty] | 1500.3062 | 0.012485
[16 empty] | 1500.3066 | 0.012125
[16 empty] | 1500.3069 | 0.011774
[16 empty] | 1500.3072 | 0.011433
[16 empty] | 1500.3075 | 0.011101
[16 empty] | 1500.3078 | 0.010778
[16 empty] | 1500.3082 | 0.010463
[16 empty] | 1500.3085 | 0.010158
[16 empty] | 1500.3088 | 0.0098601
[16 empty] | 1500.3091 | 0.0095706
[16 empty] | 1500.3094 | 0.0092891
[16 empty] | 1500.3098 | 0.0090153
[16 empty] | 1500.3101 | 0.008749
[16 empty] | 1500.3104 | 0.00849
[16 empty] | 1500.3107 | 0.0082382
[16 empty] | 1500.311 | 0.0079933
[16 empty] | 1500.3114 | 0.0077552
[16 empty] | 1500.3117 | 0.0075237
[16 empty] | 1500.312 | 0.0072987
[16 empty] | 1500.3123 | 0.0070799
[16 empty] | 1500.3126 | 0.0068673
[16 empty] | 1500.313 | 0.0066606
[16 empty] | 1500.3133 | 0.0064597
[16 empty] | 1500.3136 | 0.0062645
[16 empty] | 1500.3139 | 0.0060747
[16 empty] | 1500.3142 | 0.0058904
[16 empty] | 1500.3146 | 0.0057113
[16 empty] | 1500.3149 | 0.0055373
[16 empty] | 1500.3152 | 0.0053682
[16 empty] | 1500.3155 | 0.005204
[16 empty] | 1500.3158 | 0.0050444
[16 empty] | 1500.3162 | 0.0048895
[16 empty] | 1500.3165 | 0.004739
[16 empty] | 1500.3168 | 0.0045928
[16 empty] | 1500.3171 | 0.0044509
[16 empty] | 1500.3174 | 0.004313
[16 empty] | 1500.3178 | 0.0041792
[16 empty] | 1500.3181 | 0.0040493
[16 empty] | 1500.3184 | 0.0039231
[16 empty] | 1500.3187 | 0.0038007
[16 empty] | 1500.319 | 0.0036818
[16 empty] | 1500.3194 | 0.0035664
[16 empty] | 1500.3197 | 0.0034544

## GR Sample Size
Determing In-Tol Probability from historic data
| Reliability Target (Rt) = | 95%
| Confidence about the reliability = | 95%
| Sample Size Required = | 59
How this works:
If you want to know how much historical data is required to make a determination about "In-Tolerance" probabilities, it's really simple and it's the same numeric method for determining reliability of any instrument using interval analysis.
1. Enter a reliability target (usually 95% but you can select anything you'd like)
2. Enter a confidence level surrounding your reliability target. 90% is sufficient but you can enter anything you'd like. Just be aware the more "certain" you want to be about your decision of accepting EOPR, more samples will be required.
For example, if you wanted to be 99% certain, your EOPR is 95.45%, then 99 samples are required to demonstrate that metric. And by taking it to the extreme, if you wanted 99.9% EOPR at 99.9% confidence, then 6,905 samples are required.
Also worth mentioning, this need not apply to a single asset (although it could), it can be applied to either a make/model, or an instrument class

## Specific RISK
A1 | B | C | D | SELECTION OF GUARDBAND METHOD
2 | E = mc3 Solutions | Reported Result | Acceptance Limit | Choose Decision Rule >>>>>>> |  |  | Simple Acceptance | 1 | Custom Acceptance Limit | 0.5 | SR | ILAC G8:2009 Decision Rule(95%) | 0.804 | 0.38 x ln(TUR) - 0.54) | M2% | Acceptance Limit (A2%) - Method 6
3 | Nominal Value | 100 | 100 | [7 empty] | SR | ANSI Z540.3 Method 5(95.45%) | 0.8 | 0.071586 | -3.421% | 1.0068
4 | Lower Specification Limit | 99 | 99 | [7 empty] | SR | Simple Acceptance | 1
5 | Upper Specification Limit | 101 | 101 | [7 empty] | SR | Custom | 0.5
6 | Measured Value | 100 | 100 | [7 empty] | SR | U95 SINGLE SIDED (+ USL) | 0.8
7 | Std. Uncert. (k=1) | 0.1 | 0.1 | [7 empty] | SR | U99 | 0.7
8 | Total Risk | 7.62e-22% | Method 6 < 2 % PFA | [7 empty] | GR | RSS | 0.9798
9 | Upper Limit Risk | 7.62e-22% | 7.62e-22% | [7 empty] | GR | ANSI Z540.3 Method 6 | 1.0068
10 | Lower Limit Risk | 0% | 0%
11 | Test Uncertainty Ratio (TUR) = | 5 | 5
12 | Measurement Capability Index (Cm)= | 5 | 3.3333
13 | Process Capability (Cpk) | 3.3333
14 | Probability of Conformance (pc) | 100%
15 | Probability of NonConformance (1 - pc) | 0%
Area below for calculations | [10 empty] | SR
| Sample | Measurement
| 1 | 0
| 2 | 0
| 3 | 0
| 4 | 0
| 5 | 0
| Sample Mean | 0
| Sample Standard Deviation | 0
| Total Risk | =IF(Q16="SR",SUM(C9:C10),"Global Risk")
| Upper Limit Risk | =IF(Q16="SR",(NORMDIST(C6,C5,C7,1)),"Global Risk")
| Lower Limit Risk | =IF(Q16="SR",1-NORMDIST(C6,C4,C7,1),"Global Risk")
| Test Uncertainty Ratio (TUR) = | =((C5-C4))/(4*C7)
| Measurement Capability Index (Cm)= | =((C5-C4))/(4*C7)
| Process Capability (Cpk) | =MIN(((C5-C6)/(3*C7)),((C6-C4)/(3*C7)))
| Probability of Conformance (pc) | =NORM.DIST(C5,C6,C7,TRUE)-NORM.DIST(C4,C6,C7,TRUE)
| Probability of NonConformance (1 - pc) | =1-C14
| GuardBand | ± GB | Lower Acceptence Limit | Upper Acceptence Limit
| ILAC G8:2009 Decision Rule(95%) | 0.804 | 99.196 | 100.804
| ANSI Z540.3 Method 5(95.45%) | 0.8 | 99.2 | 100.8
| Simple Acceptance | 1 | 99 | 101
| Custom | 0.5 | 99.5 | 100.5
| U95 SINGLE SIDED (+ USL) | 0.8 | 99.2 | 100.8
| U99 | 0.7 | 99.3 | 100.7
| RSS | 0.9798 |  | 100.9798
| ANSI Z540.3 Method 6 | 1.0068 | 98.9932 | 101.0068
Risk Below Min. Allow. |  | 0 |  |  |  |  | As-Found
Risk Above Max. Allow. |  | 0 |  | Nominal Value | LSL | USL | MV
|  |  |  | 100 | 99 | 101 | 100 | 0
|  |  |  | 100 | 99 | 101 | 100 | 3.9894
[5 empty] | LAL | UAL
| Increment | 0.0008 |  |  | 99 | 101 |  | 0
[5 empty] | 99 | 101 |  | 3.9894
Norm. Dist.
| x | y
| 99.6 | 0.0013383
| 99.6008 | 0.0013818
| 99.6016 | 0.0014266
| 99.6024 | 0.0014727
| 99.6032 | 0.0015203 |  | NORMAL DISTRIBUTION FUNCTION
| 99.604 | 0.0015693
| 99.6048 | 0.0016197 |  | x = | x-axis value
| 99.6056 | 0.0016717 |  | u = | Centerline of distribution
| 99.6064 | 0.0017252 |  | σ = | Std. Dev.
| 99.6072 | 0.0017803
| 99.608 | 0.0018371
| 99.6088 | 0.0018956
| 99.6096 | 0.0019558
| 99.6104 | 0.0020178
| 99.6112 | 0.0020816
| 99.612 | 0.0021473
| 99.6128 | 0.0022149
| 99.6136 | 0.0022845
| 99.6144 | 0.0023562
| 99.6152 | 0.0024299
| 99.616 | 0.0025058
| 99.6168 | 0.0025839
| 99.6176 | 0.0026642
| 99.6184 | 0.0027469
| 99.6192 | 0.002832
| 99.62 | 0.0029195
| 99.6208 | 0.0030095
| 99.6216 | 0.0031021
| 99.6224 | 0.0031973
| 99.6232 | 0.0032953
| 99.624 | 0.003396
| 99.6248 | 0.0034996
| 99.6256 | 0.0036061
| 99.6264 | 0.0037157
| 99.6272 | 0.0038283
| 99.628 | 0.003944
| 99.6288 | 0.004063
| 99.6296 | 0.0041854
| 99.6304 | 0.0043111
| 99.6312 | 0.0044403
| 99.632 | 0.0045731
| 99.6328 | 0.0047096
| 99.6336 | 0.0048499
| 99.6344 | 0.004994
| 99.6352 | 0.005142
| 99.636 | 0.0052941
| 99.6368 | 0.0054504
| 99.6376 | 0.0056109
| 99.6384 | 0.0057758
| 99.6392 | 0.0059451
| 99.64 | 0.006119
| 99.6408 | 0.0062976
| 99.6416 | 0.006481
| 99.6424 | 0.0066693
| 99.6432 | 0.0068626
| 99.644 | 0.0070611
| 99.6448 | 0.0072649
| 99.6456 | 0.007474
| 99.6464 | 0.0076887
| 99.6472 | 0.0079091
| 99.648 | 0.0081352
| 99.6488 | 0.0083673
| 99.6496 | 0.0086054
| 99.6504 | 0.0088498
| 99.6512 | 0.0091005
| 99.652 | 0.0093577
| 99.6528 | 0.0096216
| 99.6536 | 0.0098923
| 99.6544 | 0.01017
| 99.6552 | 0.010455
| 99.656 | 0.010747
| 99.6568 | 0.011046
| 99.6576 | 0.011353
| 99.6584 | 0.011668
| 99.6592 | 0.011991
| 99.66 | 0.012322
| 99.6608 | 0.012662
| 99.6616 | 0.013009
| 99.6624 | 0.013366
| 99.6632 | 0.013731
| 99.664 | 0.014106
| 99.6648 | 0.01449
| 99.6656 | 0.014883
| 99.6664 | 0.015286
| 99.6672 | 0.015699
| 99.668 | 0.016122
| 99.6688 | 0.016556
| 99.6696 | 0.017
| 99.6704 | 0.017454
| 99.6712 | 0.01792
| 99.672 | 0.018397
| 99.6728 | 0.018886
| 99.6736 | 0.019386
| 99.6744 | 0.019898
| 99.6752 | 0.020423
| 99.676 | 0.02096
| 99.6768 | 0.021509
| 99.6776 | 0.022072
| 99.6784 | 0.022648
| 99.6792 | 0.023238
| 99.68 | 0.023841
| 99.6808 | 0.024458
| 99.6816 | 0.02509
| 99.6824 | 0.025737
| 99.6832 | 0.026398
| 99.684 | 0.027075
| 99.6848 | 0.027767
| 99.6856 | 0.028475
| 99.6864 | 0.0292
| 99.6872 | 0.02994
| 99.688 | 0.030698
| 99.6888 | 0.031473
| 99.6896 | 0.032265
| 99.6904 | 0.033076
| 99.6912 | 0.033904
| 99.692 | 0.034751
| 99.6928 | 0.035617
| 99.6936 | 0.036502
| 99.6944 | 0.037406
| 99.6952 | 0.038331
| 99.696 | 0.039276
| 99.6968 | 0.040241
| 99.6976 | 0.041228
| 99.6984 | 0.042236
| 99.6992 | 0.043266
| 99.7 | 0.044318
| 99.7008 | 0.045394
| 99.7016 | 0.046492
| 99.7024 | 0.047613
| 99.7032 | 0.048759
| 99.704 | 0.049929
| 99.7048 | 0.051124
| 99.7056 | 0.052344
| 99.7064 | 0.05359
| 99.7072 | 0.054861
| 99.708 | 0.05616
| 99.7088 | 0.057485
| 99.7096 | 0.058838
| 99.7104 | 0.060219
| 99.7112 | 0.061629
| 99.712 | 0.063067
| 99.7128 | 0.064535
| 99.7136 | 0.066033
| 99.7144 | 0.067561
| 99.7152 | 0.06912
| 99.716 | 0.070711
| 99.7168 | 0.072334
| 99.7176 | 0.073989
| 99.7184 | 0.075677
| 99.7192 | 0.077399
| 99.72 | 0.079155
| 99.7208 | 0.080945
| 99.7216 | 0.082771
| 99.7224 | 0.084632
| 99.7232 | 0.08653
| 99.724 | 0.088465
| 99.7248 | 0.090437
| 99.7256 | 0.092447
| 99.7264 | 0.094496
| 99.7272 | 0.096584
| 99.728 | 0.098712
| 99.7288 | 0.10088
| 99.7296 | 0.10309
| 99.7304 | 0.10534
| 99.7312 | 0.10763
| 99.732 | 0.10997
| 99.7328 | 0.11235
| 99.7336 | 0.11477
| 99.7344 | 0.11724
| 99.7352 | 0.11976
| 99.736 | 0.12232
| 99.7368 | 0.12492
| 99.7376 | 0.12758
| 99.7384 | 0.13028
| 99.7392 | 0.13303
| 99.74 | 0.13583
| 99.7408 | 0.13868
| 99.7416 | 0.14158
| 99.7424 | 0.14453
| 99.7432 | 0.14754
| 99.744 | 0.1506
| 99.7448 | 0.15371
| 99.7456 | 0.15687
| 99.7464 | 0.16009
| 99.7472 | 0.16337
| 99.748 | 0.1667
| 99.7488 | 0.17009
| 99.7496 | 0.17354
| 99.7504 | 0.17704
| 99.7512 | 0.18061
| 99.752 | 0.18423
| 99.7528 | 0.18792
| 99.7536 | 0.19167
| 99.7544 | 0.19548
| 99.7552 | 0.19935
| 99.756 | 0.20328
| 99.7568 | 0.20728
| 99.7576 | 0.21135
| 99.7584 | 0.21548
| 99.7592 | 0.21968
| 99.76 | 0.22395
| 99.7608 | 0.22828
| 99.7616 | 0.23268
| 99.7624 | 0.23715
| 99.7632 | 0.2417
| 99.764 | 0.24631
| 99.7648 | 0.251
| 99.7656 | 0.25576
| 99.7664 | 0.26059
| 99.7672 | 0.2655
| 99.768 | 0.27048
| 99.7688 | 0.27554
| 99.7696 | 0.28067
| 99.7704 | 0.28589
| 99.7712 | 0.29118
| 99.772 | 0.29655
| 99.7728 | 0.30199
| 99.7736 | 0.30752
| 99.7744 | 0.31313
| 99.7752 | 0.31883
| 99.776 | 0.3246
| 99.7768 | 0.33046
| 99.7776 | 0.3364
| 99.7784 | 0.34243
| 99.7792 | 0.34855
| 99.78 | 0.35475
| 99.7808 | 0.36103
| 99.7816 | 0.36741
| 99.7824 | 0.37387
| 99.7832 | 0.38043
| 99.784 | 0.38707
| 99.7848 | 0.3938
| 99.7856 | 0.40063
| 99.7864 | 0.40755
| 99.7872 | 0.41456
| 99.788 | 0.42166
| 99.7888 | 0.42886
| 99.7896 | 0.43615
| 99.7904 | 0.44354
| 99.7912 | 0.45103
| 99.792 | 0.45861
| 99.7928 | 0.46629
| 99.7936 | 0.47407
| 99.7944 | 0.48195
| 99.7952 | 0.48992
| 99.796 | 0.498
| 99.7968 | 0.50618
| 99.7976 | 0.51446
| 99.7984 | 0.52284
| 99.7992 | 0.53132
| 99.8 | 0.53991
| 99.8008 | 0.5486
| 99.8016 | 0.55739
| 99.8024 | 0.56629
| 99.8032 | 0.5753
| 99.804 | 0.58441
| 99.8048 | 0.59363
| 99.8056 | 0.60295
| 99.8064 | 0.61238
| 99.8072 | 0.62192
| 99.808 | 0.63157
| 99.8088 | 0.64132
| 99.8096 | 0.65119
| 99.8104 | 0.66116
| 99.8112 | 0.67124
| 99.812 | 0.68144
| 99.8128 | 0.69174
| 99.8136 | 0.70215
| 99.8144 | 0.71268
| 99.8152 | 0.72332
| 99.816 | 0.73407
| 99.8168 | 0.74493
| 99.8176 | 0.7559
| 99.8184 | 0.76699
| 99.8192 | 0.77819
| 99.82 | 0.7895
| 99.8208 | 0.80093
| 99.8216 | 0.81247
| 99.8224 | 0.82412
| 99.8232 | 0.83588
| 99.824 | 0.84776
| 99.8248 | 0.85976
| 99.8256 | 0.87186
| 99.8264 | 0.88409
| 99.8272 | 0.89642
| 99.828 | 0.90887
| 99.8288 | 0.92143
| 99.8296 | 0.93411
| 99.8304 | 0.9469
| 99.8312 | 0.9598
| 99.832 | 0.97282
| 99.8328 | 0.98595
| 99.8336 | 0.9992
| 99.8344 | 1.0126
| 99.8352 | 1.026
| 99.836 | 1.0396
| 99.8368 | 1.0533
| 99.8376 | 1.0671
| 99.8384 | 1.081
| 99.8392 | 1.0951
| 99.84 | 1.1092
| 99.8408 | 1.1235
| 99.8416 | 1.1378
| 99.8424 | 1.1523
| 99.8432 | 1.1669
| 99.844 | 1.1816
| 99.8448 | 1.1964
| 99.8456 | 1.2113
| 99.8464 | 1.2263
| 99.8472 | 1.2414
| 99.848 | 1.2566
| 99.8488 | 1.272
| 99.8496 | 1.2874
| 99.8504 | 1.303
| 99.8512 | 1.3186
| 99.852 | 1.3344
| 99.8528 | 1.3502
| 99.8536 | 1.3662
| 99.8544 | 1.3822
| 99.8552 | 1.3984
| 99.856 | 1.4146
| 99.8568 | 1.4309
| 99.8576 | 1.4474
| 99.8584 | 1.4639
| 99.8592 | 1.4806
| 99.86 | 1.4973
| 99.8608 | 1.5141
| 99.8616 | 1.531
| 99.8624 | 1.548
| 99.8632 | 1.5651
| 99.864 | 1.5822
| 99.8648 | 1.5995
| 99.8656 | 1.6168
| 99.8664 | 1.6343
| 99.8672 | 1.6518
| 99.868 | 1.6694
| 99.8688 | 1.687
| 99.8696 | 1.7048
| 99.8704 | 1.7226
| 99.8712 | 1.7405
| 99.872 | 1.7585
| 99.8728 | 1.7765
| 99.8736 | 1.7946
| 99.8744 | 1.8128
| 99.8752 | 1.8311
| 99.876 | 1.8494
| 99.8768 | 1.8678
| 99.8776 | 1.8862
| 99.8784 | 1.9047
| 99.8792 | 1.9232
| 99.88 | 1.9419
| 99.8808 | 1.9605
| 99.8816 | 1.9793
| 99.8824 | 1.998
| 99.8832 | 2.0168
| 99.884 | 2.0357
| 99.8848 | 2.0546
| 99.8856 | 2.0736
| 99.8864 | 2.0926
| 99.8872 | 2.1116
| 99.888 | 2.1307
| 99.8888 | 2.1498
| 99.8896 | 2.1689
| 99.8904 | 2.1881
| 99.8912 | 2.2073
| 99.892 | 2.2265
| 99.8928 | 2.2458
| 99.8936 | 2.2651
| 99.8944 | 2.2843
| 99.8952 | 2.3036
| 99.896 | 2.323
| 99.8968 | 2.3423
| 99.8976 | 2.3616
| 99.8984 | 2.381
| 99.8992 | 2.4003
| 99.9 | 2.4197
| 99.9008 | 2.4391
| 99.9016 | 2.4584
| 99.9024 | 2.4778
| 99.9032 | 2.4971
| 99.904 | 2.5164
| 99.9048 | 2.5358
| 99.9056 | 2.5551
| 99.9064 | 2.5744
| 99.9072 | 2.5936
| 99.908 | 2.6129
| 99.9088 | 2.6321
| 99.9096 | 2.6513
| 99.9104 | 2.6704
| 99.9112 | 2.6896
| 99.912 | 2.7086
| 99.9128 | 2.7277
| 99.9136 | 2.7467
| 99.9144 | 2.7657
| 99.9152 | 2.7846
| 99.916 | 2.8034
| 99.9168 | 2.8223
| 99.9176 | 2.841
| 99.9184 | 2.8597
| 99.9192 | 2.8783
| 99.92 | 2.8969
| 99.9208 | 2.9154
| 99.9216 | 2.9339
| 99.9224 | 2.9522
| 99.9232 | 2.9705
| 99.924 | 2.9887
| 99.9248 | 3.0069
| 99.9256 | 3.0249
| 99.9264 | 3.0429
| 99.9272 | 3.0607
| 99.928 | 3.0785
| 99.9288 | 3.0962
| 99.9296 | 3.1138
| 99.9304 | 3.1313
| 99.9312 | 3.1487
| 99.932 | 3.1659
| 99.9328 | 3.1831
| 99.9336 | 3.2002
| 99.9344 | 3.2171
| 99.9352 | 3.2339
| 99.936 | 3.2506
| 99.9368 | 3.2672
| 99.9376 | 3.2837
| 99.9384 | 3.3
| 99.9392 | 3.3162
| 99.94 | 3.3322
| 99.9408 | 3.3482
| 99.9416 | 3.364
| 99.9424 | 3.3796
| 99.9432 | 3.3951
| 99.944 | 3.4105
| 99.9448 | 3.4257
| 99.9456 | 3.4407
| 99.9464 | 3.4556
| 99.9472 | 3.4703
| 99.948 | 3.4849
| 99.9488 | 3.4993
| 99.9496 | 3.5136
| 99.9504 | 3.5277
| 99.9512 | 3.5416
| 99.952 | 3.5553
| 99.9528 | 3.5689
| 99.9536 | 3.5823
| 99.9544 | 3.5955
| 99.9552 | 3.6085
| 99.956 | 3.6213
| 99.9568 | 3.634
| 99.9576 | 3.6465
| 99.9584 | 3.6587
| 99.9592 | 3.6708
| 99.96 | 3.6827
| 99.9608 | 3.6944
| 99.9616 | 3.7059
| 99.9624 | 3.7172
| 99.9632 | 3.7282
| 99.964 | 3.7391
| 99.9648 | 3.7498
| 99.9656 | 3.7602
| 99.9664 | 3.7705
| 99.9672 | 3.7805
| 99.968 | 3.7903
| 99.9688 | 3.7999
| 99.9696 | 3.8093
| 99.9704 | 3.8184
| 99.9712 | 3.8274
| 99.972 | 3.8361
| 99.9728 | 3.8445
| 99.9736 | 3.8528
| 99.9744 | 3.8608
| 99.9752 | 3.8686
| 99.976 | 3.8762
| 99.9768 | 3.8835
| 99.9776 | 3.8906
| 99.9784 | 3.8974
| 99.9792 | 3.9041
| 99.98 | 3.9104
| 99.9808 | 3.9166
| 99.9816 | 3.9225
| 99.9824 | 3.9281
| 99.9832 | 3.9335
| 99.984 | 3.9387
| 99.9848 | 3.9436
| 99.9856 | 3.9483
| 99.9864 | 3.9527
| 99.9872 | 3.9569
| 99.988 | 3.9608
| 99.9888 | 3.9645
| 99.9896 | 3.9679
| 99.9904 | 3.9711
| 99.9912 | 3.974
| 99.992 | 3.9767
| 99.9928 | 3.9791
| 99.9936 | 3.9813
| 99.9944 | 3.9832
| 99.9952 | 3.9848
| 99.996 | 3.9862
| 99.9968 | 3.9874
| 99.9976 | 3.9883
| 99.9984 | 3.9889
| 99.9992 | 3.9893
| 100 | 3.9894
| 100.0008 | 3.9893
| 100.0016 | 3.9889
| 100.0024 | 3.9883
| 100.0032 | 3.9874
| 100.004 | 3.9862
| 100.0048 | 3.9848
| 100.0056 | 3.9832
| 100.0064 | 3.9813
| 100.0072 | 3.9791
| 100.008 | 3.9767
| 100.0088 | 3.974
| 100.0096 | 3.9711
| 100.0104 | 3.9679
| 100.0112 | 3.9645
| 100.012 | 3.9608
| 100.0128 | 3.9569
| 100.0136 | 3.9527
| 100.0144 | 3.9483
| 100.0152 | 3.9436
| 100.016 | 3.9387
| 100.0168 | 3.9335
| 100.0176 | 3.9281
| 100.0184 | 3.9225
| 100.0192 | 3.9166
| 100.02 | 3.9104
| 100.0208 | 3.9041
| 100.0216 | 3.8974
| 100.0224 | 3.8906
| 100.0232 | 3.8835
| 100.024 | 3.8762
| 100.0248 | 3.8686
| 100.0256 | 3.8608
| 100.0264 | 3.8528
| 100.0272 | 3.8445
| 100.028 | 3.8361
| 100.0288 | 3.8274
| 100.0296 | 3.8184
| 100.0304 | 3.8093
| 100.0312 | 3.7999
| 100.032 | 3.7903
| 100.0328 | 3.7805
| 100.0336 | 3.7705
| 100.0344 | 3.7602
| 100.0352 | 3.7498
| 100.036 | 3.7391
| 100.0368 | 3.7282
| 100.0376 | 3.7172
| 100.0384 | 3.7059
| 100.0392 | 3.6944
| 100.04 | 3.6827
| 100.0408 | 3.6708
| 100.0416 | 3.6587
| 100.0424 | 3.6465
| 100.0432 | 3.634
| 100.044 | 3.6213
| 100.0448 | 3.6085
| 100.0456 | 3.5955
| 100.0464 | 3.5823
| 100.0472 | 3.5689
| 100.048 | 3.5553
| 100.0488 | 3.5416
| 100.0496 | 3.5277
| 100.0504 | 3.5136
| 100.0512 | 3.4993
| 100.052 | 3.4849
| 100.0528 | 3.4703
| 100.0536 | 3.4556
| 100.0544 | 3.4407
| 100.0552 | 3.4257
| 100.056 | 3.4105
| 100.0568 | 3.3951
| 100.0576 | 3.3796
| 100.0584 | 3.364
| 100.0592 | 3.3482
| 100.06 | 3.3322
| 100.0608 | 3.3162
| 100.0616 | 3.3
| 100.0624 | 3.2837
| 100.0632 | 3.2672
| 100.064 | 3.2506
| 100.0648 | 3.2339
| 100.0656 | 3.2171
| 100.0664 | 3.2002
| 100.0672 | 3.1831
| 100.068 | 3.1659
| 100.0688 | 3.1487
| 100.0696 | 3.1313
| 100.0704 | 3.1138
| 100.0712 | 3.0962
| 100.072 | 3.0785
| 100.0728 | 3.0607
| 100.0736 | 3.0429
| 100.0744 | 3.0249
| 100.0752 | 3.0069
| 100.076 | 2.9887
| 100.0768 | 2.9705
| 100.0776 | 2.9522
| 100.0784 | 2.9339
| 100.0792 | 2.9154
| 100.08 | 2.8969
| 100.0808 | 2.8783
| 100.0816 | 2.8597
| 100.0824 | 2.841
| 100.0832 | 2.8223
| 100.084 | 2.8034
| 100.0848 | 2.7846
| 100.0856 | 2.7657
| 100.0864 | 2.7467
| 100.0872 | 2.7277
| 100.088 | 2.7086
| 100.0888 | 2.6896
| 100.0896 | 2.6704
| 100.0904 | 2.6513
| 100.0912 | 2.6321
| 100.092 | 2.6129
| 100.0928 | 2.5936
| 100.0936 | 2.5744
| 100.0944 | 2.5551
| 100.0952 | 2.5358
| 100.096 | 2.5164
| 100.0968 | 2.4971
| 100.0976 | 2.4778
| 100.0984 | 2.4584
| 100.0992 | 2.4391
| 100.1 | 2.4197
| 100.1008 | 2.4003
| 100.1016 | 2.381
| 100.1024 | 2.3616
| 100.1032 | 2.3423
| 100.104 | 2.323
| 100.1048 | 2.3036
| 100.1056 | 2.2843
| 100.1064 | 2.2651
| 100.1072 | 2.2458
| 100.108 | 2.2265
| 100.1088 | 2.2073
| 100.1096 | 2.1881
| 100.1104 | 2.1689
| 100.1112 | 2.1498
| 100.112 | 2.1307
| 100.1128 | 2.1116
| 100.1136 | 2.0926
| 100.1144 | 2.0736
| 100.1152 | 2.0546
| 100.116 | 2.0357
| 100.1168 | 2.0168
| 100.1176 | 1.998
| 100.1184 | 1.9793
| 100.1192 | 1.9605
| 100.12 | 1.9419
| 100.1208 | 1.9232
| 100.1216 | 1.9047
| 100.1224 | 1.8862
| 100.1232 | 1.8678
| 100.124 | 1.8494
| 100.1248 | 1.8311
| 100.1256 | 1.8128
| 100.1264 | 1.7946
| 100.1272 | 1.7765
| 100.128 | 1.7585
| 100.1288 | 1.7405
| 100.1296 | 1.7226
| 100.1304 | 1.7048
| 100.1312 | 1.687
| 100.132 | 1.6694
| 100.1328 | 1.6518
| 100.1336 | 1.6343
| 100.1344 | 1.6168
| 100.1352 | 1.5995
| 100.136 | 1.5822
| 100.1368 | 1.5651
| 100.1376 | 1.548
| 100.1384 | 1.531
| 100.1392 | 1.5141
| 100.14 | 1.4973
| 100.1408 | 1.4806
| 100.1416 | 1.4639
| 100.1424 | 1.4474
| 100.1432 | 1.4309
| 100.144 | 1.4146
| 100.1448 | 1.3984
| 100.1456 | 1.3822
| 100.1464 | 1.3662
| 100.1472 | 1.3502
| 100.148 | 1.3344
| 100.1488 | 1.3186
| 100.1496 | 1.303
| 100.1504 | 1.2874
| 100.1512 | 1.272
| 100.152 | 1.2566
| 100.1528 | 1.2414
| 100.1536 | 1.2263
| 100.1544 | 1.2113
| 100.1552 | 1.1964
| 100.156 | 1.1816
| 100.1568 | 1.1669
| 100.1576 | 1.1523
| 100.1584 | 1.1378
| 100.1592 | 1.1235
| 100.16 | 1.1092
| 100.1608 | 1.0951
| 100.1616 | 1.081
| 100.1624 | 1.0671
| 100.1632 | 1.0533
| 100.164 | 1.0396
| 100.1648 | 1.026
| 100.1656 | 1.0126
| 100.1664 | 0.9992
| 100.1672 | 0.98595
| 100.168 | 0.97282
| 100.1688 | 0.9598
| 100.1696 | 0.9469
| 100.1704 | 0.93411
| 100.1712 | 0.92143
| 100.172 | 0.90887
| 100.1728 | 0.89642
| 100.1736 | 0.88409
| 100.1744 | 0.87186
| 100.1752 | 0.85976
| 100.176 | 0.84776
| 100.1768 | 0.83588
| 100.1776 | 0.82412
| 100.1784 | 0.81247
| 100.1792 | 0.80093
| 100.18 | 0.7895
| 100.1808 | 0.77819
| 100.1816 | 0.76699
| 100.1824 | 0.7559
| 100.1832 | 0.74493
| 100.184 | 0.73407
| 100.1848 | 0.72332
| 100.1856 | 0.71268
| 100.1864 | 0.70215
| 100.1872 | 0.69174
| 100.188 | 0.68144
| 100.1888 | 0.67124
| 100.1896 | 0.66116
| 100.1904 | 0.65119
| 100.1912 | 0.64132
| 100.192 | 0.63157
| 100.1928 | 0.62192
| 100.1936 | 0.61238
| 100.1944 | 0.60295
| 100.1952 | 0.59363
| 100.196 | 0.58441
| 100.1968 | 0.5753
| 100.1976 | 0.56629
| 100.1984 | 0.55739
| 100.1992 | 0.5486
| 100.2 | 0.53991
| 100.2008 | 0.53132
| 100.2016 | 0.52284
| 100.2024 | 0.51446
| 100.2032 | 0.50618
| 100.204 | 0.498
| 100.2048 | 0.48992
| 100.2056 | 0.48195
| 100.2064 | 0.47407
| 100.2072 | 0.46629
| 100.208 | 0.45861
| 100.2088 | 0.45103
| 100.2096 | 0.44354
| 100.2104 | 0.43615
| 100.2112 | 0.42886
| 100.212 | 0.42166
| 100.2128 | 0.41456
| 100.2136 | 0.40755
| 100.2144 | 0.40063
| 100.2152 | 0.3938
| 100.216 | 0.38707
| 100.2168 | 0.38043
| 100.2176 | 0.37387
| 100.2184 | 0.36741
| 100.2192 | 0.36103
| 100.22 | 0.35475
| 100.2208 | 0.34855
| 100.2216 | 0.34243
| 100.2224 | 0.3364
| 100.2232 | 0.33046
| 100.224 | 0.3246
| 100.2248 | 0.31883
| 100.2256 | 0.31313
| 100.2264 | 0.30752
| 100.2272 | 0.30199
| 100.228 | 0.29655
| 100.2288 | 0.29118
| 100.2296 | 0.28589
| 100.2304 | 0.28067
| 100.2312 | 0.27554
| 100.232 | 0.27048
| 100.2328 | 0.2655
| 100.2336 | 0.26059
| 100.2344 | 0.25576
| 100.2352 | 0.251
| 100.236 | 0.24631
| 100.2368 | 0.2417
| 100.2376 | 0.23715
| 100.2384 | 0.23268
| 100.2392 | 0.22828
| 100.24 | 0.22395
| 100.2408 | 0.21968
| 100.2416 | 0.21548
| 100.2424 | 0.21135
| 100.2432 | 0.20728
| 100.244 | 0.20328
| 100.2448 | 0.19935
| 100.2456 | 0.19548
| 100.2464 | 0.19167
| 100.2472 | 0.18792
| 100.248 | 0.18423
| 100.2488 | 0.18061
| 100.2496 | 0.17704
| 100.2504 | 0.17354
| 100.2512 | 0.17009
| 100.252 | 0.1667
| 100.2528 | 0.16337
| 100.2536 | 0.16009
| 100.2544 | 0.15687
| 100.2552 | 0.15371
| 100.256 | 0.1506
| 100.2568 | 0.14754
| 100.2576 | 0.14453
| 100.2584 | 0.14158
| 100.2592 | 0.13868
| 100.26 | 0.13583
| 100.2608 | 0.13303
| 100.2616 | 0.13028
| 100.2624 | 0.12758
| 100.2632 | 0.12492
| 100.264 | 0.12232
| 100.2648 | 0.11976
| 100.2656 | 0.11724
| 100.2664 | 0.11477
| 100.2672 | 0.11235
| 100.268 | 0.10997
| 100.2688 | 0.10763
| 100.2696 | 0.10534
| 100.2704 | 0.10309
| 100.2712 | 0.10088
| 100.272 | 0.098712
| 100.2728 | 0.096584
| 100.2736 | 0.094496
| 100.2744 | 0.092447
| 100.2752 | 0.090437
| 100.276 | 0.088465
| 100.2768 | 0.08653
| 100.2776 | 0.084632
| 100.2784 | 0.082771
| 100.2792 | 0.080945
| 100.28 | 0.079155
| 100.2808 | 0.077399
| 100.2816 | 0.075677
| 100.2824 | 0.073989
| 100.2832 | 0.072334
| 100.284 | 0.070711
| 100.2848 | 0.06912
| 100.2856 | 0.067561
| 100.2864 | 0.066033
| 100.2872 | 0.064535
| 100.288 | 0.063067
| 100.2888 | 0.061629
| 100.2896 | 0.060219
| 100.2904 | 0.058838
| 100.2912 | 0.057485
| 100.292 | 0.05616
| 100.2928 | 0.054861
| 100.2936 | 0.05359
| 100.2944 | 0.052344
| 100.2952 | 0.051124
| 100.296 | 0.049929
| 100.2968 | 0.048759
| 100.2976 | 0.047613
| 100.2984 | 0.046492
| 100.2992 | 0.045394
| 100.3 | 0.044318
| 100.3008 | 0.043266
| 100.3016 | 0.042236
| 100.3024 | 0.041228
| 100.3032 | 0.040241
| 100.304 | 0.039276
| 100.3048 | 0.038331
| 100.3056 | 0.037406
| 100.3064 | 0.036502
| 100.3072 | 0.035617
| 100.308 | 0.034751
| 100.3088 | 0.033904
| 100.3096 | 0.033076
| 100.3104 | 0.032265
| 100.3112 | 0.031473
| 100.312 | 0.030698
| 100.3128 | 0.02994
| 100.3136 | 0.0292
| 100.3144 | 0.028475
| 100.3152 | 0.027767
| 100.316 | 0.027075
| 100.3168 | 0.026398
| 100.3176 | 0.025737
| 100.3184 | 0.02509
| 100.3192 | 0.024458
| 100.32 | 0.023841
| 100.3208 | 0.023238
| 100.3216 | 0.022648
| 100.3224 | 0.022072
| 100.3232 | 0.021509
| 100.324 | 0.02096
| 100.3248 | 0.020423
| 100.3256 | 0.019898
| 100.3264 | 0.019386
| 100.3272 | 0.018886
| 100.328 | 0.018397
| 100.3288 | 0.01792
| 100.3296 | 0.017454
| 100.3304 | 0.017
| 100.3312 | 0.016556
| 100.332 | 0.016122
| 100.3328 | 0.015699
| 100.3336 | 0.015286
| 100.3344 | 0.014883
| 100.3352 | 0.01449
| 100.336 | 0.014106
| 100.3368 | 0.013731
| 100.3376 | 0.013366
| 100.3384 | 0.013009
| 100.3392 | 0.012662
| 100.34 | 0.012322
| 100.3408 | 0.011991
| 100.3416 | 0.011668
| 100.3424 | 0.011353
| 100.3432 | 0.011046
| 100.344 | 0.010747
| 100.3448 | 0.010455
| 100.3456 | 0.01017
| 100.3464 | 0.0098923
| 100.3472 | 0.0096216
| 100.348 | 0.0093577
| 100.3488 | 0.0091005
| 100.3496 | 0.0088498
| 100.3504 | 0.0086054
| 100.3512 | 0.0083673
| 100.352 | 0.0081352
| 100.3528 | 0.0079091
| 100.3536 | 0.0076887
| 100.3544 | 0.007474
| 100.3552 | 0.0072649
| 100.356 | 0.0070611
| 100.3568 | 0.0068626
| 100.3576 | 0.0066693
| 100.3584 | 0.006481
| 100.3592 | 0.0062976
| 100.36 | 0.006119
| 100.3608 | 0.0059451
| 100.3616 | 0.0057758
| 100.3624 | 0.0056109
| 100.3632 | 0.0054504
| 100.364 | 0.0052941
| 100.3648 | 0.005142
| 100.3656 | 0.004994
| 100.3664 | 0.0048499
| 100.3672 | 0.0047096
| 100.368 | 0.0045731
| 100.3688 | 0.0044403
| 100.3696 | 0.0043111
| 100.3704 | 0.0041854
| 100.3712 | 0.004063
| 100.372 | 0.003944
| 100.3728 | 0.0038283
| 100.3736 | 0.0037157
| 100.3744 | 0.0036061
| 100.3752 | 0.0034996
| 100.376 | 0.003396
| 100.3768 | 0.0032953
| 100.3776 | 0.0031973
| 100.3784 | 0.0031021
| 100.3792 | 0.0030095
| 100.38 | 0.0029195
| 100.3808 | 0.002832
| 100.3816 | 0.0027469
| 100.3824 | 0.0026642
| 100.3832 | 0.0025839
| 100.384 | 0.0025058
| 100.3848 | 0.0024299
| 100.3856 | 0.0023562
| 100.3864 | 0.0022845
| 100.3872 | 0.0022149
| 100.388 | 0.0021473
| 100.3888 | 0.0020816
| 100.3896 | 0.0020178
| 100.3904 | 0.0019558
| 100.3912 | 0.0018956
| 100.392 | 0.0018371
| 100.3928 | 0.0017803
| 100.3936 | 0.0017252
| 100.3944 | 0.0016717
| 100.3952 | 0.0016197
| 100.396 | 0.0015693
| 100.3968 | 0.0015203
| 100.3976 | 0.0014727
| 100.3984 | 0.0014266
| 100.3992 | 0.0013818
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
Note on B2: Dilip Shah: This Worksheet is the intellectual Property of E= mc3 Solutions. Permission is granted to the E=mc3 Solution Clients who have validly purchased this spreadsheet. All others should contact at emc3solu@aol.com (call 330-328-4400) for purchasing a license.
Note on H2: Dilip Shah: Choose the Guardband method from the Dropdown menu.
Note on I2: Dilip Shah: Choose the Guardband calculation method from the dropdown menu on the left.
Note on A16: Dilip Shah: The yellow background area is there for you to do calculations.

## GB Multiplier
Risk Calculator |  | Conformance Probability Table
Uppper Tolerance TU | 1500.2 | Conformance Probability , Pc | Guard Band Multipier, r Two Sided
Lower Tolerance TL | 1499.8
Nominal Value | 1500 | 0.0668 | -0.75003
Measurement Unc um | 0.04 | 0.159 | -0.49929
Measured Value xm | 1500.0887 | 0.3085 | -0.25005
Tolerance T | 0.4 | 0.5 | 0
|  | 0.6914 | 0.24991
|  | 0.8 | 0.42081
Probability of Conformance (pc) | 99.73% | 0.85 | 0.51822
Probability of NonConformance (1 - pc) | 0.2697% | 0.9 | 0.64078
|  | 0.95 | 0.82243
Setting the Guard Band Upper and Lower pc |  | 0.9545 | 0.84507
Select Desired Conformance Probability | 95% | 0.97725 | 1
Risk if within GU & GL | 5% | 0.98 | 1.0269
hU (GB Multiplier) | 0.82243 | 0.99 | 1.1632 |  | 0.82243
Guard Band Upper Gu | 1500.1342 | 0.9973 | 1.3911
Guard Band Lower GL | 1499.8658 | 0.999 | 3.0902
|  | 1 | 5.9978
0.065794
[5 empty] | 0 |  |  |  | As-Found
[5 empty] | 0 | Nominal Value | LSL | USL | MV
[6 empty] | 1500 | 1499.8 | 1500.2 | 1500.0887 | 0
[6 empty] | 1500 | 1499.8 | 1500.2 | 1500.0887 | 9.9736
[7 empty] | LAL | UAL
|  |  |  | Increment | 0.00032 |  | 1499.8658 | 1500.1342 |  | 0
[7 empty] | 1499.8658 | 1500.1342 |  | 9.9736
Norm. Dist.
|  |  |  | x | y
|  |  |  | 1499.9287 | 0.0033458
|  |  |  | 1499.929 | 0.0034544
|  |  |  | 1499.9293 | 0.0035664
|  |  |  | 1499.9297 | 0.0036818
|  |  |  | 1499.93 | 0.0038007 | NORMAL DISTRIBUTION FUNCTION
|  |  |  | 1499.9303 | 0.0039231
|  |  |  | 1499.9306 | 0.0040493 | x = | x-axis value
|  |  |  | 1499.9309 | 0.0041792 | u = | Centerline of distribution
|  |  |  | 1499.9313 | 0.004313 | σ = | Std. Dev.
|  |  |  | 1499.9316 | 0.0044509
|  |  |  | 1499.9319 | 0.0045928
|  |  |  | 1499.9322 | 0.004739
|  |  |  | 1499.9325 | 0.0048895
|  |  |  | 1499.9329 | 0.0050444
|  |  |  | 1499.9332 | 0.005204
|  |  |  | 1499.9335 | 0.0053682
|  |  |  | 1499.9338 | 0.0055373
|  |  |  | 1499.9341 | 0.0057113
|  |  |  | 1499.9345 | 0.0058904
|  |  |  | 1499.9348 | 0.0060747
|  |  |  | 1499.9351 | 0.0062645
|  |  |  | 1499.9354 | 0.0064597
|  |  |  | 1499.9357 | 0.0066606
|  |  |  | 1499.9361 | 0.0068673
|  |  |  | 1499.9364 | 0.0070799
|  |  |  | 1499.9367 | 0.0072987
|  |  |  | 1499.937 | 0.0075237
|  |  |  | 1499.9373 | 0.0077552
|  |  |  | 1499.9377 | 0.0079933
|  |  |  | 1499.938 | 0.0082382
|  |  |  | 1499.9383 | 0.00849
|  |  |  | 1499.9386 | 0.008749
|  |  |  | 1499.9389 | 0.0090153
|  |  |  | 1499.9393 | 0.0092891
|  |  |  | 1499.9396 | 0.0095706
|  |  |  | 1499.9399 | 0.0098601
|  |  |  | 1499.9402 | 0.010158
|  |  |  | 1499.9405 | 0.010463
|  |  |  | 1499.9409 | 0.010778
|  |  |  | 1499.9412 | 0.011101
|  |  |  | 1499.9415 | 0.011433
|  |  |  | 1499.9418 | 0.011774
|  |  |  | 1499.9421 | 0.012125
|  |  |  | 1499.9425 | 0.012485
|  |  |  | 1499.9428 | 0.012855
|  |  |  | 1499.9431 | 0.013235
|  |  |  | 1499.9434 | 0.013626
|  |  |  | 1499.9437 | 0.014027
|  |  |  | 1499.9441 | 0.014439
|  |  |  | 1499.9444 | 0.014863
|  |  |  | 1499.9447 | 0.015298
|  |  |  | 1499.945 | 0.015744
|  |  |  | 1499.9453 | 0.016202
|  |  |  | 1499.9457 | 0.016673
|  |  |  | 1499.946 | 0.017157
|  |  |  | 1499.9463 | 0.017653
|  |  |  | 1499.9466 | 0.018162
|  |  |  | 1499.9469 | 0.018685
|  |  |  | 1499.9473 | 0.019222
|  |  |  | 1499.9476 | 0.019773
|  |  |  | 1499.9479 | 0.020338
|  |  |  | 1499.9482 | 0.020918
|  |  |  | 1499.9485 | 0.021514
|  |  |  | 1499.9489 | 0.022124
|  |  |  | 1499.9492 | 0.022751
|  |  |  | 1499.9495 | 0.023394
|  |  |  | 1499.9498 | 0.024054
|  |  |  | 1499.9501 | 0.024731
|  |  |  | 1499.9505 | 0.025425
|  |  |  | 1499.9508 | 0.026137
|  |  |  | 1499.9511 | 0.026867
|  |  |  | 1499.9514 | 0.027616
|  |  |  | 1499.9517 | 0.028383
|  |  |  | 1499.9521 | 0.029171
|  |  |  | 1499.9524 | 0.029978
|  |  |  | 1499.9527 | 0.030805
|  |  |  | 1499.953 | 0.031654
|  |  |  | 1499.9533 | 0.032524
|  |  |  | 1499.9537 | 0.033415
|  |  |  | 1499.954 | 0.034329
|  |  |  | 1499.9543 | 0.035265
|  |  |  | 1499.9546 | 0.036225
|  |  |  | 1499.9549 | 0.037208
|  |  |  | 1499.9553 | 0.038216
|  |  |  | 1499.9556 | 0.039248
|  |  |  | 1499.9559 | 0.040306
|  |  |  | 1499.9562 | 0.041389
|  |  |  | 1499.9565 | 0.042499
|  |  |  | 1499.9569 | 0.043636
|  |  |  | 1499.9572 | 0.044801
|  |  |  | 1499.9575 | 0.045993
|  |  |  | 1499.9578 | 0.047214
|  |  |  | 1499.9581 | 0.048465
|  |  |  | 1499.9585 | 0.049746
|  |  |  | 1499.9588 | 0.051057
|  |  |  | 1499.9591 | 0.052399
|  |  |  | 1499.9594 | 0.053773
|  |  |  | 1499.9597 | 0.05518
|  |  |  | 1499.9601 | 0.05662
|  |  |  | 1499.9604 | 0.058094
|  |  |  | 1499.9607 | 0.059602
|  |  |  | 1499.961 | 0.061146
|  |  |  | 1499.9613 | 0.062725
|  |  |  | 1499.9617 | 0.064341
|  |  |  | 1499.962 | 0.065995
|  |  |  | 1499.9623 | 0.067687
|  |  |  | 1499.9626 | 0.069418
|  |  |  | 1499.9629 | 0.071188
|  |  |  | 1499.9633 | 0.072999
|  |  |  | 1499.9636 | 0.074851
|  |  |  | 1499.9639 | 0.076745
|  |  |  | 1499.9642 | 0.078682
|  |  |  | 1499.9645 | 0.080663
|  |  |  | 1499.9649 | 0.082689
|  |  |  | 1499.9652 | 0.08476
|  |  |  | 1499.9655 | 0.086877
|  |  |  | 1499.9658 | 0.089041
|  |  |  | 1499.9661 | 0.091254
|  |  |  | 1499.9665 | 0.093515
|  |  |  | 1499.9668 | 0.095827
|  |  |  | 1499.9671 | 0.098189
|  |  |  | 1499.9674 | 0.1006
|  |  |  | 1499.9677 | 0.10307
|  |  |  | 1499.9681 | 0.10559
|  |  |  | 1499.9684 | 0.10817
|  |  |  | 1499.9687 | 0.1108
|  |  |  | 1499.969 | 0.11348
|  |  |  | 1499.9693 | 0.11623
|  |  |  | 1499.9697 | 0.11903
|  |  |  | 1499.97 | 0.1219
|  |  |  | 1499.9703 | 0.12482
|  |  |  | 1499.9706 | 0.12781
|  |  |  | 1499.9709 | 0.13086
|  |  |  | 1499.9713 | 0.13397
|  |  |  | 1499.9716 | 0.13715
|  |  |  | 1499.9719 | 0.1404
|  |  |  | 1499.9722 | 0.14371
|  |  |  | 1499.9725 | 0.1471
|  |  |  | 1499.9729 | 0.15055
|  |  |  | 1499.9732 | 0.15407
|  |  |  | 1499.9735 | 0.15767
|  |  |  | 1499.9738 | 0.16134
|  |  |  | 1499.9741 | 0.16508
|  |  |  | 1499.9745 | 0.1689
|  |  |  | 1499.9748 | 0.1728
|  |  |  | 1499.9751 | 0.17678
|  |  |  | 1499.9754 | 0.18083
|  |  |  | 1499.9757 | 0.18497
|  |  |  | 1499.9761 | 0.18919
|  |  |  | 1499.9764 | 0.1935
|  |  |  | 1499.9767 | 0.19789
|  |  |  | 1499.977 | 0.20236
|  |  |  | 1499.9773 | 0.20693
|  |  |  | 1499.9777 | 0.21158
|  |  |  | 1499.978 | 0.21632
|  |  |  | 1499.9783 | 0.22116
|  |  |  | 1499.9786 | 0.22609
|  |  |  | 1499.9789 | 0.23112
|  |  |  | 1499.9793 | 0.23624
|  |  |  | 1499.9796 | 0.24146
|  |  |  | 1499.9799 | 0.24678
|  |  |  | 1499.9802 | 0.2522
|  |  |  | 1499.9805 | 0.25772
|  |  |  | 1499.9809 | 0.26335
|  |  |  | 1499.9812 | 0.26908
|  |  |  | 1499.9815 | 0.27492
|  |  |  | 1499.9818 | 0.28087
|  |  |  | 1499.9821 | 0.28693
|  |  |  | 1499.9825 | 0.2931
|  |  |  | 1499.9828 | 0.29939
|  |  |  | 1499.9831 | 0.30579
|  |  |  | 1499.9834 | 0.31231
|  |  |  | 1499.9837 | 0.31894
|  |  |  | 1499.9841 | 0.3257
|  |  |  | 1499.9844 | 0.33257
|  |  |  | 1499.9847 | 0.33957
|  |  |  | 1499.985 | 0.3467
|  |  |  | 1499.9853 | 0.35395
|  |  |  | 1499.9857 | 0.36133
|  |  |  | 1499.986 | 0.36885
|  |  |  | 1499.9863 | 0.37649
|  |  |  | 1499.9866 | 0.38427
|  |  |  | 1499.9869 | 0.39218
|  |  |  | 1499.9873 | 0.40023
|  |  |  | 1499.9876 | 0.40842
|  |  |  | 1499.9879 | 0.41675
|  |  |  | 1499.9882 | 0.42523
|  |  |  | 1499.9885 | 0.43384
|  |  |  | 1499.9889 | 0.44261
|  |  |  | 1499.9892 | 0.45152
|  |  |  | 1499.9895 | 0.46058
|  |  |  | 1499.9898 | 0.4698
|  |  |  | 1499.9901 | 0.47916
|  |  |  | 1499.9905 | 0.48869
|  |  |  | 1499.9908 | 0.49837
|  |  |  | 1499.9911 | 0.50821
|  |  |  | 1499.9914 | 0.51821
|  |  |  | 1499.9917 | 0.52837
|  |  |  | 1499.9921 | 0.5387
|  |  |  | 1499.9924 | 0.5492
|  |  |  | 1499.9927 | 0.55986
|  |  |  | 1499.993 | 0.5707
|  |  |  | 1499.9933 | 0.58171
|  |  |  | 1499.9937 | 0.59289
|  |  |  | 1499.994 | 0.60425
|  |  |  | 1499.9943 | 0.61578
|  |  |  | 1499.9946 | 0.6275
|  |  |  | 1499.9949 | 0.6394
|  |  |  | 1499.9953 | 0.65148
|  |  |  | 1499.9956 | 0.66375
|  |  |  | 1499.9959 | 0.6762
|  |  |  | 1499.9962 | 0.68885
|  |  |  | 1499.9965 | 0.70168
|  |  |  | 1499.9969 | 0.71472
|  |  |  | 1499.9972 | 0.72794
|  |  |  | 1499.9975 | 0.74136
|  |  |  | 1499.9978 | 0.75499
|  |  |  | 1499.9981 | 0.76881
|  |  |  | 1499.9985 | 0.78284
|  |  |  | 1499.9988 | 0.79707
|  |  |  | 1499.9991 | 0.81151
|  |  |  | 1499.9994 | 0.82615
|  |  |  | 1499.9997 | 0.84101
|  |  |  | 1500.0001 | 0.85608
|  |  |  | 1500.0004 | 0.87136
|  |  |  | 1500.0007 | 0.88686
|  |  |  | 1500.001 | 0.90258
|  |  |  | 1500.0013 | 0.91852
|  |  |  | 1500.0017 | 0.93468
|  |  |  | 1500.002 | 0.95106
|  |  |  | 1500.0023 | 0.96767
|  |  |  | 1500.0026 | 0.98451
|  |  |  | 1500.0029 | 1.0016
|  |  |  | 1500.0033 | 1.0189
|  |  |  | 1500.0036 | 1.0364
|  |  |  | 1500.0039 | 1.0542
|  |  |  | 1500.0042 | 1.0721
|  |  |  | 1500.0045 | 1.0904
|  |  |  | 1500.0049 | 1.1089
|  |  |  | 1500.0052 | 1.1276
|  |  |  | 1500.0055 | 1.1465
|  |  |  | 1500.0058 | 1.1657
|  |  |  | 1500.0061 | 1.1852
|  |  |  | 1500.0065 | 1.2049
|  |  |  | 1500.0068 | 1.2248
|  |  |  | 1500.0071 | 1.245
|  |  |  | 1500.0074 | 1.2654
|  |  |  | 1500.0077 | 1.2861
|  |  |  | 1500.0081 | 1.3071
|  |  |  | 1500.0084 | 1.3283
|  |  |  | 1500.0087 | 1.3498
|  |  |  | 1500.009 | 1.3715
|  |  |  | 1500.0093 | 1.3935
|  |  |  | 1500.0097 | 1.4157
|  |  |  | 1500.01 | 1.4382
|  |  |  | 1500.0103 | 1.461
|  |  |  | 1500.0106 | 1.4841
|  |  |  | 1500.0109 | 1.5074
|  |  |  | 1500.0113 | 1.531
|  |  |  | 1500.0116 | 1.5548
|  |  |  | 1500.0119 | 1.5789
|  |  |  | 1500.0122 | 1.6033
|  |  |  | 1500.0125 | 1.628
|  |  |  | 1500.0129 | 1.6529
|  |  |  | 1500.0132 | 1.6781
|  |  |  | 1500.0135 | 1.7036
|  |  |  | 1500.0138 | 1.7293
|  |  |  | 1500.0141 | 1.7554
|  |  |  | 1500.0145 | 1.7817
|  |  |  | 1500.0148 | 1.8083
|  |  |  | 1500.0151 | 1.8352
|  |  |  | 1500.0154 | 1.8623
|  |  |  | 1500.0157 | 1.8898
|  |  |  | 1500.0161 | 1.9175
|  |  |  | 1500.0164 | 1.9455
|  |  |  | 1500.0167 | 1.9738
|  |  |  | 1500.017 | 2.0023
|  |  |  | 1500.0173 | 2.0312
|  |  |  | 1500.0177 | 2.0603
|  |  |  | 1500.018 | 2.0897
|  |  |  | 1500.0183 | 2.1194
|  |  |  | 1500.0186 | 2.1494
|  |  |  | 1500.0189 | 2.1797
|  |  |  | 1500.0193 | 2.2102
|  |  |  | 1500.0196 | 2.2411
|  |  |  | 1500.0199 | 2.2722
|  |  |  | 1500.0202 | 2.3036
|  |  |  | 1500.0205 | 2.3353
|  |  |  | 1500.0209 | 2.3673
|  |  |  | 1500.0212 | 2.3995
|  |  |  | 1500.0215 | 2.4321
|  |  |  | 1500.0218 | 2.4649
|  |  |  | 1500.0221 | 2.498
|  |  |  | 1500.0225 | 2.5314
|  |  |  | 1500.0228 | 2.5651
|  |  |  | 1500.0231 | 2.599
|  |  |  | 1500.0234 | 2.6333
|  |  |  | 1500.0237 | 2.6678
|  |  |  | 1500.0241 | 2.7026
|  |  |  | 1500.0244 | 2.7377
|  |  |  | 1500.0247 | 2.773
|  |  |  | 1500.025 | 2.8087
|  |  |  | 1500.0253 | 2.8446
|  |  |  | 1500.0257 | 2.8807
|  |  |  | 1500.026 | 2.9172
|  |  |  | 1500.0263 | 2.9539
|  |  |  | 1500.0266 | 2.9909
|  |  |  | 1500.0269 | 3.0282
|  |  |  | 1500.0273 | 3.0657
|  |  |  | 1500.0276 | 3.1035
|  |  |  | 1500.0279 | 3.1416
|  |  |  | 1500.0282 | 3.1799
|  |  |  | 1500.0285 | 3.2185
|  |  |  | 1500.0289 | 3.2574
|  |  |  | 1500.0292 | 3.2965
|  |  |  | 1500.0295 | 3.3359
|  |  |  | 1500.0298 | 3.3755
|  |  |  | 1500.0301 | 3.4154
|  |  |  | 1500.0305 | 3.4555
|  |  |  | 1500.0308 | 3.4959
|  |  |  | 1500.0311 | 3.5365
|  |  |  | 1500.0314 | 3.5774
|  |  |  | 1500.0317 | 3.6185
|  |  |  | 1500.0321 | 3.6598
|  |  |  | 1500.0324 | 3.7014
|  |  |  | 1500.0327 | 3.7432
|  |  |  | 1500.033 | 3.7852
|  |  |  | 1500.0333 | 3.8275
|  |  |  | 1500.0337 | 3.87
|  |  |  | 1500.034 | 3.9127
|  |  |  | 1500.0343 | 3.9556
|  |  |  | 1500.0346 | 3.9988
|  |  |  | 1500.0349 | 4.0421
|  |  |  | 1500.0353 | 4.0857
|  |  |  | 1500.0356 | 4.1295
|  |  |  | 1500.0359 | 4.1734
|  |  |  | 1500.0362 | 4.2176
|  |  |  | 1500.0365 | 4.262
|  |  |  | 1500.0369 | 4.3065
|  |  |  | 1500.0372 | 4.3513
|  |  |  | 1500.0375 | 4.3962
|  |  |  | 1500.0378 | 4.4413
|  |  |  | 1500.0381 | 4.4866
|  |  |  | 1500.0385 | 4.532
|  |  |  | 1500.0388 | 4.5776
|  |  |  | 1500.0391 | 4.6234
|  |  |  | 1500.0394 | 4.6694
|  |  |  | 1500.0397 | 4.7155
|  |  |  | 1500.0401 | 4.7617
|  |  |  | 1500.0404 | 4.8081
|  |  |  | 1500.0407 | 4.8547
|  |  |  | 1500.041 | 4.9013
|  |  |  | 1500.0413 | 4.9481
|  |  |  | 1500.0417 | 4.9951
|  |  |  | 1500.042 | 5.0421
|  |  |  | 1500.0423 | 5.0893
|  |  |  | 1500.0426 | 5.1366
|  |  |  | 1500.0429 | 5.184
|  |  |  | 1500.0433 | 5.2315
|  |  |  | 1500.0436 | 5.279
|  |  |  | 1500.0439 | 5.3267
|  |  |  | 1500.0442 | 5.3745
|  |  |  | 1500.0445 | 5.4223
|  |  |  | 1500.0449 | 5.4703
|  |  |  | 1500.0452 | 5.5183
|  |  |  | 1500.0455 | 5.5663
|  |  |  | 1500.0458 | 5.6145
|  |  |  | 1500.0461 | 5.6626
|  |  |  | 1500.0465 | 5.7109
|  |  |  | 1500.0468 | 5.7591
|  |  |  | 1500.0471 | 5.8074
|  |  |  | 1500.0474 | 5.8558
|  |  |  | 1500.0477 | 5.9041
|  |  |  | 1500.0481 | 5.9525
|  |  |  | 1500.0484 | 6.0009
|  |  |  | 1500.0487 | 6.0493
|  |  |  | 1500.049 | 6.0977
|  |  |  | 1500.0493 | 6.146
|  |  |  | 1500.0497 | 6.1944
|  |  |  | 1500.05 | 6.2428
|  |  |  | 1500.0503 | 6.2911
|  |  |  | 1500.0506 | 6.3394
|  |  |  | 1500.0509 | 6.3877
|  |  |  | 1500.0513 | 6.4359
|  |  |  | 1500.0516 | 6.484
|  |  |  | 1500.0519 | 6.5322
|  |  |  | 1500.0522 | 6.5802
|  |  |  | 1500.0525 | 6.6282
|  |  |  | 1500.0529 | 6.6761
|  |  |  | 1500.0532 | 6.7239
|  |  |  | 1500.0535 | 6.7716
|  |  |  | 1500.0538 | 6.8192
|  |  |  | 1500.0541 | 6.8667
|  |  |  | 1500.0545 | 6.9141
|  |  |  | 1500.0548 | 6.9614
|  |  |  | 1500.0551 | 7.0086
|  |  |  | 1500.0554 | 7.0556
|  |  |  | 1500.0557 | 7.1025
|  |  |  | 1500.0561 | 7.1493
|  |  |  | 1500.0564 | 7.1959
|  |  |  | 1500.0567 | 7.2423
|  |  |  | 1500.057 | 7.2886
|  |  |  | 1500.0573 | 7.3346
|  |  |  | 1500.0577 | 7.3806
|  |  |  | 1500.058 | 7.4263
|  |  |  | 1500.0583 | 7.4718
|  |  |  | 1500.0586 | 7.5171
|  |  |  | 1500.0589 | 7.5623
|  |  |  | 1500.0593 | 7.6072
|  |  |  | 1500.0596 | 7.6518
|  |  |  | 1500.0599 | 7.6963
|  |  |  | 1500.0602 | 7.7405
|  |  |  | 1500.0605 | 7.7845
|  |  |  | 1500.0609 | 7.8282
|  |  |  | 1500.0612 | 7.8716
|  |  |  | 1500.0615 | 7.9148
|  |  |  | 1500.0618 | 7.9577
|  |  |  | 1500.0621 | 8.0004
|  |  |  | 1500.0625 | 8.0427
|  |  |  | 1500.0628 | 8.0848
|  |  |  | 1500.0631 | 8.1266
|  |  |  | 1500.0634 | 8.168
|  |  |  | 1500.0637 | 8.2091
|  |  |  | 1500.0641 | 8.25
|  |  |  | 1500.0644 | 8.2905
|  |  |  | 1500.0647 | 8.3306
|  |  |  | 1500.065 | 8.3704
|  |  |  | 1500.0653 | 8.4099
|  |  |  | 1500.0657 | 8.449
|  |  |  | 1500.066 | 8.4878
|  |  |  | 1500.0663 | 8.5261
|  |  |  | 1500.0666 | 8.5642
|  |  |  | 1500.0669 | 8.6018
|  |  |  | 1500.0673 | 8.639
|  |  |  | 1500.0676 | 8.6759
|  |  |  | 1500.0679 | 8.7123
|  |  |  | 1500.0682 | 8.7484
|  |  |  | 1500.0685 | 8.784
|  |  |  | 1500.0689 | 8.8192
|  |  |  | 1500.0692 | 8.854
|  |  |  | 1500.0695 | 8.8883
|  |  |  | 1500.0698 | 8.9222
|  |  |  | 1500.0701 | 8.9557
|  |  |  | 1500.0705 | 8.9887
|  |  |  | 1500.0708 | 9.0213
|  |  |  | 1500.0711 | 9.0534
|  |  |  | 1500.0714 | 9.085
|  |  |  | 1500.0717 | 9.1162
|  |  |  | 1500.0721 | 9.1468
|  |  |  | 1500.0724 | 9.177
|  |  |  | 1500.0727 | 9.2068
|  |  |  | 1500.073 | 9.236
|  |  |  | 1500.0733 | 9.2647
|  |  |  | 1500.0737 | 9.2929
|  |  |  | 1500.074 | 9.3206
|  |  |  | 1500.0743 | 9.3478
|  |  |  | 1500.0746 | 9.3744
|  |  |  | 1500.0749 | 9.4006
|  |  |  | 1500.0753 | 9.4262
|  |  |  | 1500.0756 | 9.4512
|  |  |  | 1500.0759 | 9.4758
|  |  |  | 1500.0762 | 9.4997
|  |  |  | 1500.0765 | 9.5232
|  |  |  | 1500.0769 | 9.5461
|  |  |  | 1500.0772 | 9.5684
|  |  |  | 1500.0775 | 9.5902
|  |  |  | 1500.0778 | 9.6114
|  |  |  | 1500.0781 | 9.632
|  |  |  | 1500.0785 | 9.652
|  |  |  | 1500.0788 | 9.6715
|  |  |  | 1500.0791 | 9.6904
|  |  |  | 1500.0794 | 9.7087
|  |  |  | 1500.0797 | 9.7265
|  |  |  | 1500.0801 | 9.7436
|  |  |  | 1500.0804 | 9.7601
|  |  |  | 1500.0807 | 9.7761
|  |  |  | 1500.081 | 9.7914
|  |  |  | 1500.0813 | 9.8061
|  |  |  | 1500.0817 | 9.8203
|  |  |  | 1500.082 | 9.8338
|  |  |  | 1500.0823 | 9.8467
|  |  |  | 1500.0826 | 9.859
|  |  |  | 1500.0829 | 9.8707
|  |  |  | 1500.0833 | 9.8817
|  |  |  | 1500.0836 | 9.8922
|  |  |  | 1500.0839 | 9.902
|  |  |  | 1500.0842 | 9.9112
|  |  |  | 1500.0845 | 9.9198
|  |  |  | 1500.0849 | 9.9277
|  |  |  | 1500.0852 | 9.935
|  |  |  | 1500.0855 | 9.9417
|  |  |  | 1500.0858 | 9.9477
|  |  |  | 1500.0861 | 9.9532
|  |  |  | 1500.0865 | 9.9579
|  |  |  | 1500.0868 | 9.9621
|  |  |  | 1500.0871 | 9.9656
|  |  |  | 1500.0874 | 9.9685
|  |  |  | 1500.0877 | 9.9707
|  |  |  | 1500.0881 | 9.9723
|  |  |  | 1500.0884 | 9.9732
|  |  |  | 1500.0887 | 9.9736
|  |  |  | 1500.089 | 9.9732
|  |  |  | 1500.0893 | 9.9723
|  |  |  | 1500.0897 | 9.9707
|  |  |  | 1500.09 | 9.9685
|  |  |  | 1500.0903 | 9.9656
|  |  |  | 1500.0906 | 9.9621
|  |  |  | 1500.0909 | 9.9579
|  |  |  | 1500.0913 | 9.9532
|  |  |  | 1500.0916 | 9.9477
|  |  |  | 1500.0919 | 9.9417
|  |  |  | 1500.0922 | 9.935
|  |  |  | 1500.0925 | 9.9277
|  |  |  | 1500.0929 | 9.9198
|  |  |  | 1500.0932 | 9.9112
|  |  |  | 1500.0935 | 9.902
|  |  |  | 1500.0938 | 9.8922
|  |  |  | 1500.0941 | 9.8817
|  |  |  | 1500.0945 | 9.8707
|  |  |  | 1500.0948 | 9.859
|  |  |  | 1500.0951 | 9.8467
|  |  |  | 1500.0954 | 9.8338
|  |  |  | 1500.0957 | 9.8203
|  |  |  | 1500.0961 | 9.8061
|  |  |  | 1500.0964 | 9.7914
|  |  |  | 1500.0967 | 9.7761
|  |  |  | 1500.097 | 9.7601
|  |  |  | 1500.0973 | 9.7436
|  |  |  | 1500.0977 | 9.7265
|  |  |  | 1500.098 | 9.7087
|  |  |  | 1500.0983 | 9.6904
|  |  |  | 1500.0986 | 9.6715
|  |  |  | 1500.0989 | 9.652
|  |  |  | 1500.0993 | 9.632
|  |  |  | 1500.0996 | 9.6114
|  |  |  | 1500.0999 | 9.5902
|  |  |  | 1500.1002 | 9.5684
|  |  |  | 1500.1005 | 9.5461
|  |  |  | 1500.1009 | 9.5232
|  |  |  | 1500.1012 | 9.4997
|  |  |  | 1500.1015 | 9.4758
|  |  |  | 1500.1018 | 9.4512
|  |  |  | 1500.1021 | 9.4262
|  |  |  | 1500.1025 | 9.4006
|  |  |  | 1500.1028 | 9.3744
|  |  |  | 1500.1031 | 9.3478
|  |  |  | 1500.1034 | 9.3206
|  |  |  | 1500.1037 | 9.2929
|  |  |  | 1500.1041 | 9.2647
|  |  |  | 1500.1044 | 9.236
|  |  |  | 1500.1047 | 9.2068
|  |  |  | 1500.105 | 9.177
|  |  |  | 1500.1053 | 9.1468
|  |  |  | 1500.1057 | 9.1162
|  |  |  | 1500.106 | 9.085
|  |  |  | 1500.1063 | 9.0534
|  |  |  | 1500.1066 | 9.0213
|  |  |  | 1500.1069 | 8.9887
|  |  |  | 1500.1073 | 8.9557
|  |  |  | 1500.1076 | 8.9222
|  |  |  | 1500.1079 | 8.8883
|  |  |  | 1500.1082 | 8.854
|  |  |  | 1500.1085 | 8.8192
|  |  |  | 1500.1089 | 8.784
|  |  |  | 1500.1092 | 8.7484
|  |  |  | 1500.1095 | 8.7123
|  |  |  | 1500.1098 | 8.6759
|  |  |  | 1500.1101 | 8.639
|  |  |  | 1500.1105 | 8.6018
|  |  |  | 1500.1108 | 8.5642
|  |  |  | 1500.1111 | 8.5261
|  |  |  | 1500.1114 | 8.4878
|  |  |  | 1500.1117 | 8.449
|  |  |  | 1500.1121 | 8.4099
|  |  |  | 1500.1124 | 8.3704
|  |  |  | 1500.1127 | 8.3306
|  |  |  | 1500.113 | 8.2905
|  |  |  | 1500.1133 | 8.25
|  |  |  | 1500.1137 | 8.2091
|  |  |  | 1500.114 | 8.168
|  |  |  | 1500.1143 | 8.1266
|  |  |  | 1500.1146 | 8.0848
|  |  |  | 1500.1149 | 8.0427
|  |  |  | 1500.1153 | 8.0004
|  |  |  | 1500.1156 | 7.9577
|  |  |  | 1500.1159 | 7.9148
|  |  |  | 1500.1162 | 7.8716
|  |  |  | 1500.1165 | 7.8282
|  |  |  | 1500.1169 | 7.7845
|  |  |  | 1500.1172 | 7.7405
|  |  |  | 1500.1175 | 7.6963
|  |  |  | 1500.1178 | 7.6518
|  |  |  | 1500.1181 | 7.6072
|  |  |  | 1500.1185 | 7.5623
|  |  |  | 1500.1188 | 7.5171
|  |  |  | 1500.1191 | 7.4718
|  |  |  | 1500.1194 | 7.4263
|  |  |  | 1500.1197 | 7.3806
|  |  |  | 1500.1201 | 7.3346
|  |  |  | 1500.1204 | 7.2886
|  |  |  | 1500.1207 | 7.2423
|  |  |  | 1500.121 | 7.1959
|  |  |  | 1500.1213 | 7.1493
|  |  |  | 1500.1217 | 7.1025
|  |  |  | 1500.122 | 7.0556
|  |  |  | 1500.1223 | 7.0086
|  |  |  | 1500.1226 | 6.9614
|  |  |  | 1500.1229 | 6.9141
|  |  |  | 1500.1233 | 6.8667
|  |  |  | 1500.1236 | 6.8192
|  |  |  | 1500.1239 | 6.7716
|  |  |  | 1500.1242 | 6.7239
|  |  |  | 1500.1245 | 6.6761
|  |  |  | 1500.1249 | 6.6282
|  |  |  | 1500.1252 | 6.5802
|  |  |  | 1500.1255 | 6.5322
|  |  |  | 1500.1258 | 6.484
|  |  |  | 1500.1261 | 6.4359
|  |  |  | 1500.1265 | 6.3877
|  |  |  | 1500.1268 | 6.3394
|  |  |  | 1500.1271 | 6.2911
|  |  |  | 1500.1274 | 6.2428
|  |  |  | 1500.1277 | 6.1944
|  |  |  | 1500.1281 | 6.146
|  |  |  | 1500.1284 | 6.0977
|  |  |  | 1500.1287 | 6.0493
|  |  |  | 1500.129 | 6.0009
|  |  |  | 1500.1293 | 5.9525
|  |  |  | 1500.1297 | 5.9041
|  |  |  | 1500.13 | 5.8558
|  |  |  | 1500.1303 | 5.8074
|  |  |  | 1500.1306 | 5.7591
|  |  |  | 1500.1309 | 5.7109
|  |  |  | 1500.1313 | 5.6626
|  |  |  | 1500.1316 | 5.6145
|  |  |  | 1500.1319 | 5.5663
|  |  |  | 1500.1322 | 5.5183
|  |  |  | 1500.1325 | 5.4703
|  |  |  | 1500.1329 | 5.4223
|  |  |  | 1500.1332 | 5.3745
|  |  |  | 1500.1335 | 5.3267
|  |  |  | 1500.1338 | 5.279
|  |  |  | 1500.1341 | 5.2315
|  |  |  | 1500.1345 | 5.184
|  |  |  | 1500.1348 | 5.1366
|  |  |  | 1500.1351 | 5.0893
|  |  |  | 1500.1354 | 5.0421
|  |  |  | 1500.1357 | 4.9951
|  |  |  | 1500.1361 | 4.9481
|  |  |  | 1500.1364 | 4.9013
|  |  |  | 1500.1367 | 4.8547
|  |  |  | 1500.137 | 4.8081
|  |  |  | 1500.1373 | 4.7617
|  |  |  | 1500.1377 | 4.7155
|  |  |  | 1500.138 | 4.6694
|  |  |  | 1500.1383 | 4.6234
|  |  |  | 1500.1386 | 4.5776
|  |  |  | 1500.1389 | 4.532
|  |  |  | 1500.1393 | 4.4866
|  |  |  | 1500.1396 | 4.4413
|  |  |  | 1500.1399 | 4.3962
|  |  |  | 1500.1402 | 4.3513
|  |  |  | 1500.1405 | 4.3065
|  |  |  | 1500.1409 | 4.262
|  |  |  | 1500.1412 | 4.2176
|  |  |  | 1500.1415 | 4.1734
|  |  |  | 1500.1418 | 4.1295
|  |  |  | 1500.1421 | 4.0857
|  |  |  | 1500.1425 | 4.0421
|  |  |  | 1500.1428 | 3.9988
|  |  |  | 1500.1431 | 3.9556
|  |  |  | 1500.1434 | 3.9127
|  |  |  | 1500.1437 | 3.87
|  |  |  | 1500.1441 | 3.8275
|  |  |  | 1500.1444 | 3.7852
|  |  |  | 1500.1447 | 3.7432
|  |  |  | 1500.145 | 3.7014
|  |  |  | 1500.1453 | 3.6598
|  |  |  | 1500.1457 | 3.6185
|  |  |  | 1500.146 | 3.5774
|  |  |  | 1500.1463 | 3.5365
|  |  |  | 1500.1466 | 3.4959
|  |  |  | 1500.1469 | 3.4555
|  |  |  | 1500.1473 | 3.4154
|  |  |  | 1500.1476 | 3.3755
|  |  |  | 1500.1479 | 3.3359
|  |  |  | 1500.1482 | 3.2965
|  |  |  | 1500.1485 | 3.2574
|  |  |  | 1500.1489 | 3.2185
|  |  |  | 1500.1492 | 3.1799
|  |  |  | 1500.1495 | 3.1416
|  |  |  | 1500.1498 | 3.1035
|  |  |  | 1500.1501 | 3.0657
|  |  |  | 1500.1505 | 3.0282
|  |  |  | 1500.1508 | 2.9909
|  |  |  | 1500.1511 | 2.9539
|  |  |  | 1500.1514 | 2.9172
|  |  |  | 1500.1517 | 2.8807
|  |  |  | 1500.1521 | 2.8446
|  |  |  | 1500.1524 | 2.8087
|  |  |  | 1500.1527 | 2.773
|  |  |  | 1500.153 | 2.7377
|  |  |  | 1500.1533 | 2.7026
|  |  |  | 1500.1537 | 2.6678
|  |  |  | 1500.154 | 2.6333
|  |  |  | 1500.1543 | 2.599
|  |  |  | 1500.1546 | 2.5651
|  |  |  | 1500.1549 | 2.5314
|  |  |  | 1500.1553 | 2.498
|  |  |  | 1500.1556 | 2.4649
|  |  |  | 1500.1559 | 2.4321
|  |  |  | 1500.1562 | 2.3995
|  |  |  | 1500.1565 | 2.3673
|  |  |  | 1500.1569 | 2.3353
|  |  |  | 1500.1572 | 2.3036
|  |  |  | 1500.1575 | 2.2722
|  |  |  | 1500.1578 | 2.2411
|  |  |  | 1500.1581 | 2.2102
|  |  |  | 1500.1585 | 2.1797
|  |  |  | 1500.1588 | 2.1494
|  |  |  | 1500.1591 | 2.1194
|  |  |  | 1500.1594 | 2.0897
|  |  |  | 1500.1597 | 2.0603
|  |  |  | 1500.1601 | 2.0312
|  |  |  | 1500.1604 | 2.0023
|  |  |  | 1500.1607 | 1.9738
|  |  |  | 1500.161 | 1.9455
|  |  |  | 1500.1613 | 1.9175
|  |  |  | 1500.1617 | 1.8898
|  |  |  | 1500.162 | 1.8623
|  |  |  | 1500.1623 | 1.8352
|  |  |  | 1500.1626 | 1.8083
|  |  |  | 1500.1629 | 1.7817
|  |  |  | 1500.1633 | 1.7554
|  |  |  | 1500.1636 | 1.7293
|  |  |  | 1500.1639 | 1.7036
|  |  |  | 1500.1642 | 1.6781
|  |  |  | 1500.1645 | 1.6529
|  |  |  | 1500.1649 | 1.628
|  |  |  | 1500.1652 | 1.6033
|  |  |  | 1500.1655 | 1.5789
|  |  |  | 1500.1658 | 1.5548
|  |  |  | 1500.1661 | 1.531
|  |  |  | 1500.1665 | 1.5074
|  |  |  | 1500.1668 | 1.4841
|  |  |  | 1500.1671 | 1.461
|  |  |  | 1500.1674 | 1.4382
|  |  |  | 1500.1677 | 1.4157
|  |  |  | 1500.1681 | 1.3935
|  |  |  | 1500.1684 | 1.3715
|  |  |  | 1500.1687 | 1.3498
|  |  |  | 1500.169 | 1.3283
|  |  |  | 1500.1693 | 1.3071
|  |  |  | 1500.1697 | 1.2861
|  |  |  | 1500.17 | 1.2654
|  |  |  | 1500.1703 | 1.245
|  |  |  | 1500.1706 | 1.2248
|  |  |  | 1500.1709 | 1.2049
|  |  |  | 1500.1713 | 1.1852
|  |  |  | 1500.1716 | 1.1657
|  |  |  | 1500.1719 | 1.1465
|  |  |  | 1500.1722 | 1.1276
|  |  |  | 1500.1725 | 1.1089
|  |  |  | 1500.1729 | 1.0904
|  |  |  | 1500.1732 | 1.0721
|  |  |  | 1500.1735 | 1.0542
|  |  |  | 1500.1738 | 1.0364
|  |  |  | 1500.1741 | 1.0189
|  |  |  | 1500.1745 | 1.0016
|  |  |  | 1500.1748 | 0.98451
|  |  |  | 1500.1751 | 0.96767
|  |  |  | 1500.1754 | 0.95106
|  |  |  | 1500.1757 | 0.93468
|  |  |  | 1500.1761 | 0.91852
|  |  |  | 1500.1764 | 0.90258
|  |  |  | 1500.1767 | 0.88686
|  |  |  | 1500.177 | 0.87136
|  |  |  | 1500.1773 | 0.85608
|  |  |  | 1500.1777 | 0.84101
|  |  |  | 1500.178 | 0.82615
|  |  |  | 1500.1783 | 0.81151
|  |  |  | 1500.1786 | 0.79707
|  |  |  | 1500.1789 | 0.78284
|  |  |  | 1500.1793 | 0.76881
|  |  |  | 1500.1796 | 0.75499
|  |  |  | 1500.1799 | 0.74136
|  |  |  | 1500.1802 | 0.72794
|  |  |  | 1500.1805 | 0.71472
|  |  |  | 1500.1809 | 0.70168
|  |  |  | 1500.1812 | 0.68885
|  |  |  | 1500.1815 | 0.6762
|  |  |  | 1500.1818 | 0.66375
|  |  |  | 1500.1821 | 0.65148
|  |  |  | 1500.1825 | 0.6394
|  |  |  | 1500.1828 | 0.6275
|  |  |  | 1500.1831 | 0.61578
|  |  |  | 1500.1834 | 0.60425
|  |  |  | 1500.1837 | 0.59289
|  |  |  | 1500.1841 | 0.58171
|  |  |  | 1500.1844 | 0.5707
|  |  |  | 1500.1847 | 0.55986
|  |  |  | 1500.185 | 0.5492
|  |  |  | 1500.1853 | 0.5387
|  |  |  | 1500.1857 | 0.52837
|  |  |  | 1500.186 | 0.51821
|  |  |  | 1500.1863 | 0.50821
|  |  |  | 1500.1866 | 0.49837
|  |  |  | 1500.1869 | 0.48869
|  |  |  | 1500.1873 | 0.47916
|  |  |  | 1500.1876 | 0.4698
|  |  |  | 1500.1879 | 0.46058
|  |  |  | 1500.1882 | 0.45152
|  |  |  | 1500.1885 | 0.44261
|  |  |  | 1500.1889 | 0.43384
|  |  |  | 1500.1892 | 0.42523
|  |  |  | 1500.1895 | 0.41675
|  |  |  | 1500.1898 | 0.40842
|  |  |  | 1500.1901 | 0.40023
|  |  |  | 1500.1905 | 0.39218
|  |  |  | 1500.1908 | 0.38427
|  |  |  | 1500.1911 | 0.37649
|  |  |  | 1500.1914 | 0.36885
|  |  |  | 1500.1917 | 0.36133
|  |  |  | 1500.1921 | 0.35395
|  |  |  | 1500.1924 | 0.3467
|  |  |  | 1500.1927 | 0.33957
|  |  |  | 1500.193 | 0.33257
|  |  |  | 1500.1933 | 0.3257
|  |  |  | 1500.1937 | 0.31894
|  |  |  | 1500.194 | 0.31231
|  |  |  | 1500.1943 | 0.30579
|  |  |  | 1500.1946 | 0.29939
|  |  |  | 1500.1949 | 0.2931
|  |  |  | 1500.1953 | 0.28693
|  |  |  | 1500.1956 | 0.28087
|  |  |  | 1500.1959 | 0.27492
|  |  |  | 1500.1962 | 0.26908
|  |  |  | 1500.1965 | 0.26335
|  |  |  | 1500.1969 | 0.25772
|  |  |  | 1500.1972 | 0.2522
|  |  |  | 1500.1975 | 0.24678
|  |  |  | 1500.1978 | 0.24146
|  |  |  | 1500.1981 | 0.23624
|  |  |  | 1500.1985 | 0.23112
|  |  |  | 1500.1988 | 0.22609
|  |  |  | 1500.1991 | 0.22116
|  |  |  | 1500.1994 | 0.21632
|  |  |  | 1500.1997 | 0.21158
|  |  |  | 1500.2001 | 0.20693
|  |  |  | 1500.2004 | 0.20236
|  |  |  | 1500.2007 | 0.19789
|  |  |  | 1500.201 | 0.1935
|  |  |  | 1500.2013 | 0.18919
|  |  |  | 1500.2017 | 0.18497
|  |  |  | 1500.202 | 0.18083
|  |  |  | 1500.2023 | 0.17678
|  |  |  | 1500.2026 | 0.1728
|  |  |  | 1500.2029 | 0.1689
|  |  |  | 1500.2033 | 0.16508
|  |  |  | 1500.2036 | 0.16134
|  |  |  | 1500.2039 | 0.15767
|  |  |  | 1500.2042 | 0.15407
|  |  |  | 1500.2045 | 0.15055
|  |  |  | 1500.2049 | 0.1471
|  |  |  | 1500.2052 | 0.14371
|  |  |  | 1500.2055 | 0.1404
|  |  |  | 1500.2058 | 0.13715
|  |  |  | 1500.2061 | 0.13397
|  |  |  | 1500.2065 | 0.13086
|  |  |  | 1500.2068 | 0.12781
|  |  |  | 1500.2071 | 0.12482
|  |  |  | 1500.2074 | 0.1219
|  |  |  | 1500.2077 | 0.11903
|  |  |  | 1500.2081 | 0.11623
|  |  |  | 1500.2084 | 0.11348
|  |  |  | 1500.2087 | 0.1108
|  |  |  | 1500.209 | 0.10817
|  |  |  | 1500.2093 | 0.10559
|  |  |  | 1500.2097 | 0.10307
|  |  |  | 1500.21 | 0.1006
|  |  |  | 1500.2103 | 0.098189
|  |  |  | 1500.2106 | 0.095827
|  |  |  | 1500.2109 | 0.093515
|  |  |  | 1500.2113 | 0.091254
|  |  |  | 1500.2116 | 0.089041
|  |  |  | 1500.2119 | 0.086877
|  |  |  | 1500.2122 | 0.08476
|  |  |  | 1500.2125 | 0.082689
|  |  |  | 1500.2129 | 0.080663
|  |  |  | 1500.2132 | 0.078682
|  |  |  | 1500.2135 | 0.076745
|  |  |  | 1500.2138 | 0.074851
|  |  |  | 1500.2141 | 0.072999
|  |  |  | 1500.2145 | 0.071188
|  |  |  | 1500.2148 | 0.069418
|  |  |  | 1500.2151 | 0.067687
|  |  |  | 1500.2154 | 0.065995
|  |  |  | 1500.2157 | 0.064341
|  |  |  | 1500.2161 | 0.062725
|  |  |  | 1500.2164 | 0.061146
|  |  |  | 1500.2167 | 0.059602
|  |  |  | 1500.217 | 0.058094
|  |  |  | 1500.2173 | 0.05662
|  |  |  | 1500.2177 | 0.05518
|  |  |  | 1500.218 | 0.053773
|  |  |  | 1500.2183 | 0.052399
|  |  |  | 1500.2186 | 0.051057
|  |  |  | 1500.2189 | 0.049746
|  |  |  | 1500.2193 | 0.048465
|  |  |  | 1500.2196 | 0.047214
|  |  |  | 1500.2199 | 0.045993
|  |  |  | 1500.2202 | 0.044801
|  |  |  | 1500.2205 | 0.043636
|  |  |  | 1500.2209 | 0.042499
|  |  |  | 1500.2212 | 0.041389
|  |  |  | 1500.2215 | 0.040306
|  |  |  | 1500.2218 | 0.039248
|  |  |  | 1500.2221 | 0.038216
|  |  |  | 1500.2225 | 0.037208
|  |  |  | 1500.2228 | 0.036225
|  |  |  | 1500.2231 | 0.035265
|  |  |  | 1500.2234 | 0.034329
|  |  |  | 1500.2237 | 0.033415
|  |  |  | 1500.2241 | 0.032524
|  |  |  | 1500.2244 | 0.031654
|  |  |  | 1500.2247 | 0.030805
|  |  |  | 1500.225 | 0.029978
|  |  |  | 1500.2253 | 0.029171
|  |  |  | 1500.2257 | 0.028383
|  |  |  | 1500.226 | 0.027616
|  |  |  | 1500.2263 | 0.026867
|  |  |  | 1500.2266 | 0.026137
|  |  |  | 1500.2269 | 0.025425
|  |  |  | 1500.2273 | 0.024731
|  |  |  | 1500.2276 | 0.024054
|  |  |  | 1500.2279 | 0.023394
|  |  |  | 1500.2282 | 0.022751
|  |  |  | 1500.2285 | 0.022124
|  |  |  | 1500.2289 | 0.021514
|  |  |  | 1500.2292 | 0.020918
|  |  |  | 1500.2295 | 0.020338
|  |  |  | 1500.2298 | 0.019773
|  |  |  | 1500.2301 | 0.019222
|  |  |  | 1500.2305 | 0.018685
|  |  |  | 1500.2308 | 0.018162
|  |  |  | 1500.2311 | 0.017653
|  |  |  | 1500.2314 | 0.017157
|  |  |  | 1500.2317 | 0.016673
|  |  |  | 1500.2321 | 0.016202
|  |  |  | 1500.2324 | 0.015744
|  |  |  | 1500.2327 | 0.015298
|  |  |  | 1500.233 | 0.014863
|  |  |  | 1500.2333 | 0.014439
|  |  |  | 1500.2337 | 0.014027
|  |  |  | 1500.234 | 0.013626
|  |  |  | 1500.2343 | 0.013235
|  |  |  | 1500.2346 | 0.012855
|  |  |  | 1500.2349 | 0.012485
|  |  |  | 1500.2353 | 0.012125
|  |  |  | 1500.2356 | 0.011774
|  |  |  | 1500.2359 | 0.011433
|  |  |  | 1500.2362 | 0.011101
|  |  |  | 1500.2365 | 0.010778
|  |  |  | 1500.2369 | 0.010463
|  |  |  | 1500.2372 | 0.010158
|  |  |  | 1500.2375 | 0.0098601
|  |  |  | 1500.2378 | 0.0095706
|  |  |  | 1500.2381 | 0.0092891
|  |  |  | 1500.2385 | 0.0090153
|  |  |  | 1500.2388 | 0.008749
|  |  |  | 1500.2391 | 0.00849
|  |  |  | 1500.2394 | 0.0082382
|  |  |  | 1500.2397 | 0.0079933
|  |  |  | 1500.2401 | 0.0077552
|  |  |  | 1500.2404 | 0.0075237
|  |  |  | 1500.2407 | 0.0072987
|  |  |  | 1500.241 | 0.0070799
|  |  |  | 1500.2413 | 0.0068673
|  |  |  | 1500.2417 | 0.0066606
|  |  |  | 1500.242 | 0.0064597
|  |  |  | 1500.2423 | 0.0062645
|  |  |  | 1500.2426 | 0.0060747
|  |  |  | 1500.2429 | 0.0058904
|  |  |  | 1500.2433 | 0.0057113
|  |  |  | 1500.2436 | 0.0055373
|  |  |  | 1500.2439 | 0.0053682
|  |  |  | 1500.2442 | 0.005204
|  |  |  | 1500.2445 | 0.0050444
|  |  |  | 1500.2449 | 0.0048895
|  |  |  | 1500.2452 | 0.004739
|  |  |  | 1500.2455 | 0.0045928
|  |  |  | 1500.2458 | 0.0044509
|  |  |  | 1500.2461 | 0.004313
|  |  |  | 1500.2465 | 0.0041792
|  |  |  | 1500.2468 | 0.0040493
|  |  |  | 1500.2471 | 0.0039231
|  |  |  | 1500.2474 | 0.0038007
|  |  |  | 1500.2477 | 0.0036818
|  |  |  | 1500.2481 | 0.0035664
|  |  |  | 1500.2484 | 0.0034544

## Guarded Rejection
Guarded Rejection |  | Probability Table (single sided)
Probability (%) | 99.9% | Probability | Z-Value
Nominal Value | 75 | 0.5 | 0
Measurement Unc um | 0.02 | 0.7 | 0.5244
99.9 % Acceptence Threshold | 79.9407 | 0.75 | 0.67449
Probability of Making a Wrong Decision | 0.1% | 0.85 | 1.0364
|  | 0.9 | 1.2816
|  | 0.95 | 1.6449
|  | 0.9545 | 1.6901
|  | 0.97725 | 2
|  | 0.98 | 2.0537
|  | 0.99 | 2.3263
|  | 0.999 | 3.0902
|  | 1 | 5.9978
3.0902

## Method 5 & 6
PFA CALCULATION FOR AN ENTIRE LOADING RANGE METHOD 5
Instrument Capacity | 10000
Resolution UUT | 0.1
Tol UUT | 0.05% | Reading
Plus 1 Count | no
CMC | 0.02%
k = for CMC | 2
[7 empty] | Manufacturer's | [8 empty] | Acceptance
Force or Torque Applied | Instrument Reading 1 | Instrument Reading 2 | Instrument Reading 3 | Avg Instrument Reading | Req Tolerance | Plus 1 Count | LSL | USL | Res UUT | Rep UUT | Std Unc | TUR | Lower Risk | Upper Risk | Total Risk | Limit |  | LSL | USL
1000 | 1000 | 1000.1 | 1000 | 1000.0333 | 0.5 | No | 999.5 | 1000.5 | 0.1 | 0.057735 | 0.11902 | 2.1004 | 0.0003717% | 0.004413% | 0.004785% | 0.25555 |  | 999.7444 | 1000.2556 |  |  | CMC | 0% |  | Tol | 0.005% | yes | 1 | Reading | Max Force | 10000
2000 | 2000 | 2000 | 2000.1 | 2000.0333 | 1 | No | 1999 | 2001 | 0.1 | 0.057735 | 0.21016 | 2.3792 | 4.396e-05% | 0.0002116% | 0.0002555% | 0.56839 |  | 1999.4316 | 2000.5684 |  |  |  | 0.0016% |  | of applied | 0.01% | no | 2 | Full Scale
3000 | 3000 | 3000 | 3000 | 3000 | 1.5 | No | 2998.5 | 3001.5 | 0.1 | 0 | 0.30139 | 2.4885 | 3.229e-05% | 3.229e-05% | 6.457e-05% | 0.88103 |  | 2999.119 | 3000.881 |  |  |  | 0.003% |  |  | 0.02% |  | 3
4000 | 4000 | 4001 | 4000.8 | 4000.6 | 2 | No | 3998 | 4002 | 0.1 | 0.52915 | 0.66395 | 1.5061 | 0.004503% | 1.749% | 1.754% | 0.63641 |  | 3999.3636 | 4000.6364 |  |  |  | 0.005% |  |  | 0.025% |  | 4
5000 | 5000 | 5000 | 5000.5 | 5000.1667 | 2.5 | No | 4997.5 | 5002.5 | 0.1 | 0.28868 | 0.57807 | 2.1624 | 0.0001984% | 0.002714% | 0.002912% | 1.3128 |  | 4998.6872 | 5001.3128 |  |  |  | 0.01% |  |  | 0.03% |  | 5
6000 | 6000 | 6000 | 6000 | 6000 | 3 | No | 5997 | 6003 | 0.1 | 0 | 0.60069 | 2.4971 | 2.954e-05% | 2.954e-05% | 5.907e-05% | 1.7663 |  | 5998.2337 | 6001.7663 |  |  |  | 0.02% |  |  | 0.05% |  | 6
7000 | 7000 | 7000 | 7000 | 7000 | 3.5 | No | 6996.5 | 7003.5 | 0.1 | 0 | 0.7006 | 2.4979 | 2.93e-05% | 2.93e-05% | 5.861e-05% | 2.0612 |  | 6997.9388 | 7002.0612 |  |  |  | 0.025% |  |  | 0.1% |  | 7
8000 | 8000 | 8000 | 8000 | 8000 | 4 | No | 7996 | 8004 | 0.1 | 0 | 0.80052 | 2.4984 | 2.915e-05% | 2.915e-05% | 5.831e-05% | 2.3559 |  | 7997.6441 | 8002.3559 |  |  |  | 0.03% |  |  | 0.2% |  | 8
9000 | 9000 | 9000 | 9000 | 9000 | 4.5 | No | 8995.5 | 9004.5 | 0.1 | 0 | 0.90046 | 2.4987 | 2.905e-05% | 2.905e-05% | 5.81e-05% | 2.6507 |  | 8997.3493 | 9002.6507 |  |  |  | 0.05% |  |  | 0.25% |  | 9
10000 | 10000 | 10001.5 | 10000.8 | 10000.7667 | 5 | No | 9995 | 10005 | 0.1 | 0.75056 | 1.2507 | 1.9989 | 0.0002005% | 0.03561% | 0.03581% | 2.4314 |  | 9997.5686 | 10002.4314 |  |  |  | 0.1% |  |  | 0.5% |  | 10
[23 empty] | 0.15% |  |  | 1%
[23 empty] | 0.2% |  |  | 4%
0.3%
PFA CALCULATION FOR AN ENTIRE LOADING RANGE METHOD 6 | [22 empty] | 0.5%
1%
Instrument Capacity | 10000
Resolution UUT | 0.1
Tol UUT | 0.05% | Reading
Plus 1 Count | no
CMC | 0.02%
K = for CMC | 2
[7 empty] | Manufacturer's | [8 empty] | Acceptance Limit (A2%) - Method 6 | [7 empty] | 2
Force or Torque Applied | Instrument Reading 1 | Instrument Reading 2 | Instrument Reading 3 | Avg Instrument Reading | Req Tolerance | Plus 1 Count | LSL | USL | Res UUT | Rep UUT | U95% | TUR | (0.38xln(TUR)-0.54) | M2% | TOL (+/-L) |  |  | LSL | USL | Pass/Fail |  |  |  | 3
1000 | 1000 | 1000.1 | 1000 | 1000.0333 | 0.5 | No | 999.5 | 1000.5 | 0.1 | 0.057735 | 0.23329 | 2.1433 | -0.25031 | 26.14% | 0.5 | 0.43901 |  | 999.561 | 1000.439 | PASS | 2.1433 |  | U95%
2000 | 2000 | 2000 | 2000.1 | 2000.0333 | 1 | No | 1999 | 2001 | 0.1 | 0.057735 | 0.41191 | 2.4277 | -0.20296 | 22.37% | 1 | 0.90786 |  | 1999.0921 | 2000.9079 | PASS | 2.4277 |  | U98%
3000 | 3000 | 3000 | 3000 | 3000 | 1.5 | No | 2998.5 | 3001.5 | 0.1 | 0 | 0.59072 | 2.5393 | -0.18588 | 20.96% | 1.5 | 1.3762 |  | 2998.6238 | 3001.3762 | PASS | 2.5393
4000 | 4000 | 4001 | 4000.8 | 4000.6 | 2 | No | 3998 | 4002 | 0.1 | 0.52915 | 1.3013 | 1.5369 | -0.3767 | 35.39% | 2 | 1.5395 |  | 3998.4605 | 4001.5395 | PASS | 1.5369
5000 | 5000 | 5000 | 5000.5 | 5000.1667 | 2.5 | No | 4997.5 | 5002.5 | 0.1 | 0.28868 | 1.133 | 2.2065 | -0.23927 | 25.28% | 2.5 | 2.2136 |  | 4997.7864 | 5002.2136 | PASS | 2.2065
6000 | 6000 | 6000 | 6000 | 6000 | 3 | No | 5997 | 6003 | 0.1 | 0 | 1.1774 | 2.5481 | -0.18457 | 20.85% | 3 | 2.7545 |  | 5997.2455 | 6002.7545 | PASS | 2.5481
7000 | 7000 | 7000 | 7000 | 7000 | 3.5 | No | 6996.5 | 7003.5 | 0.1 | 0 | 1.3732 | 2.5489 | -0.18446 | 20.84% | 3.5 | 3.2138 |  | 6996.7862 | 7003.2138 | PASS | 2.5489
8000 | 8000 | 8000 | 8000 | 8000 | 4 | No | 7996 | 8004 | 0.1 | 0 | 1.569 | 2.5494 | -0.18438 | 20.84% | 4 | 3.673 |  | 7996.327 | 8003.673 | PASS | 2.5494
9000 | 9000 | 9000 | 9000 | 9000 | 4.5 | No | 8995.5 | 9004.5 | 0.1 | 0 | 1.7649 | 2.5497 | -0.18433 | 20.83% | 4.5 | 4.1323 |  | 8995.8677 | 9004.1323 | PASS | 2.5497
10000 | 10000 | 10001.5 | 10000.8 | 10000.7667 | 5 | No | 9995 | 10005 | 0.1 | 0.75056 | 2.4513 | 2.0397 | -0.26913 | 27.6% | 5 | 4.3235 |  | 9995.6765 | 10004.3235 | PASS | 2.0397
DIFFERENCE IN ACCEPTANCE LIMITS METHOD 5 VS METHOD 6
Force or Torque Applied | Avg Instrument Reading | Method 6 AL | PASS/FAIL | Method 5 AL | PASS/FAIL | % Diff in AL
1000 | 1000.0333 | 0.43901 | PASS | 0.25555 | PASS | 41.79%
2000 | 2000.0333 | 0.90786 | PASS | 0.56839 | PASS | 37.39%
3000 | 3000 | 1.3762 | PASS | 0.88103 | PASS | 35.98%
4000 | 4000.6 | 1.5395 | PASS | 0.63641 | PASS | 58.66%
5000 | 5000.1667 | 2.2136 | PASS | 1.3128 | PASS | 40.69%
6000 | 6000 | 2.7545 | PASS | 1.7663 | PASS | 35.87%
7000 | 7000 | 3.2138 | PASS | 2.0612 | PASS | 35.86%
8000 | 8000 | 3.673 | PASS | 2.3559 | PASS | 35.86%
9000 | 9000 | 4.1323 | PASS | 2.6507 | PASS | 35.85%
10000 | 10000.7667 | 4.3235 | PASS | 2.4314 | PASS | 43.76%
TEST LIMIT = SQUARE ROOT (SPEC^2 - UNCERTAINTY^2)
Instrument Capacity | 10000
Resolution UUT | 0.1
Tol UUT | 0.05% | Reading
Plus 1 Count | no
CMC | 0.02%
K = for CMC | 2
Manufacturer's
Force or Torque Applied | Instrument Reading 1 | Instrument Reading 2 | Instrument Reading 3 | Avg Instrument Reading | Req Tolerance | Plus 1 Count | LSL | USL | Res UUT | Rep UUT | U95% | TUR | TOL (+/-L) | TEST LIMIT |  | LSL | USL | PASS/FAIL
1000 | 1000 | 1000.1 | 1000 | 1000.0333 | 0.5 | No | 999.5 | 1000.5 | 0.1 | 0.057735 | 0.23329 | 2.1433 | 0.5 | 0.44224 |  | 999.5578 | 1000.4422 | PASS
2000 | 2000 | 2000 | 2000.1 | 2000.0333 | 1 | No | 1999 | 2001 | 0.1 | 0.057735 | 0.41191 | 2.4277 | 1 | 0.91122 |  | 1999.0888 | 2000.9112 | PASS
3000 | 3000 | 3000 | 3000 | 3000 | 1.5 | No | 2998.5 | 3001.5 | 0.1 | 0 | 0.59072 | 2.5393 | 1.5 | 1.3788 |  | 2998.6212 | 3001.3788 | PASS
4000 | 4000 | 4001 | 4000.8 | 4000.6 | 2 | No | 3998 | 4002 | 0.1 | 0.52915 | 1.3013 | 1.5369 | 2 | 1.5187 |  | 3998.4813 | 4001.5187 | PASS
5000 | 5000 | 5000 | 5000.5 | 5000.1667 | 2.5 | No | 4997.5 | 5002.5 | 0.1 | 0.28868 | 1.133 | 2.2065 | 2.5 | 2.2285 |  | 4997.7715 | 5002.2285 | PASS
6000 | 6000 | 6000 | 6000 | 6000 | 3 | No | 5997 | 6003 | 0.1 | 0 | 1.1774 | 2.5481 | 3 | 2.7593 |  | 5997.2407 | 6002.7593 | PASS
7000 | 7000 | 7000 | 7000 | 7000 | 3.5 | No | 6996.5 | 7003.5 | 0.1 | 0 | 1.3732 | 2.5489 | 3.5 | 3.2194 |  | 6996.7806 | 7003.2194 | PASS
8000 | 8000 | 8000 | 8000 | 8000 | 4 | No | 7996 | 8004 | 0.1 | 0 | 1.569 | 2.5494 | 4 | 3.6794 |  | 7996.3206 | 8003.6794 | PASS
9000 | 9000 | 9000 | 9000 | 9000 | 4.5 | No | 8995.5 | 9004.5 | 0.1 | 0 | 1.7649 | 2.5497 | 4.5 | 4.1395 |  | 8995.8605 | 9004.1395 | PASS
10000 | 10000 | 10001.5 | 10000.8 | 10000.7667 | 5 | No | 9995 | 10005 | 0.1 | 0.75056 | 2.4513 | 2.0397 | 5 | 4.3579 |  | 9995.6421 | 10004.3579 | PASS

## Z- Dist
Risk Calculator
Upper Tolerance TU | 12 |  | Conformance Probability Table
Lower Tolerance TL | 8 |  | Conformance Probability , Pc | Guard Band Multiplier, r Two - Sided
Nominal Value | 10
Measurement Unc um | 0.25 |  | 0.0668 | -0.75003
Measured Value xm | 11.5 |  | 0.159 | -0.49929
Tolerance T | 4 | Area Outside of USL | 0.3085 | -0.25005
Z Upper | 2 | 2.275% | 0.5 | 0
|  | Area Outside of LSL | 0.6914 | 0.24991
Z Lower | -14 | 7.794e-43% | 0.8 | 0.42081
|  |  | 0.85 | 0.51822
Setting the Guard Band Upper and Lower pc |  |  | 0.9 | 0.64078
Select Desired Conformance Probability | 95% |  | 0.95 | 0.82243
Maximum Risk if within GU & GL | 5% |  | 0.9545 | 0.84507
hU (GB Multiplier) | 0.82243 |  | 0.97725 | 1
Guard Band Upper Gu | 11.5888 |  | 0.98 | 1.0269
Guard Band Lower GL | 8.4112 |  | 0.99 | 1.1632 |  | 0.82243
|  |  | 0.999 | 3.0902
|  |  | 1 | 5.9978
Z-Score | Probability | Area Under Curve Single Sided
0 | 0.5 | 50%
0.1 | 0.53983 | 46.02%
0.2 | 0.57926 | 42.07%
0.3 | 0.61791 | 38.21%
0.4 | 0.65542 | 34.46%
0.5 | 0.69146 | 30.85%
0.6 | 0.72575 | 27.43%
0.7 | 0.75804 | 24.2%
0.8 | 0.78814 | 21.19%
0.9 | 0.81594 | 18.41%
1 | 0.84134 | 15.87%
1.1 | 0.86433 | 13.57%
1.2 | 0.88493 | 11.51%
1.3 | 0.9032 | 9.68%
1.4 | 0.91924 | 8.076%
1.5 | 0.93319 | 6.681%
1.6 | 0.9452 | 5.48%
1.7 | 0.95543 | 4.457%
1.8 | 0.96407 | 3.593%
1.9 | 0.97128 | 2.872%
2 | 0.97725 | 2.275%
2.1 | 0.98214 | 1.786%
2.2 | 0.9861 | 1.39%
2.3 | 0.98928 | 1.072%
2.4 | 0.9918 | 0.8198%
2.5 | 0.99379 | 0.621%
2.6 | 0.99534 | 0.4661%
2.7 | 0.99653 | 0.3467%
2.8 | 0.99744 | 0.2555%
2.9 | 0.99813 | 0.1866%
3 | 0.99865 | 0.135%
3.1 | 0.99903 | 0.09676%
3.2 | 0.99931 | 0.06871%
3.3 | 0.99952 | 0.04834%
3.4 | 0.99966 | 0.03369% |  |  |  | 0 |  |  |  | As-Found
3.5 | 0.99977 | 0.02326% |  |  |  | 0 | Nominal Value | LSL | USL | MV
3.6 | 0.99984 | 0.01591% |  |  |  |  | 10 | 8 | 12 | 11.5 | 0
3.7 | 0.99989 | 0.01078% |  |  |  |  | 10 | 8 | 12 | 11.5 | 1.5958
3.8 | 0.99993 | 0.007235%
3.9 | 0.99995 | 0.00481% | [5 empty] | LAL | UAL
4 | 0.99997 | 0.003167% |  |  | Increment | 0.002 |  | 8.4112 | 11.5888 |  | 0
4.1 | 0.99998 | 0.002066% | [5 empty] | 8.4112 | 11.5888 |  | 1.5958
4.2 | 0.99999 | 0.001335% |  |  | Norm. Dist.
4.3 | 0.99999 | 0.000854% |  |  | x | y
4.4 | 0.99999 | 0.0005413% |  |  | 10.5 | 0.00053532
4.5 | 1 | 0.0003398% |  |  | 10.502 | 0.00055271
4.6 | 1 | 0.0002112% |  |  | 10.504 | 0.00057063
4.7 | 1 | 0.0001301% |  |  | 10.506 | 0.00058909
4.8 | 1 | 7.933e-05% |  |  | 10.508 | 0.00060811 | NORMAL DISTRIBUTION FUNCTION
4.9 | 1 | 4.792e-05% |  |  | 10.51 | 0.0006277
5 | 1 | 2.867e-05% |  |  | 10.512 | 0.00064789 | x = | x-axis value
5.1 | 1 | 1.698e-05% |  |  | 10.514 | 0.00066868 | u = | Centerline of distribution
5.2 | 1 | 9.964e-06% |  |  | 10.516 | 0.00069009 | σ = | Std. Dev.
5.3 | 1 | 5.79e-06% |  |  | 10.518 | 0.00071214
5.4 | 1 | 3.332e-06% |  |  | 10.52 | 0.00073485
5.5 | 1 | 1.899e-06% |  |  | 10.522 | 0.00075824
5.6 | 1 | 1.072e-06% |  |  | 10.524 | 0.00078232
5.7 | 1 | 5.99e-07% |  |  | 10.526 | 0.00080711
5.8 | 1 | 3.316e-07% |  |  | 10.528 | 0.00083263
5.9 | 1 | 1.818e-07% |  |  | 10.53 | 0.00085891
6 | 1 | 9.866e-08% |  |  | 10.532 | 0.00088596
[5 empty] | 10.534 | 0.00091381
[5 empty] | 10.536 | 0.00094246
[5 empty] | 10.538 | 0.00097196
[5 empty] | 10.54 | 0.0010023
[5 empty] | 10.542 | 0.0010335
[5 empty] | 10.544 | 0.0010657
[5 empty] | 10.546 | 0.0010988
[5 empty] | 10.548 | 0.0011328
[5 empty] | 10.55 | 0.0011678
[5 empty] | 10.552 | 0.0012038
[5 empty] | 10.554 | 0.0012408
[5 empty] | 10.556 | 0.0012789
[5 empty] | 10.558 | 0.0013181
[5 empty] | 10.56 | 0.0013584
[5 empty] | 10.562 | 0.0013998
[5 empty] | 10.564 | 0.0014425
[5 empty] | 10.566 | 0.0014863
[5 empty] | 10.568 | 0.0015313
[5 empty] | 10.57 | 0.0015776
[5 empty] | 10.572 | 0.0016252
[5 empty] | 10.574 | 0.0016741
[5 empty] | 10.576 | 0.0017244
[5 empty] | 10.578 | 0.0017761
[5 empty] | 10.58 | 0.0018293
[5 empty] | 10.582 | 0.0018839
[5 empty] | 10.584 | 0.00194
[5 empty] | 10.586 | 0.0019976
[5 empty] | 10.588 | 0.0020568
[5 empty] | 10.59 | 0.0021177
[5 empty] | 10.592 | 0.0021802
[5 empty] | 10.594 | 0.0022444
[5 empty] | 10.596 | 0.0023103
[5 empty] | 10.598 | 0.002378
[5 empty] | 10.6 | 0.0024476
[5 empty] | 10.602 | 0.002519
[5 empty] | 10.604 | 0.0025924
[5 empty] | 10.606 | 0.0026677
[5 empty] | 10.608 | 0.002745
[5 empty] | 10.61 | 0.0028244
[5 empty] | 10.612 | 0.0029059
[5 empty] | 10.614 | 0.0029896
[5 empty] | 10.616 | 0.0030755
[5 empty] | 10.618 | 0.0031636
[5 empty] | 10.62 | 0.0032541
[5 empty] | 10.622 | 0.0033469
[5 empty] | 10.624 | 0.0034422
[5 empty] | 10.626 | 0.0035399
[5 empty] | 10.628 | 0.0036402
[5 empty] | 10.63 | 0.0037431
[5 empty] | 10.632 | 0.0038486
[5 empty] | 10.634 | 0.0039569
[5 empty] | 10.636 | 0.004068
[5 empty] | 10.638 | 0.0041819
[5 empty] | 10.64 | 0.0042987
[5 empty] | 10.642 | 0.0044185
[5 empty] | 10.644 | 0.0045413
[5 empty] | 10.646 | 0.0046673
[5 empty] | 10.648 | 0.0047965
[5 empty] | 10.65 | 0.0049289
[5 empty] | 10.652 | 0.0050646
[5 empty] | 10.654 | 0.0052038
[5 empty] | 10.656 | 0.0053464
[5 empty] | 10.658 | 0.0054926
[5 empty] | 10.66 | 0.0056424
[5 empty] | 10.662 | 0.0057959
[5 empty] | 10.664 | 0.0059533
[5 empty] | 10.666 | 0.0061145
[5 empty] | 10.668 | 0.0062797
[5 empty] | 10.67 | 0.0064489
[5 empty] | 10.672 | 0.0066223
[5 empty] | 10.674 | 0.0067999
[5 empty] | 10.676 | 0.0069818
[5 empty] | 10.678 | 0.0071681
[5 empty] | 10.68 | 0.0073589
[5 empty] | 10.682 | 0.0075543
[5 empty] | 10.684 | 0.0077544
[5 empty] | 10.686 | 0.0079593
[5 empty] | 10.688 | 0.0081691
[5 empty] | 10.69 | 0.0083839
[5 empty] | 10.692 | 0.0086038
[5 empty] | 10.694 | 0.0088288
[5 empty] | 10.696 | 0.0090592
[5 empty] | 10.698 | 0.009295
[5 empty] | 10.7 | 0.0095364
[5 empty] | 10.702 | 0.0097833
[5 empty] | 10.704 | 0.010036
[5 empty] | 10.706 | 0.010295
[5 empty] | 10.708 | 0.010559
[5 empty] | 10.71 | 0.01083
[5 empty] | 10.712 | 0.011107
[5 empty] | 10.714 | 0.01139
[5 empty] | 10.716 | 0.01168
[5 empty] | 10.718 | 0.011976
[5 empty] | 10.72 | 0.012279
[5 empty] | 10.722 | 0.012589
[5 empty] | 10.724 | 0.012906
[5 empty] | 10.726 | 0.01323
[5 empty] | 10.728 | 0.013562
[5 empty] | 10.73 | 0.0139
[5 empty] | 10.732 | 0.014247
[5 empty] | 10.734 | 0.014601
[5 empty] | 10.736 | 0.014962
[5 empty] | 10.738 | 0.015332
[5 empty] | 10.74 | 0.01571
[5 empty] | 10.742 | 0.016096
[5 empty] | 10.744 | 0.016491
[5 empty] | 10.746 | 0.016894
[5 empty] | 10.748 | 0.017306
[5 empty] | 10.75 | 0.017727
[5 empty] | 10.752 | 0.018157
[5 empty] | 10.754 | 0.018597
[5 empty] | 10.756 | 0.019045
[5 empty] | 10.758 | 0.019504
[5 empty] | 10.76 | 0.019972
[5 empty] | 10.762 | 0.02045
[5 empty] | 10.764 | 0.020938
[5 empty] | 10.766 | 0.021436
[5 empty] | 10.768 | 0.021945
[5 empty] | 10.77 | 0.022464
[5 empty] | 10.772 | 0.022994
[5 empty] | 10.774 | 0.023535
[5 empty] | 10.776 | 0.024088
[5 empty] | 10.778 | 0.024652
[5 empty] | 10.78 | 0.025227
[5 empty] | 10.782 | 0.025814
[5 empty] | 10.784 | 0.026413
[5 empty] | 10.786 | 0.027024
[5 empty] | 10.788 | 0.027648
[5 empty] | 10.79 | 0.028284
[5 empty] | 10.792 | 0.028933
[5 empty] | 10.794 | 0.029596
[5 empty] | 10.796 | 0.030271
[5 empty] | 10.798 | 0.030959
[5 empty] | 10.8 | 0.031662
[5 empty] | 10.802 | 0.032378
[5 empty] | 10.804 | 0.033108
[5 empty] | 10.806 | 0.033853
[5 empty] | 10.808 | 0.034612
[5 empty] | 10.81 | 0.035386
[5 empty] | 10.812 | 0.036175
[5 empty] | 10.814 | 0.036979
[5 empty] | 10.816 | 0.037798
[5 empty] | 10.818 | 0.038633
[5 empty] | 10.82 | 0.039485
[5 empty] | 10.822 | 0.040352
[5 empty] | 10.824 | 0.041236
[5 empty] | 10.826 | 0.042136
[5 empty] | 10.828 | 0.043053
[5 empty] | 10.83 | 0.043988
[5 empty] | 10.832 | 0.04494
[5 empty] | 10.834 | 0.045909
[5 empty] | 10.836 | 0.046897
[5 empty] | 10.838 | 0.047902
[5 empty] | 10.84 | 0.048926
[5 empty] | 10.842 | 0.049969
[5 empty] | 10.844 | 0.05103
[5 empty] | 10.846 | 0.052111
[5 empty] | 10.848 | 0.053212
[5 empty] | 10.85 | 0.054332
[5 empty] | 10.852 | 0.055472
[5 empty] | 10.854 | 0.056633
[5 empty] | 10.856 | 0.057814
[5 empty] | 10.858 | 0.059015
[5 empty] | 10.86 | 0.060238
[5 empty] | 10.862 | 0.061483
[5 empty] | 10.864 | 0.062749
[5 empty] | 10.866 | 0.064037
[5 empty] | 10.868 | 0.065347
[5 empty] | 10.87 | 0.06668
[5 empty] | 10.872 | 0.068036
[5 empty] | 10.874 | 0.069415
[5 empty] | 10.876 | 0.070817
[5 empty] | 10.878 | 0.072243
[5 empty] | 10.88 | 0.073693
[5 empty] | 10.882 | 0.075168
[5 empty] | 10.884 | 0.076666
[5 empty] | 10.886 | 0.07819
[5 empty] | 10.888 | 0.079739
[5 empty] | 10.89 | 0.081313
[5 empty] | 10.892 | 0.082914
[5 empty] | 10.894 | 0.08454
[5 empty] | 10.896 | 0.086192
[5 empty] | 10.898 | 0.087872
[5 empty] | 10.9 | 0.089578
[5 empty] | 10.902 | 0.091312
[5 empty] | 10.904 | 0.093073
[5 empty] | 10.906 | 0.094862
[5 empty] | 10.908 | 0.096679
[5 empty] | 10.91 | 0.098525
[5 empty] | 10.912 | 0.1004
[5 empty] | 10.914 | 0.1023
[5 empty] | 10.916 | 0.10424
[5 empty] | 10.918 | 0.1062
[5 empty] | 10.92 | 0.10819
[5 empty] | 10.922 | 0.11022
[5 empty] | 10.924 | 0.11227
[5 empty] | 10.926 | 0.11435
[5 empty] | 10.928 | 0.11647
[5 empty] | 10.93 | 0.11862
[5 empty] | 10.932 | 0.1208
[5 empty] | 10.934 | 0.12301
[5 empty] | 10.936 | 0.12525
[5 empty] | 10.938 | 0.12753
[5 empty] | 10.94 | 0.12984
[5 empty] | 10.942 | 0.13218
[5 empty] | 10.944 | 0.13456
[5 empty] | 10.946 | 0.13697
[5 empty] | 10.948 | 0.13942
[5 empty] | 10.95 | 0.1419
[5 empty] | 10.952 | 0.14441
[5 empty] | 10.954 | 0.14696
[5 empty] | 10.956 | 0.14955
[5 empty] | 10.958 | 0.15217
[5 empty] | 10.96 | 0.15483
[5 empty] | 10.962 | 0.15752
[5 empty] | 10.964 | 0.16025
[5 empty] | 10.966 | 0.16302
[5 empty] | 10.968 | 0.16582
[5 empty] | 10.97 | 0.16866
[5 empty] | 10.972 | 0.17154
[5 empty] | 10.974 | 0.17446
[5 empty] | 10.976 | 0.17742
[5 empty] | 10.978 | 0.18041
[5 empty] | 10.98 | 0.18344
[5 empty] | 10.982 | 0.18652
[5 empty] | 10.984 | 0.18963
[5 empty] | 10.986 | 0.19278
[5 empty] | 10.988 | 0.19597
[5 empty] | 10.99 | 0.1992
[5 empty] | 10.992 | 0.20247
[5 empty] | 10.994 | 0.20578
[5 empty] | 10.996 | 0.20914
[5 empty] | 10.998 | 0.21253
[5 empty] | 11 | 0.21596
[5 empty] | 11.002 | 0.21944
[5 empty] | 11.004 | 0.22296
[5 empty] | 11.006 | 0.22652
[5 empty] | 11.008 | 0.23012
[5 empty] | 11.01 | 0.23376
[5 empty] | 11.012 | 0.23745
[5 empty] | 11.014 | 0.24118
[5 empty] | 11.016 | 0.24495
[5 empty] | 11.018 | 0.24877
[5 empty] | 11.02 | 0.25263
[5 empty] | 11.022 | 0.25653
[5 empty] | 11.024 | 0.26047
[5 empty] | 11.026 | 0.26446
[5 empty] | 11.028 | 0.2685
[5 empty] | 11.03 | 0.27257
[5 empty] | 11.032 | 0.2767
[5 empty] | 11.034 | 0.28086
[5 empty] | 11.036 | 0.28507
[5 empty] | 11.038 | 0.28933
[5 empty] | 11.04 | 0.29363
[5 empty] | 11.042 | 0.29797
[5 empty] | 11.044 | 0.30236
[5 empty] | 11.046 | 0.3068
[5 empty] | 11.048 | 0.31128
[5 empty] | 11.05 | 0.3158
[5 empty] | 11.052 | 0.32037
[5 empty] | 11.054 | 0.32499
[5 empty] | 11.056 | 0.32965
[5 empty] | 11.058 | 0.33435
[5 empty] | 11.06 | 0.33911
[5 empty] | 11.062 | 0.3439
[5 empty] | 11.064 | 0.34875
[5 empty] | 11.066 | 0.35363
[5 empty] | 11.068 | 0.35857
[5 empty] | 11.07 | 0.36355
[5 empty] | 11.072 | 0.36857
[5 empty] | 11.074 | 0.37364
[5 empty] | 11.076 | 0.37876
[5 empty] | 11.078 | 0.38392
[5 empty] | 11.08 | 0.38913
[5 empty] | 11.082 | 0.39438
[5 empty] | 11.084 | 0.39968
[5 empty] | 11.086 | 0.40502
[5 empty] | 11.088 | 0.41041
[5 empty] | 11.09 | 0.41584
[5 empty] | 11.092 | 0.42132
[5 empty] | 11.094 | 0.42685
[5 empty] | 11.096 | 0.43241
[5 empty] | 11.098 | 0.43803
[5 empty] | 11.1 | 0.44368
[5 empty] | 11.102 | 0.44938
[5 empty] | 11.104 | 0.45513
[5 empty] | 11.106 | 0.46092
[5 empty] | 11.108 | 0.46675
[5 empty] | 11.11 | 0.47263
[5 empty] | 11.112 | 0.47855
[5 empty] | 11.114 | 0.48451
[5 empty] | 11.116 | 0.49052
[5 empty] | 11.118 | 0.49657
[5 empty] | 11.12 | 0.50266
[5 empty] | 11.122 | 0.50879
[5 empty] | 11.124 | 0.51497
[5 empty] | 11.126 | 0.52118
[5 empty] | 11.128 | 0.52744
[5 empty] | 11.13 | 0.53374
[5 empty] | 11.132 | 0.54008
[5 empty] | 11.134 | 0.54646
[5 empty] | 11.136 | 0.55288
[5 empty] | 11.138 | 0.55934
[5 empty] | 11.14 | 0.56584
[5 empty] | 11.142 | 0.57238
[5 empty] | 11.144 | 0.57895
[5 empty] | 11.146 | 0.58557
[5 empty] | 11.148 | 0.59222
[5 empty] | 11.15 | 0.59891
[5 empty] | 11.152 | 0.60564
[5 empty] | 11.154 | 0.6124
[5 empty] | 11.156 | 0.6192
[5 empty] | 11.158 | 0.62603
[5 empty] | 11.16 | 0.6329
[5 empty] | 11.162 | 0.6398
[5 empty] | 11.164 | 0.64674
[5 empty] | 11.166 | 0.65371
[5 empty] | 11.168 | 0.66071
[5 empty] | 11.17 | 0.66775
[5 empty] | 11.172 | 0.67482
[5 empty] | 11.174 | 0.68191
[5 empty] | 11.176 | 0.68904
[5 empty] | 11.178 | 0.6962
[5 empty] | 11.18 | 0.70339
[5 empty] | 11.182 | 0.71061
[5 empty] | 11.184 | 0.71785
[5 empty] | 11.186 | 0.72512
[5 empty] | 11.188 | 0.73242
[5 empty] | 11.19 | 0.73975
[5 empty] | 11.192 | 0.7471
[5 empty] | 11.194 | 0.75448
[5 empty] | 11.196 | 0.76188
[5 empty] | 11.198 | 0.7693
[5 empty] | 11.2 | 0.77674
[5 empty] | 11.202 | 0.78421
[5 empty] | 11.204 | 0.7917
[5 empty] | 11.206 | 0.79921
[5 empty] | 11.208 | 0.80674
[5 empty] | 11.21 | 0.81429
[5 empty] | 11.212 | 0.82185
[5 empty] | 11.214 | 0.82943
[5 empty] | 11.216 | 0.83703
[5 empty] | 11.218 | 0.84465
[5 empty] | 11.22 | 0.85228
[5 empty] | 11.222 | 0.85992
[5 empty] | 11.224 | 0.86758
[5 empty] | 11.226 | 0.87524
[5 empty] | 11.228 | 0.88292
[5 empty] | 11.23 | 0.89061
[5 empty] | 11.232 | 0.89831
[5 empty] | 11.234 | 0.90602
[5 empty] | 11.236 | 0.91374
[5 empty] | 11.238 | 0.92146
[5 empty] | 11.24 | 0.92919
[5 empty] | 11.242 | 0.93692
[5 empty] | 11.244 | 0.94466
[5 empty] | 11.246 | 0.9524
[5 empty] | 11.248 | 0.96014
[5 empty] | 11.25 | 0.96788
[5 empty] | 11.252 | 0.97563
[5 empty] | 11.254 | 0.98337
[5 empty] | 11.256 | 0.99111
[5 empty] | 11.258 | 0.99884
[5 empty] | 11.26 | 1.0066
[5 empty] | 11.262 | 1.0143
[5 empty] | 11.264 | 1.022
[5 empty] | 11.266 | 1.0297
[5 empty] | 11.268 | 1.0374
[5 empty] | 11.27 | 1.0451
[5 empty] | 11.272 | 1.0528
[5 empty] | 11.274 | 1.0605
[5 empty] | 11.276 | 1.0682
[5 empty] | 11.278 | 1.0758
[5 empty] | 11.28 | 1.0835
[5 empty] | 11.282 | 1.0911
[5 empty] | 11.284 | 1.0987
[5 empty] | 11.286 | 1.1063
[5 empty] | 11.288 | 1.1138
[5 empty] | 11.29 | 1.1214
[5 empty] | 11.292 | 1.1289
[5 empty] | 11.294 | 1.1364
[5 empty] | 11.296 | 1.1439
[5 empty] | 11.298 | 1.1513
[5 empty] | 11.3 | 1.1588
[5 empty] | 11.302 | 1.1662
[5 empty] | 11.304 | 1.1735
[5 empty] | 11.306 | 1.1809
[5 empty] | 11.308 | 1.1882
[5 empty] | 11.31 | 1.1955
[5 empty] | 11.312 | 1.2027
[5 empty] | 11.314 | 1.21
[5 empty] | 11.316 | 1.2171
[5 empty] | 11.318 | 1.2243
[5 empty] | 11.32 | 1.2314
[5 empty] | 11.322 | 1.2385
[5 empty] | 11.324 | 1.2455
[5 empty] | 11.326 | 1.2525
[5 empty] | 11.328 | 1.2595
[5 empty] | 11.33 | 1.2664
[5 empty] | 11.332 | 1.2732
[5 empty] | 11.334 | 1.2801
[5 empty] | 11.336 | 1.2868
[5 empty] | 11.338 | 1.2936
[5 empty] | 11.34 | 1.3002
[5 empty] | 11.342 | 1.3069
[5 empty] | 11.344 | 1.3135
[5 empty] | 11.346 | 1.32
[5 empty] | 11.348 | 1.3265
[5 empty] | 11.35 | 1.3329
[5 empty] | 11.352 | 1.3393
[5 empty] | 11.354 | 1.3456
[5 empty] | 11.356 | 1.3518
[5 empty] | 11.358 | 1.358
[5 empty] | 11.36 | 1.3642
[5 empty] | 11.362 | 1.3703
[5 empty] | 11.364 | 1.3763
[5 empty] | 11.366 | 1.3822
[5 empty] | 11.368 | 1.3881
[5 empty] | 11.37 | 1.394
[5 empty] | 11.372 | 1.3997
[5 empty] | 11.374 | 1.4054
[5 empty] | 11.376 | 1.4111
[5 empty] | 11.378 | 1.4166
[5 empty] | 11.38 | 1.4221
[5 empty] | 11.382 | 1.4276
[5 empty] | 11.384 | 1.4329
[5 empty] | 11.386 | 1.4382
[5 empty] | 11.388 | 1.4434
[5 empty] | 11.39 | 1.4485
[5 empty] | 11.392 | 1.4536
[5 empty] | 11.394 | 1.4586
[5 empty] | 11.396 | 1.4635
[5 empty] | 11.398 | 1.4683
[5 empty] | 11.4 | 1.4731
[5 empty] | 11.402 | 1.4778
[5 empty] | 11.404 | 1.4823
[5 empty] | 11.406 | 1.4869
[5 empty] | 11.408 | 1.4913
[5 empty] | 11.41 | 1.4956
[5 empty] | 11.412 | 1.4999
[5 empty] | 11.414 | 1.5041
[5 empty] | 11.416 | 1.5082
[5 empty] | 11.418 | 1.5122
[5 empty] | 11.42 | 1.5161
[5 empty] | 11.422 | 1.52
[5 empty] | 11.424 | 1.5237
[5 empty] | 11.426 | 1.5274
[5 empty] | 11.428 | 1.5309
[5 empty] | 11.43 | 1.5344
[5 empty] | 11.432 | 1.5378
[5 empty] | 11.434 | 1.5411
[5 empty] | 11.436 | 1.5443
[5 empty] | 11.438 | 1.5474
[5 empty] | 11.44 | 1.5505
[5 empty] | 11.442 | 1.5534
[5 empty] | 11.444 | 1.5562
[5 empty] | 11.446 | 1.559
[5 empty] | 11.448 | 1.5616
[5 empty] | 11.45 | 1.5642
[5 empty] | 11.452 | 1.5666
[5 empty] | 11.454 | 1.569
[5 empty] | 11.456 | 1.5712
[5 empty] | 11.458 | 1.5734
[5 empty] | 11.46 | 1.5755
[5 empty] | 11.462 | 1.5774
[5 empty] | 11.464 | 1.5793
[5 empty] | 11.466 | 1.5811
[5 empty] | 11.468 | 1.5827
[5 empty] | 11.47 | 1.5843
[5 empty] | 11.472 | 1.5858
[5 empty] | 11.474 | 1.5872
[5 empty] | 11.476 | 1.5884
[5 empty] | 11.478 | 1.5896
[5 empty] | 11.48 | 1.5907
[5 empty] | 11.482 | 1.5916
[5 empty] | 11.484 | 1.5925
[5 empty] | 11.486 | 1.5933
[5 empty] | 11.488 | 1.5939
[5 empty] | 11.49 | 1.5945
[5 empty] | 11.492 | 1.595
[5 empty] | 11.494 | 1.5953
[5 empty] | 11.496 | 1.5956
[5 empty] | 11.498 | 1.5957
[5 empty] | 11.5 | 1.5958
[5 empty] | 11.502 | 1.5957
[5 empty] | 11.504 | 1.5956
[5 empty] | 11.506 | 1.5953
[5 empty] | 11.508 | 1.595
[5 empty] | 11.51 | 1.5945
[5 empty] | 11.512 | 1.5939
[5 empty] | 11.514 | 1.5933
[5 empty] | 11.516 | 1.5925
[5 empty] | 11.518 | 1.5916
[5 empty] | 11.52 | 1.5907
[5 empty] | 11.522 | 1.5896
[5 empty] | 11.524 | 1.5884
[5 empty] | 11.526 | 1.5872
[5 empty] | 11.528 | 1.5858
[5 empty] | 11.53 | 1.5843
[5 empty] | 11.532 | 1.5827
[5 empty] | 11.534 | 1.5811
[5 empty] | 11.536 | 1.5793
[5 empty] | 11.538 | 1.5774
[5 empty] | 11.54 | 1.5755
[5 empty] | 11.542 | 1.5734
[5 empty] | 11.544 | 1.5712
[5 empty] | 11.546 | 1.569
[5 empty] | 11.548 | 1.5666
[5 empty] | 11.55 | 1.5642
[5 empty] | 11.552 | 1.5616
[5 empty] | 11.554 | 1.559
[5 empty] | 11.556 | 1.5562
[5 empty] | 11.558 | 1.5534
[5 empty] | 11.56 | 1.5505
[5 empty] | 11.562 | 1.5474
[5 empty] | 11.564 | 1.5443
[5 empty] | 11.566 | 1.5411
[5 empty] | 11.568 | 1.5378
[5 empty] | 11.57 | 1.5344
[5 empty] | 11.572 | 1.5309
[5 empty] | 11.574 | 1.5274
[5 empty] | 11.576 | 1.5237
[5 empty] | 11.578 | 1.52
[5 empty] | 11.58 | 1.5161
[5 empty] | 11.582 | 1.5122
[5 empty] | 11.584 | 1.5082
[5 empty] | 11.586 | 1.5041
[5 empty] | 11.588 | 1.4999
[5 empty] | 11.59 | 1.4956
[5 empty] | 11.592 | 1.4913
[5 empty] | 11.594 | 1.4869
[5 empty] | 11.596 | 1.4823
[5 empty] | 11.598 | 1.4778
[5 empty] | 11.6 | 1.4731
[5 empty] | 11.602 | 1.4683
[5 empty] | 11.604 | 1.4635
[5 empty] | 11.606 | 1.4586
[5 empty] | 11.608 | 1.4536
[5 empty] | 11.61 | 1.4485
[5 empty] | 11.612 | 1.4434
[5 empty] | 11.614 | 1.4382
[5 empty] | 11.616 | 1.4329
[5 empty] | 11.618 | 1.4276
[5 empty] | 11.62 | 1.4221
[5 empty] | 11.622 | 1.4166
[5 empty] | 11.624 | 1.4111
[5 empty] | 11.626 | 1.4054
[5 empty] | 11.628 | 1.3997
[5 empty] | 11.63 | 1.394
[5 empty] | 11.632 | 1.3881
[5 empty] | 11.634 | 1.3822
[5 empty] | 11.636 | 1.3763
[5 empty] | 11.638 | 1.3703
[5 empty] | 11.64 | 1.3642
[5 empty] | 11.642 | 1.358
[5 empty] | 11.644 | 1.3518
[5 empty] | 11.646 | 1.3456
[5 empty] | 11.648 | 1.3393
[5 empty] | 11.65 | 1.3329
[5 empty] | 11.652 | 1.3265
[5 empty] | 11.654 | 1.32
[5 empty] | 11.656 | 1.3135
[5 empty] | 11.658 | 1.3069
[5 empty] | 11.66 | 1.3002
[5 empty] | 11.662 | 1.2936
[5 empty] | 11.664 | 1.2868
[5 empty] | 11.666 | 1.2801
[5 empty] | 11.668 | 1.2732
[5 empty] | 11.67 | 1.2664
[5 empty] | 11.672 | 1.2595
[5 empty] | 11.674 | 1.2525
[5 empty] | 11.676 | 1.2455
[5 empty] | 11.678 | 1.2385
[5 empty] | 11.68 | 1.2314
[5 empty] | 11.682 | 1.2243
[5 empty] | 11.684 | 1.2171
[5 empty] | 11.686 | 1.21
[5 empty] | 11.688 | 1.2027
[5 empty] | 11.69 | 1.1955
[5 empty] | 11.692 | 1.1882
[5 empty] | 11.694 | 1.1809
[5 empty] | 11.696 | 1.1735
[5 empty] | 11.698 | 1.1662
[5 empty] | 11.7 | 1.1588
[5 empty] | 11.702 | 1.1513
[5 empty] | 11.704 | 1.1439
[5 empty] | 11.706 | 1.1364
[5 empty] | 11.708 | 1.1289
[5 empty] | 11.71 | 1.1214
[5 empty] | 11.712 | 1.1138
[5 empty] | 11.714 | 1.1063
[5 empty] | 11.716 | 1.0987
[5 empty] | 11.718 | 1.0911
[5 empty] | 11.72 | 1.0835
[5 empty] | 11.722 | 1.0758
[5 empty] | 11.724 | 1.0682
[5 empty] | 11.726 | 1.0605
[5 empty] | 11.728 | 1.0528
[5 empty] | 11.73 | 1.0451
[5 empty] | 11.732 | 1.0374
[5 empty] | 11.734 | 1.0297
[5 empty] | 11.736 | 1.022
[5 empty] | 11.738 | 1.0143
[5 empty] | 11.74 | 1.0066
[5 empty] | 11.742 | 0.99884
[5 empty] | 11.744 | 0.99111
[5 empty] | 11.746 | 0.98337
[5 empty] | 11.748 | 0.97563
[5 empty] | 11.75 | 0.96788
[5 empty] | 11.752 | 0.96014
[5 empty] | 11.754 | 0.9524
[5 empty] | 11.756 | 0.94466
[5 empty] | 11.758 | 0.93692
[5 empty] | 11.76 | 0.92919
[5 empty] | 11.762 | 0.92146
[5 empty] | 11.764 | 0.91374
[5 empty] | 11.766 | 0.90602
[5 empty] | 11.768 | 0.89831
[5 empty] | 11.77 | 0.89061
[5 empty] | 11.772 | 0.88292
[5 empty] | 11.774 | 0.87524
[5 empty] | 11.776 | 0.86758
[5 empty] | 11.778 | 0.85992
[5 empty] | 11.78 | 0.85228
[5 empty] | 11.782 | 0.84465
[5 empty] | 11.784 | 0.83703
[5 empty] | 11.786 | 0.82943
[5 empty] | 11.788 | 0.82185
[5 empty] | 11.79 | 0.81429
[5 empty] | 11.792 | 0.80674
[5 empty] | 11.794 | 0.79921
[5 empty] | 11.796 | 0.7917
[5 empty] | 11.798 | 0.78421
[5 empty] | 11.8 | 0.77674
[5 empty] | 11.802 | 0.7693
[5 empty] | 11.804 | 0.76188
[5 empty] | 11.806 | 0.75448
[5 empty] | 11.808 | 0.7471
[5 empty] | 11.81 | 0.73975
[5 empty] | 11.812 | 0.73242
[5 empty] | 11.814 | 0.72512
[5 empty] | 11.816 | 0.71785
[5 empty] | 11.818 | 0.71061
[5 empty] | 11.82 | 0.70339
[5 empty] | 11.822 | 0.6962
[5 empty] | 11.824 | 0.68904
[5 empty] | 11.826 | 0.68191
[5 empty] | 11.828 | 0.67482
[5 empty] | 11.83 | 0.66775
[5 empty] | 11.832 | 0.66071
[5 empty] | 11.834 | 0.65371
[5 empty] | 11.836 | 0.64674
[5 empty] | 11.838 | 0.6398
[5 empty] | 11.84 | 0.6329
[5 empty] | 11.842 | 0.62603
[5 empty] | 11.844 | 0.6192
[5 empty] | 11.846 | 0.6124
[5 empty] | 11.848 | 0.60564
[5 empty] | 11.85 | 0.59891
[5 empty] | 11.852 | 0.59222
[5 empty] | 11.854 | 0.58557
[5 empty] | 11.856 | 0.57895
[5 empty] | 11.858 | 0.57238
[5 empty] | 11.86 | 0.56584
[5 empty] | 11.862 | 0.55934
[5 empty] | 11.864 | 0.55288
[5 empty] | 11.866 | 0.54646
[5 empty] | 11.868 | 0.54008
[5 empty] | 11.87 | 0.53374
[5 empty] | 11.872 | 0.52744
[5 empty] | 11.874 | 0.52118
[5 empty] | 11.876 | 0.51497
[5 empty] | 11.878 | 0.50879
[5 empty] | 11.88 | 0.50266
[5 empty] | 11.882 | 0.49657
[5 empty] | 11.884 | 0.49052
[5 empty] | 11.886 | 0.48451
[5 empty] | 11.888 | 0.47855
[5 empty] | 11.89 | 0.47263
[5 empty] | 11.892 | 0.46675
[5 empty] | 11.894 | 0.46092
[5 empty] | 11.896 | 0.45513
[5 empty] | 11.898 | 0.44938
[5 empty] | 11.9 | 0.44368
[5 empty] | 11.902 | 0.43803
[5 empty] | 11.904 | 0.43241
[5 empty] | 11.906 | 0.42685
[5 empty] | 11.908 | 0.42132
[5 empty] | 11.91 | 0.41584
[5 empty] | 11.912 | 0.41041
[5 empty] | 11.914 | 0.40502
[5 empty] | 11.916 | 0.39968
[5 empty] | 11.918 | 0.39438
[5 empty] | 11.92 | 0.38913
[5 empty] | 11.922 | 0.38392
[5 empty] | 11.924 | 0.37876
[5 empty] | 11.926 | 0.37364
[5 empty] | 11.928 | 0.36857
[5 empty] | 11.93 | 0.36355
[5 empty] | 11.932 | 0.35857
[5 empty] | 11.934 | 0.35363
[5 empty] | 11.936 | 0.34875
[5 empty] | 11.938 | 0.3439
[5 empty] | 11.94 | 0.33911
[5 empty] | 11.942 | 0.33435
[5 empty] | 11.944 | 0.32965
[5 empty] | 11.946 | 0.32499
[5 empty] | 11.948 | 0.32037
[5 empty] | 11.95 | 0.3158
[5 empty] | 11.952 | 0.31128
[5 empty] | 11.954 | 0.3068
[5 empty] | 11.956 | 0.30236
[5 empty] | 11.958 | 0.29797
[5 empty] | 11.96 | 0.29363
[5 empty] | 11.962 | 0.28933
[5 empty] | 11.964 | 0.28507
[5 empty] | 11.966 | 0.28086
[5 empty] | 11.968 | 0.2767
[5 empty] | 11.97 | 0.27257
[5 empty] | 11.972 | 0.2685
[5 empty] | 11.974 | 0.26446
[5 empty] | 11.976 | 0.26047
[5 empty] | 11.978 | 0.25653
[5 empty] | 11.98 | 0.25263
[5 empty] | 11.982 | 0.24877
[5 empty] | 11.984 | 0.24495
[5 empty] | 11.986 | 0.24118
[5 empty] | 11.988 | 0.23745
[5 empty] | 11.99 | 0.23376
[5 empty] | 11.992 | 0.23012
[5 empty] | 11.994 | 0.22652
[5 empty] | 11.996 | 0.22296
[5 empty] | 11.998 | 0.21944
[5 empty] | 12 | 0.21596
[5 empty] | 12.002 | 0.21253
[5 empty] | 12.004 | 0.20914
[5 empty] | 12.006 | 0.20578
[5 empty] | 12.008 | 0.20247
[5 empty] | 12.01 | 0.1992
[5 empty] | 12.012 | 0.19597
[5 empty] | 12.014 | 0.19278
[5 empty] | 12.016 | 0.18963
[5 empty] | 12.018 | 0.18652
[5 empty] | 12.02 | 0.18344
[5 empty] | 12.022 | 0.18041
[5 empty] | 12.024 | 0.17742
[5 empty] | 12.026 | 0.17446
[5 empty] | 12.028 | 0.17154
[5 empty] | 12.03 | 0.16866
[5 empty] | 12.032 | 0.16582
[5 empty] | 12.034 | 0.16302
[5 empty] | 12.036 | 0.16025
[5 empty] | 12.038 | 0.15752
[5 empty] | 12.04 | 0.15483
[5 empty] | 12.042 | 0.15217
[5 empty] | 12.044 | 0.14955
[5 empty] | 12.046 | 0.14696
[5 empty] | 12.048 | 0.14441
[5 empty] | 12.05 | 0.1419
[5 empty] | 12.052 | 0.13942
[5 empty] | 12.054 | 0.13697
[5 empty] | 12.056 | 0.13456
[5 empty] | 12.058 | 0.13218
[5 empty] | 12.06 | 0.12984
[5 empty] | 12.062 | 0.12753
[5 empty] | 12.064 | 0.12525
[5 empty] | 12.066 | 0.12301
[5 empty] | 12.068 | 0.1208
[5 empty] | 12.07 | 0.11862
[5 empty] | 12.072 | 0.11647
[5 empty] | 12.074 | 0.11435
[5 empty] | 12.076 | 0.11227
[5 empty] | 12.078 | 0.11022
[5 empty] | 12.08 | 0.10819
[5 empty] | 12.082 | 0.1062
[5 empty] | 12.084 | 0.10424
[5 empty] | 12.086 | 0.1023
[5 empty] | 12.088 | 0.1004
[5 empty] | 12.09 | 0.098525
[5 empty] | 12.092 | 0.096679
[5 empty] | 12.094 | 0.094862
[5 empty] | 12.096 | 0.093073
[5 empty] | 12.098 | 0.091312
[5 empty] | 12.1 | 0.089578
[5 empty] | 12.102 | 0.087872
[5 empty] | 12.104 | 0.086192
[5 empty] | 12.106 | 0.08454
[5 empty] | 12.108 | 0.082914
[5 empty] | 12.11 | 0.081313
[5 empty] | 12.112 | 0.079739
[5 empty] | 12.114 | 0.07819
[5 empty] | 12.116 | 0.076666
[5 empty] | 12.118 | 0.075168
[5 empty] | 12.12 | 0.073693
[5 empty] | 12.122 | 0.072243
[5 empty] | 12.124 | 0.070817
[5 empty] | 12.126 | 0.069415
[5 empty] | 12.128 | 0.068036
[5 empty] | 12.13 | 0.06668
[5 empty] | 12.132 | 0.065347
[5 empty] | 12.134 | 0.064037
[5 empty] | 12.136 | 0.062749
[5 empty] | 12.138 | 0.061483
[5 empty] | 12.14 | 0.060238
[5 empty] | 12.142 | 0.059015
[5 empty] | 12.144 | 0.057814
[5 empty] | 12.146 | 0.056633
[5 empty] | 12.148 | 0.055472
[5 empty] | 12.15 | 0.054332
[5 empty] | 12.152 | 0.053212
[5 empty] | 12.154 | 0.052111
[5 empty] | 12.156 | 0.05103
[5 empty] | 12.158 | 0.049969
[5 empty] | 12.16 | 0.048926
[5 empty] | 12.162 | 0.047902
[5 empty] | 12.164 | 0.046897
[5 empty] | 12.166 | 0.045909
[5 empty] | 12.168 | 0.04494
[5 empty] | 12.17 | 0.043988
[5 empty] | 12.172 | 0.043053
[5 empty] | 12.174 | 0.042136
[5 empty] | 12.176 | 0.041236
[5 empty] | 12.178 | 0.040352
[5 empty] | 12.18 | 0.039485
[5 empty] | 12.182 | 0.038633
[5 empty] | 12.184 | 0.037798
[5 empty] | 12.186 | 0.036979
[5 empty] | 12.188 | 0.036175
[5 empty] | 12.19 | 0.035386
[5 empty] | 12.192 | 0.034612
[5 empty] | 12.194 | 0.033853
[5 empty] | 12.196 | 0.033108
[5 empty] | 12.198 | 0.032378
[5 empty] | 12.2 | 0.031662
[5 empty] | 12.202 | 0.030959
[5 empty] | 12.204 | 0.030271
[5 empty] | 12.206 | 0.029596
[5 empty] | 12.208 | 0.028933
[5 empty] | 12.21 | 0.028284
[5 empty] | 12.212 | 0.027648
[5 empty] | 12.214 | 0.027024
[5 empty] | 12.216 | 0.026413
[5 empty] | 12.218 | 0.025814
[5 empty] | 12.22 | 0.025227
[5 empty] | 12.222 | 0.024652
[5 empty] | 12.224 | 0.024088
[5 empty] | 12.226 | 0.023535
[5 empty] | 12.228 | 0.022994
[5 empty] | 12.23 | 0.022464
[5 empty] | 12.232 | 0.021945
[5 empty] | 12.234 | 0.021436
[5 empty] | 12.236 | 0.020938
[5 empty] | 12.238 | 0.02045
[5 empty] | 12.24 | 0.019972
[5 empty] | 12.242 | 0.019504
[5 empty] | 12.244 | 0.019045
[5 empty] | 12.246 | 0.018597
[5 empty] | 12.248 | 0.018157
[5 empty] | 12.25 | 0.017727
[5 empty] | 12.252 | 0.017306
[5 empty] | 12.254 | 0.016894
[5 empty] | 12.256 | 0.016491
[5 empty] | 12.258 | 0.016096
[5 empty] | 12.26 | 0.01571
[5 empty] | 12.262 | 0.015332
[5 empty] | 12.264 | 0.014962
[5 empty] | 12.266 | 0.014601
[5 empty] | 12.268 | 0.014247
[5 empty] | 12.27 | 0.0139
[5 empty] | 12.272 | 0.013562
[5 empty] | 12.274 | 0.01323
[5 empty] | 12.276 | 0.012906
[5 empty] | 12.278 | 0.012589
[5 empty] | 12.28 | 0.012279
[5 empty] | 12.282 | 0.011976
[5 empty] | 12.284 | 0.01168
[5 empty] | 12.286 | 0.01139
[5 empty] | 12.288 | 0.011107
[5 empty] | 12.29 | 0.01083
[5 empty] | 12.292 | 0.010559
[5 empty] | 12.294 | 0.010295
[5 empty] | 12.296 | 0.010036
[5 empty] | 12.298 | 0.0097833
[5 empty] | 12.3 | 0.0095364
[5 empty] | 12.302 | 0.009295
[5 empty] | 12.304 | 0.0090592
[5 empty] | 12.306 | 0.0088288
[5 empty] | 12.308 | 0.0086038
[5 empty] | 12.31 | 0.0083839
[5 empty] | 12.312 | 0.0081691
[5 empty] | 12.314 | 0.0079593
[5 empty] | 12.316 | 0.0077544
[5 empty] | 12.318 | 0.0075543
[5 empty] | 12.32 | 0.0073589
[5 empty] | 12.322 | 0.0071681
[5 empty] | 12.324 | 0.0069818
[5 empty] | 12.326 | 0.0067999
[5 empty] | 12.328 | 0.0066223
[5 empty] | 12.33 | 0.0064489
[5 empty] | 12.332 | 0.0062797
[5 empty] | 12.334 | 0.0061145
[5 empty] | 12.336 | 0.0059533
[5 empty] | 12.338 | 0.0057959
[5 empty] | 12.34 | 0.0056424
[5 empty] | 12.342 | 0.0054926
[5 empty] | 12.344 | 0.0053464
[5 empty] | 12.346 | 0.0052038
[5 empty] | 12.348 | 0.0050646
[5 empty] | 12.35 | 0.0049289
[5 empty] | 12.352 | 0.0047965
[5 empty] | 12.354 | 0.0046673
[5 empty] | 12.356 | 0.0045413
[5 empty] | 12.358 | 0.0044185
[5 empty] | 12.36 | 0.0042987
[5 empty] | 12.362 | 0.0041819
[5 empty] | 12.364 | 0.004068
[5 empty] | 12.366 | 0.0039569
[5 empty] | 12.368 | 0.0038486
[5 empty] | 12.37 | 0.0037431
[5 empty] | 12.372 | 0.0036402
[5 empty] | 12.374 | 0.0035399
[5 empty] | 12.376 | 0.0034422
[5 empty] | 12.378 | 0.0033469
[5 empty] | 12.38 | 0.0032541
[5 empty] | 12.382 | 0.0031636
[5 empty] | 12.384 | 0.0030755
[5 empty] | 12.386 | 0.0029896
[5 empty] | 12.388 | 0.0029059
[5 empty] | 12.39 | 0.0028244
[5 empty] | 12.392 | 0.002745
[5 empty] | 12.394 | 0.0026677
[5 empty] | 12.396 | 0.0025924
[5 empty] | 12.398 | 0.002519
[5 empty] | 12.4 | 0.0024476
[5 empty] | 12.402 | 0.002378
[5 empty] | 12.404 | 0.0023103
[5 empty] | 12.406 | 0.0022444
[5 empty] | 12.408 | 0.0021802
[5 empty] | 12.41 | 0.0021177
[5 empty] | 12.412 | 0.0020568
[5 empty] | 12.414 | 0.0019976
[5 empty] | 12.416 | 0.00194
[5 empty] | 12.418 | 0.0018839
[5 empty] | 12.42 | 0.0018293
[5 empty] | 12.422 | 0.0017761
[5 empty] | 12.424 | 0.0017244
[5 empty] | 12.426 | 0.0016741
[5 empty] | 12.428 | 0.0016252
[5 empty] | 12.43 | 0.0015776
[5 empty] | 12.432 | 0.0015313
[5 empty] | 12.434 | 0.0014863
[5 empty] | 12.436 | 0.0014425
[5 empty] | 12.438 | 0.0013998
[5 empty] | 12.44 | 0.0013584
[5 empty] | 12.442 | 0.0013181
[5 empty] | 12.444 | 0.0012789
[5 empty] | 12.446 | 0.0012408
[5 empty] | 12.448 | 0.0012038
[5 empty] | 12.45 | 0.0011678
[5 empty] | 12.452 | 0.0011328
[5 empty] | 12.454 | 0.0010988
[5 empty] | 12.456 | 0.0010657
[5 empty] | 12.458 | 0.0010335
[5 empty] | 12.46 | 0.0010023
[5 empty] | 12.462 | 0.00097196
[5 empty] | 12.464 | 0.00094246
[5 empty] | 12.466 | 0.00091381
[5 empty] | 12.468 | 0.00088596
[5 empty] | 12.47 | 0.00085891
[5 empty] | 12.472 | 0.00083263
[5 empty] | 12.474 | 0.00080711
[5 empty] | 12.476 | 0.00078232
[5 empty] | 12.478 | 0.00075824
[5 empty] | 12.48 | 0.00073485
[5 empty] | 12.482 | 0.00071214
[5 empty] | 12.484 | 0.00069009
[5 empty] | 12.486 | 0.00066868
[5 empty] | 12.488 | 0.00064789
[5 empty] | 12.49 | 0.0006277
[5 empty] | 12.492 | 0.00060811
[5 empty] | 12.494 | 0.00058909
[5 empty] | 12.496 | 0.00057063
[5 empty] | 12.498 | 0.00055271

## z-table
Areas Under the Normal Curve
z - Table
z | 0 | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09
0 | 0 | 0.0039894 | 0.0079783 | 0.011966 | 0.015953 | 0.019939 | 0.023922 | 0.027903 | 0.031881 | 0.035856
0.1 | 0.039828 | 0.043795 | 0.047758 | 0.051717 | 0.05567 | 0.059618 | 0.063559 | 0.067495 | 0.071424 | 0.075345
0.2 | 0.07926 | 0.083166 | 0.087064 | 0.090954 | 0.094835 | 0.098706 | 0.10257 | 0.10642 | 0.11026 | 0.11409
0.3 | 0.11791 | 0.12172 | 0.12552 | 0.1293 | 0.13307 | 0.13683 | 0.14058 | 0.14431 | 0.14803 | 0.15173
0.4 | 0.15542 | 0.1591 | 0.16276 | 0.1664 | 0.17003 | 0.17364 | 0.17724 | 0.18082 | 0.18439 | 0.18793
0.5 | 0.19146 | 0.19497 | 0.19847 | 0.20194 | 0.2054 | 0.20884 | 0.21226 | 0.21566 | 0.21904 | 0.2224
0.6 | 0.22575 | 0.22907 | 0.23237 | 0.23565 | 0.23891 | 0.24215 | 0.24537 | 0.24857 | 0.25175 | 0.2549
0.7 | 0.25804 | 0.26115 | 0.26424 | 0.2673 | 0.27035 | 0.27337 | 0.27637 | 0.27935 | 0.2823 | 0.28524
0.8 | 0.28814 | 0.29103 | 0.29389 | 0.29673 | 0.29955 | 0.30234 | 0.30511 | 0.30785 | 0.31057 | 0.31327
0.9 | 0.31594 | 0.31859 | 0.32121 | 0.32381 | 0.32639 | 0.32894 | 0.33147 | 0.33398 | 0.33646 | 0.33891
1 | 0.34134 | 0.34375 | 0.34614 | 0.34849 | 0.35083 | 0.35314 | 0.35543 | 0.35769 | 0.35993 | 0.36214
1.1 | 0.36433 | 0.3665 | 0.36864 | 0.37076 | 0.37286 | 0.37493 | 0.37698 | 0.379 | 0.381 | 0.38298
1.2 | 0.38493 | 0.38686 | 0.38877 | 0.39065 | 0.39251 | 0.39435 | 0.39617 | 0.39796 | 0.39973 | 0.40147
1.3 | 0.4032 | 0.4049 | 0.40658 | 0.40824 | 0.40988 | 0.41149 | 0.41309 | 0.41466 | 0.41621 | 0.41774
1.4 | 0.41924 | 0.42073 | 0.4222 | 0.42364 | 0.42507 | 0.42647 | 0.42785 | 0.42922 | 0.43056 | 0.43189
1.5 | 0.43319 | 0.43448 | 0.43574 | 0.43699 | 0.43822 | 0.43943 | 0.44062 | 0.44179 | 0.44295 | 0.44408
1.6 | 0.4452 | 0.4463 | 0.44738 | 0.44845 | 0.4495 | 0.45053 | 0.45154 | 0.45254 | 0.45352 | 0.45449
1.7 | 0.45543 | 0.45637 | 0.45728 | 0.45818 | 0.45907 | 0.45994 | 0.4608 | 0.46164 | 0.46246 | 0.46327
1.8 | 0.46407 | 0.46485 | 0.46562 | 0.46638 | 0.46712 | 0.46784 | 0.46856 | 0.46926 | 0.46995 | 0.47062
1.9 | 0.47128 | 0.47193 | 0.47257 | 0.4732 | 0.47381 | 0.47441 | 0.475 | 0.47558 | 0.47615 | 0.4767
2 | 0.47725 | 0.47778 | 0.47831 | 0.47882 | 0.47932 | 0.47982 | 0.4803 | 0.48077 | 0.48124 | 0.48169 | 0.9545
2.1 | 0.48214 | 0.48257 | 0.483 | 0.48341 | 0.48382 | 0.48422 | 0.48461 | 0.485 | 0.48537 | 0.48574
2.2 | 0.4861 | 0.48645 | 0.48679 | 0.48713 | 0.48745 | 0.48778 | 0.48809 | 0.4884 | 0.4887 | 0.48899
2.3 | 0.48928 | 0.48956 | 0.48983 | 0.4901 | 0.49036 | 0.49061 | 0.49086 | 0.49111 | 0.49134 | 0.49158
2.4 | 0.4918 | 0.49202 | 0.49224 | 0.49245 | 0.49266 | 0.49286 | 0.49305 | 0.49324 | 0.49343 | 0.49361 | 0.9836
2.5 | 0.49379 | 0.49396 | 0.49413 | 0.4943 | 0.49446 | 0.49461 | 0.49477 | 0.49492 | 0.49506 | 0.4952
2.6 | 0.49534 | 0.49547 | 0.4956 | 0.49573 | 0.49585 | 0.49598 | 0.49609 | 0.49621 | 0.49632 | 0.49643
2.7 | 0.49653 | 0.49664 | 0.49674 | 0.49683 | 0.49693 | 0.49702 | 0.49711 | 0.4972 | 0.49728 | 0.49736
2.8 | 0.49744 | 0.49752 | 0.4976 | 0.49767 | 0.49774 | 0.49781 | 0.49788 | 0.49795 | 0.49801 | 0.49807
2.9 | 0.49813 | 0.49819 | 0.49825 | 0.49831 | 0.49836 | 0.49841 | 0.49846 | 0.49851 | 0.49856 | 0.49861
3 | 0.49865 | 0.49869 | 0.49874 | 0.49878 | 0.49882 | 0.49886 | 0.49889 | 0.49893 | 0.49896 | 0.499 | 0.9973
3.5 | 0.49977 | 0.49978 | 0.49978 | 0.49979 | 0.4998 | 0.49981 | 0.49981 | 0.49982 | 0.49983 | 0.49983
4 | 0.49997 | 0.49997 | 0.49997 | 0.49997 | 0.49997 | 0.49997 | 0.49998 | 0.49998 | 0.49998 | 0.49998
4.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
5.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
6 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
z | 0 | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09
Areas Under the Normal Curve
z - Table
z | 0 | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09
-3 | 0.0013499 | 0.0013949 | 0.0014412 | 0.001489 | 0.0015382 | 0.0015889 | 0.0016411 | 0.0016948 | 0.0017502 | 0.0018071
-2.9 | 0.0018658 | 0.0019262 | 0.0019884 | 0.0020524 | 0.0021182 | 0.002186 | 0.0022557 | 0.0023274 | 0.0024012 | 0.0024771
-2.8 | 0.0025551 | 0.0026354 | 0.0027179 | 0.0028028 | 0.0028901 | 0.0029798 | 0.003072 | 0.0031667 | 0.0032641 | 0.0033642
-2.7 | 0.003467 | 0.0035726 | 0.0036811 | 0.0037926 | 0.003907 | 0.0040246 | 0.0041453 | 0.0042692 | 0.0043965 | 0.0045271
-2.6 | 0.0046612 | 0.0047988 | 0.00494 | 0.0050849 | 0.0052336 | 0.0053861 | 0.0055426 | 0.0057031 | 0.0058677 | 0.0060366
-2.5 | 0.0062097 | 0.0063872 | 0.0065691 | 0.0067557 | 0.0069469 | 0.0071428 | 0.0073436 | 0.0075494 | 0.0077603 | 0.0079763
-2.4 | 0.0081975 | 0.0084242 | 0.0086563 | 0.008894 | 0.0091375 | 0.0093867 | 0.0096419 | 0.0099031 | 0.01017 | 0.010444
-2.3 | 0.010724 | 0.011011 | 0.011304 | 0.011604 | 0.011911 | 0.012224 | 0.012545 | 0.012874 | 0.013209 | 0.013553
-2.2 | 0.013903 | 0.014262 | 0.014629 | 0.015003 | 0.015386 | 0.015778 | 0.016177 | 0.016586 | 0.017003 | 0.017429
-2.1 | 0.017864 | 0.018309 | 0.018763 | 0.019226 | 0.019699 | 0.020182 | 0.020675 | 0.021178 | 0.021692 | 0.022216
-2 | 0.02275 | 0.023295 | 0.023852 | 0.024419 | 0.024998 | 0.025588 | 0.02619 | 0.026803 | 0.027429 | 0.028067
-1.9 | 0.028717 | 0.029379 | 0.030054 | 0.030742 | 0.031443 | 0.032157 | 0.032884 | 0.033625 | 0.03438 | 0.035148
-1.8 | 0.03593 | 0.036727 | 0.037538 | 0.038364 | 0.039204 | 0.040059 | 0.04093 | 0.041815 | 0.042716 | 0.043633
-1.7 | 0.044565 | 0.045514 | 0.046479 | 0.04746 | 0.048457 | 0.049471 | 0.050503 | 0.051551 | 0.052616 | 0.053699
-1.6 | 0.054799 | 0.055917 | 0.057053 | 0.058208 | 0.05938 | 0.060571 | 0.06178 | 0.063008 | 0.064255 | 0.065522
-1.5 | 0.066807 | 0.068112 | 0.069437 | 0.070781 | 0.072145 | 0.073529 | 0.074934 | 0.076359 | 0.077804 | 0.07927
-1.4 | 0.080757 | 0.082264 | 0.083793 | 0.085343 | 0.086915 | 0.088508 | 0.090123 | 0.091759 | 0.093418 | 0.095098
-1.3 | 0.0968 | 0.098525 | 0.10027 | 0.10204 | 0.10383 | 0.10565 | 0.10749 | 0.10935 | 0.11123 | 0.11314
-1.2 | 0.11507 | 0.11702 | 0.119 | 0.121 | 0.12302 | 0.12507 | 0.12714 | 0.12924 | 0.13136 | 0.1335
-1.1 | 0.13567 | 0.13786 | 0.14007 | 0.14231 | 0.14457 | 0.14686 | 0.14917 | 0.15151 | 0.15386 | 0.15625
-1 | 0.15866 | 0.16109 | 0.16354 | 0.16602 | 0.16853 | 0.17106 | 0.17361 | 0.17619 | 0.17879 | 0.18141
-0.9 | 0.18406 | 0.18673 | 0.18943 | 0.19215 | 0.19489 | 0.19766 | 0.20045 | 0.20327 | 0.20611 | 0.20897
-0.8 | 0.21186 | 0.21476 | 0.2177 | 0.22065 | 0.22363 | 0.22663 | 0.22965 | 0.2327 | 0.23576 | 0.23885
-0.7 | 0.24196 | 0.2451 | 0.24825 | 0.25143 | 0.25463 | 0.25785 | 0.26109 | 0.26435 | 0.26763 | 0.27093
-0.6 | 0.27425 | 0.2776 | 0.28096 | 0.28434 | 0.28774 | 0.29116 | 0.2946 | 0.29806 | 0.30153 | 0.30503
-0.5 | 0.30854 | 0.31207 | 0.31561 | 0.31918 | 0.32276 | 0.32636 | 0.32997 | 0.3336 | 0.33724 | 0.3409
-0.4 | 0.34458 | 0.34827 | 0.35197 | 0.35569 | 0.35942 | 0.36317 | 0.36693 | 0.3707 | 0.37448 | 0.37828
-0.3 | 0.38209 | 0.38591 | 0.38974 | 0.39358 | 0.39743 | 0.40129 | 0.40517 | 0.40905 | 0.41294 | 0.41683
-0.2 | 0.42074 | 0.42465 | 0.42858 | 0.43251 | 0.43644 | 0.44038 | 0.44433 | 0.44828 | 0.45224 | 0.4562
-0.1 | 0.46017 | 0.46414 | 0.46812 | 0.4721 | 0.47608 | 0.48006 | 0.48405 | 0.48803 | 0.49202 | 0.49601
0 | 0.5 | 0.50399 | 0.50798 | 0.51197 | 0.51595 | 0.51994 | 0.52392 | 0.5279 | 0.53188 | 0.53586
0.1 | 0.53983 | 0.5438 | 0.54776 | 0.55172 | 0.55567 | 0.55962 | 0.56356 | 0.56749 | 0.57142 | 0.57535
0.2 | 0.57926 | 0.58317 | 0.58706 | 0.59095 | 0.59483 | 0.59871 | 0.60257 | 0.60642 | 0.61026 | 0.61409
0.3 | 0.61791 | 0.62172 | 0.62552 | 0.6293 | 0.63307 | 0.63683 | 0.64058 | 0.64431 | 0.64803 | 0.65173
0.4 | 0.65542 | 0.6591 | 0.66276 | 0.6664 | 0.67003 | 0.67364 | 0.67724 | 0.68082 | 0.68439 | 0.68793
0.5 | 0.69146 | 0.69497 | 0.69847 | 0.70194 | 0.7054 | 0.70884 | 0.71226 | 0.71566 | 0.71904 | 0.7224
0.6 | 0.72575 | 0.72907 | 0.73237 | 0.73565 | 0.73891 | 0.74215 | 0.74537 | 0.74857 | 0.75175 | 0.7549
0.7 | 0.75804 | 0.76115 | 0.76424 | 0.7673 | 0.77035 | 0.77337 | 0.77637 | 0.77935 | 0.7823 | 0.78524
0.8 | 0.78814 | 0.79103 | 0.79389 | 0.79673 | 0.79955 | 0.80234 | 0.80511 | 0.80785 | 0.81057 | 0.81327
0.9 | 0.81594 | 0.81859 | 0.82121 | 0.82381 | 0.82639 | 0.82894 | 0.83147 | 0.83398 | 0.83646 | 0.83891
1 | 0.84134 | 0.84375 | 0.84614 | 0.84849 | 0.85083 | 0.85314 | 0.85543 | 0.85769 | 0.85993 | 0.86214
1.1 | 0.86433 | 0.8665 | 0.86864 | 0.87076 | 0.87286 | 0.87493 | 0.87698 | 0.879 | 0.881 | 0.88298
1.2 | 0.88493 | 0.88686 | 0.88877 | 0.89065 | 0.89251 | 0.89435 | 0.89617 | 0.89796 | 0.89973 | 0.90147
1.3 | 0.9032 | 0.9049 | 0.90658 | 0.90824 | 0.90988 | 0.91149 | 0.91309 | 0.91466 | 0.91621 | 0.91774
1.4 | 0.91924 | 0.92073 | 0.9222 | 0.92364 | 0.92507 | 0.92647 | 0.92785 | 0.92922 | 0.93056 | 0.93189
1.5 | 0.93319 | 0.93448 | 0.93574 | 0.93699 | 0.93822 | 0.93943 | 0.94062 | 0.94179 | 0.94295 | 0.94408
1.6 | 0.9452 | 0.9463 | 0.94738 | 0.94845 | 0.9495 | 0.95053 | 0.95154 | 0.95254 | 0.95352 | 0.95449
1.7 | 0.95543 | 0.95637 | 0.95728 | 0.95818 | 0.95907 | 0.95994 | 0.9608 | 0.96164 | 0.96246 | 0.96327
1.8 | 0.96407 | 0.96485 | 0.96562 | 0.96638 | 0.96712 | 0.96784 | 0.96856 | 0.96926 | 0.96995 | 0.97062
1.9 | 0.97128 | 0.97193 | 0.97257 | 0.9732 | 0.97381 | 0.97441 | 0.975 | 0.97558 | 0.97615 | 0.9767
2 | 0.97725 | 0.97778 | 0.97831 | 0.97882 | 0.97932 | 0.97982 | 0.9803 | 0.98077 | 0.98124 | 0.98169
2.1 | 0.98214 | 0.98257 | 0.983 | 0.98341 | 0.98382 | 0.98422 | 0.98461 | 0.985 | 0.98537 | 0.98574
2.2 | 0.9861 | 0.98645 | 0.98679 | 0.98713 | 0.98745 | 0.98778 | 0.98809 | 0.9884 | 0.9887 | 0.98899
2.3 | 0.98928 | 0.98956 | 0.98983 | 0.9901 | 0.99036 | 0.99061 | 0.99086 | 0.99111 | 0.99134 | 0.99158
2.4 | 0.9918 | 0.99202 | 0.99224 | 0.99245 | 0.99266 | 0.99286 | 0.99305 | 0.99324 | 0.99343 | 0.99361
2.5 | 0.99379 | 0.99396 | 0.99413 | 0.9943 | 0.99446 | 0.99461 | 0.99477 | 0.99492 | 0.99506 | 0.9952
2.6 | 0.99534 | 0.99547 | 0.9956 | 0.99573 | 0.99585 | 0.99598 | 0.99609 | 0.99621 | 0.99632 | 0.99643
2.7 | 0.99653 | 0.99664 | 0.99674 | 0.99683 | 0.99693 | 0.99702 | 0.99711 | 0.9972 | 0.99728 | 0.99736
2.8 | 0.99744 | 0.99752 | 0.9976 | 0.99767 | 0.99774 | 0.99781 | 0.99788 | 0.99795 | 0.99801 | 0.99807
2.9 | 0.99813 | 0.99819 | 0.99825 | 0.99831 | 0.99836 | 0.99841 | 0.99846 | 0.99851 | 0.99856 | 0.99861
3 | 0.99865 | 0.99869 | 0.99874 | 0.99878 | 0.99882 | 0.99886 | 0.99889 | 0.99893 | 0.99896 | 0.999
3.1 | 0.99903 | 0.99906 | 0.9991 | 0.99913 | 0.99916 | 0.99918 | 0.99921 | 0.99924 | 0.99926 | 0.99929
3.2 | 0.99931 | 0.99934 | 0.99936 | 0.99938 | 0.9994 | 0.99942 | 0.99944 | 0.99946 | 0.99948 | 0.9995
3.3 | 0.99952 | 0.99953 | 0.99955 | 0.99957 | 0.99958 | 0.9996 | 0.99961 | 0.99962 | 0.99964 | 0.99965
3.4 | 0.99966 | 0.99968 | 0.99969 | 0.9997 | 0.99971 | 0.99972 | 0.99973 | 0.99974 | 0.99975 | 0.99976
3.5 | 0.99977 | 0.99978 | 0.99978 | 0.99979 | 0.9998 | 0.99981 | 0.99981 | 0.99982 | 0.99983 | 0.99983
3.6 | 0.99984 | 0.99985 | 0.99985 | 0.99986 | 0.99986 | 0.99987 | 0.99987 | 0.99988 | 0.99988 | 0.99989
3.7 | 0.99989 | 0.9999 | 0.9999 | 0.9999 | 0.99991 | 0.99991 | 0.99992 | 0.99992 | 0.99992 | 0.99992
3.8 | 0.99993 | 0.99993 | 0.99993 | 0.99994 | 0.99994 | 0.99994 | 0.99994 | 0.99995 | 0.99995 | 0.99995
3.9 | 0.99995 | 0.99995 | 0.99996 | 0.99996 | 0.99996 | 0.99996 | 0.99996 | 0.99996 | 0.99997 | 0.99997
4 | 0.99997 | 0.99997 | 0.99997 | 0.99997 | 0.99997 | 0.99997 | 0.99998 | 0.99998 | 0.99998 | 0.99998
z | 0 | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09

## Coverage Factor k
Normal Distribution k - value & Probability |  | Probability & Coverage Factor
Coverage Factor (k) | Probability | Probability | Coverage Factor (k)
0.25 | 19.74% | 10% | 0.12566
0.5 | 38.29% | 15% | 0.18912
0.83 | 59.35% | 20% | 0.25335
1 | 68.27% | 25% | 0.31864
1.25 | 78.87% | 30% | 0.38532
1.5 | 86.64% | 35% | 0.45376
1.75 | 91.99% | 40% | 0.5244
1.96 | 95% | 45% | 0.59776
2 | 95.45% | 50% | 0.67449
2.25 | 97.56% | 55% | 0.75542
2.5 | 98.76% | 60% | 0.84162
2.75 | 99.4% | 65% | 0.93459
3 | 99.73% | 68.27% | 1
3.25 | 99.88% | 70% | 1.0364
3.5 | 99.95% | 75% | 1.1503
3.75 | 99.98% | 80% | 1.2816
4 | 99.99% | 85% | 1.4395
4.25 | 100% | 90% | 1.6449
4.5 | 100% | 95% | 1.96
4.75 | 100% | 95.45% | 2
5 | 100% | 98.36% | 2.4
5.25 | 100% | 99% | 2.5758
5.5 | 100% | 99.9% | 3.2905
5.75 | 100% | 99.99% | 3.8906
6 | 100% | 100% | 4.4172

## G8 AL = TL ± w
Risk Calculator
Upper Tolerance TU | 1500.2 | [7 empty] | r
Lower Tolerance TL | 1499.8 | [7 empty] | -1.645
Nominal Value | 1500 | [7 empty] | -1.5
Measurement Unc um | 0.04 | [7 empty] | -1.25
Measured Value xm | 1500.12 | [7 empty] | -1
Tolerance T | 0.4 | [7 empty] | -0.75
-0.5
-0.25
Probability of Conformance (pc) | 97.73% | 50% | [6 empty] | 0
Probability of NonConformance (1 - pc) | 2.275% | [7 empty] | 0.01
0.05
Setting the Guard Band Upper and Lower AL | [8 empty] | 0.1
Guard Band Upper Gu (AL= TL - w) | 1500.12 | [7 empty] | 0.25
Guard Band Lower GL (AL = TL + w) | 1499.88 | [7 empty] | 0.5
0.83
Setting AL based on Probability of Conformance | [8 empty] | 1
Probability of Conformance (pc) = | 97.72% | [7 empty] | 1.391
r | 1 | [7 empty] | 1.5
w = U95 * r | 0.08 | [7 empty] | 1.645
Cm | 2.5 | [7 empty] | 1.75
2
Setting AL based on Guard Band w | [8 empty] | 2.25
Upper Acceptance Limit | PASS | [7 empty] | 2.5
Lower Acceptance Limit | PASS | [7 empty] | 2.75
3
Area of Curve Outside of the AL | 50% | [7 empty] | 3.25
3.5
3.75
4
4.25
4.5
4.75
5
5.25
5.5
5.75
6
|  |  |  | 0 |  |  |  | As-Found
|  |  |  | 0 | Nominal Value | LSL | USL | MV
[5 empty] | 1500 | 1499.8 | 1500.2 | 1500.12 | 0
[5 empty] | 1500 | 1499.8 | 1500.2 | 1500.12 | 9.9736
[6 empty] | LAL | UAL
|  |  | Increment | 0.00032 |  | 1499.88 | 1500.12 |  | 0
[6 empty] | 1499.88 | 1500.12 |  | 9.9736
Norm. Dist.
|  |  | x | y
|  |  | 1499.96 | 0.0033458
|  |  | 1499.9603 | 0.0034544
|  |  | 1499.9606 | 0.0035664
|  |  | 1499.961 | 0.0036818
|  |  | 1499.9613 | 0.0038007 | NORMAL DISTRIBUTION FUNCTION
|  |  | 1499.9616 | 0.0039231
|  |  | 1499.9619 | 0.0040493 | x = | x-axis value
|  |  | 1499.9622 | 0.0041792 | u = | Centerline of distribution
|  |  | 1499.9626 | 0.004313 | σ = | Std. Dev.
|  |  | 1499.9629 | 0.0044509
|  |  | 1499.9632 | 0.0045928
|  |  | 1499.9635 | 0.004739
|  |  | 1499.9638 | 0.0048895
|  |  | 1499.9642 | 0.0050444
|  |  | 1499.9645 | 0.005204
|  |  | 1499.9648 | 0.0053682
|  |  | 1499.9651 | 0.0055373
|  |  | 1499.9654 | 0.0057113
|  |  | 1499.9658 | 0.0058904
|  |  | 1499.9661 | 0.0060747
|  |  | 1499.9664 | 0.0062645
|  |  | 1499.9667 | 0.0064597
|  |  | 1499.967 | 0.0066606
|  |  | 1499.9674 | 0.0068673
|  |  | 1499.9677 | 0.0070799
|  |  | 1499.968 | 0.0072987
|  |  | 1499.9683 | 0.0075237
|  |  | 1499.9686 | 0.0077552
|  |  | 1499.969 | 0.0079933
|  |  | 1499.9693 | 0.0082382
|  |  | 1499.9696 | 0.00849
|  |  | 1499.9699 | 0.008749
|  |  | 1499.9702 | 0.0090153
|  |  | 1499.9706 | 0.0092891
|  |  | 1499.9709 | 0.0095706
|  |  | 1499.9712 | 0.0098601
|  |  | 1499.9715 | 0.010158
|  |  | 1499.9718 | 0.010463
|  |  | 1499.9722 | 0.010778
|  |  | 1499.9725 | 0.011101
|  |  | 1499.9728 | 0.011433
|  |  | 1499.9731 | 0.011774
|  |  | 1499.9734 | 0.012125
|  |  | 1499.9738 | 0.012485
|  |  | 1499.9741 | 0.012855
|  |  | 1499.9744 | 0.013235
|  |  | 1499.9747 | 0.013626
|  |  | 1499.975 | 0.014027
|  |  | 1499.9754 | 0.014439
|  |  | 1499.9757 | 0.014863
|  |  | 1499.976 | 0.015298
|  |  | 1499.9763 | 0.015744
|  |  | 1499.9766 | 0.016202
|  |  | 1499.977 | 0.016673
|  |  | 1499.9773 | 0.017157
|  |  | 1499.9776 | 0.017653
|  |  | 1499.9779 | 0.018162
|  |  | 1499.9782 | 0.018685
|  |  | 1499.9786 | 0.019222
|  |  | 1499.9789 | 0.019773
|  |  | 1499.9792 | 0.020338
|  |  | 1499.9795 | 0.020918
|  |  | 1499.9798 | 0.021514
|  |  | 1499.9802 | 0.022124
|  |  | 1499.9805 | 0.022751
|  |  | 1499.9808 | 0.023394
|  |  | 1499.9811 | 0.024054
|  |  | 1499.9814 | 0.024731
|  |  | 1499.9818 | 0.025425
|  |  | 1499.9821 | 0.026137
|  |  | 1499.9824 | 0.026867
|  |  | 1499.9827 | 0.027616
|  |  | 1499.983 | 0.028383
|  |  | 1499.9834 | 0.029171
|  |  | 1499.9837 | 0.029978
|  |  | 1499.984 | 0.030805
|  |  | 1499.9843 | 0.031654
|  |  | 1499.9846 | 0.032524
|  |  | 1499.985 | 0.033415
|  |  | 1499.9853 | 0.034329
|  |  | 1499.9856 | 0.035265
|  |  | 1499.9859 | 0.036225
|  |  | 1499.9862 | 0.037208
|  |  | 1499.9866 | 0.038216
|  |  | 1499.9869 | 0.039248
|  |  | 1499.9872 | 0.040306
|  |  | 1499.9875 | 0.041389
|  |  | 1499.9878 | 0.042499
|  |  | 1499.9882 | 0.043636
|  |  | 1499.9885 | 0.044801
|  |  | 1499.9888 | 0.045993
|  |  | 1499.9891 | 0.047214
|  |  | 1499.9894 | 0.048465
|  |  | 1499.9898 | 0.049746
|  |  | 1499.9901 | 0.051057
|  |  | 1499.9904 | 0.052399
|  |  | 1499.9907 | 0.053773
|  |  | 1499.991 | 0.05518
|  |  | 1499.9914 | 0.05662
|  |  | 1499.9917 | 0.058094
|  |  | 1499.992 | 0.059602
|  |  | 1499.9923 | 0.061146
|  |  | 1499.9926 | 0.062725
|  |  | 1499.993 | 0.064341
|  |  | 1499.9933 | 0.065995
|  |  | 1499.9936 | 0.067687
|  |  | 1499.9939 | 0.069418
|  |  | 1499.9942 | 0.071188
|  |  | 1499.9946 | 0.072999
|  |  | 1499.9949 | 0.074851
|  |  | 1499.9952 | 0.076745
|  |  | 1499.9955 | 0.078682
|  |  | 1499.9958 | 0.080663
|  |  | 1499.9962 | 0.082689
|  |  | 1499.9965 | 0.08476
|  |  | 1499.9968 | 0.086877
|  |  | 1499.9971 | 0.089041
|  |  | 1499.9974 | 0.091254
|  |  | 1499.9978 | 0.093515
|  |  | 1499.9981 | 0.095827
|  |  | 1499.9984 | 0.098189
|  |  | 1499.9987 | 0.1006
|  |  | 1499.999 | 0.10307
|  |  | 1499.9994 | 0.10559
|  |  | 1499.9997 | 0.10817
|  |  | 1500 | 0.1108
|  |  | 1500.0003 | 0.11348
|  |  | 1500.0006 | 0.11623
|  |  | 1500.001 | 0.11903
|  |  | 1500.0013 | 0.1219
|  |  | 1500.0016 | 0.12482
|  |  | 1500.0019 | 0.12781
|  |  | 1500.0022 | 0.13086
|  |  | 1500.0026 | 0.13397
|  |  | 1500.0029 | 0.13715
|  |  | 1500.0032 | 0.1404
|  |  | 1500.0035 | 0.14371
|  |  | 1500.0038 | 0.1471
|  |  | 1500.0042 | 0.15055
|  |  | 1500.0045 | 0.15407
|  |  | 1500.0048 | 0.15767
|  |  | 1500.0051 | 0.16134
|  |  | 1500.0054 | 0.16508
|  |  | 1500.0058 | 0.1689
|  |  | 1500.0061 | 0.1728
|  |  | 1500.0064 | 0.17678
|  |  | 1500.0067 | 0.18083
|  |  | 1500.007 | 0.18497
|  |  | 1500.0074 | 0.18919
|  |  | 1500.0077 | 0.1935
|  |  | 1500.008 | 0.19789
|  |  | 1500.0083 | 0.20236
|  |  | 1500.0086 | 0.20693
|  |  | 1500.009 | 0.21158
|  |  | 1500.0093 | 0.21632
|  |  | 1500.0096 | 0.22116
|  |  | 1500.0099 | 0.22609
|  |  | 1500.0102 | 0.23112
|  |  | 1500.0106 | 0.23624
|  |  | 1500.0109 | 0.24146
|  |  | 1500.0112 | 0.24678
|  |  | 1500.0115 | 0.2522
|  |  | 1500.0118 | 0.25772
|  |  | 1500.0122 | 0.26335
|  |  | 1500.0125 | 0.26908
|  |  | 1500.0128 | 0.27492
|  |  | 1500.0131 | 0.28087
|  |  | 1500.0134 | 0.28693
|  |  | 1500.0138 | 0.2931
|  |  | 1500.0141 | 0.29939
|  |  | 1500.0144 | 0.30579
|  |  | 1500.0147 | 0.31231
|  |  | 1500.015 | 0.31894
|  |  | 1500.0154 | 0.3257
|  |  | 1500.0157 | 0.33257
|  |  | 1500.016 | 0.33957
|  |  | 1500.0163 | 0.3467
|  |  | 1500.0166 | 0.35395
|  |  | 1500.017 | 0.36133
|  |  | 1500.0173 | 0.36885
|  |  | 1500.0176 | 0.37649
|  |  | 1500.0179 | 0.38427
|  |  | 1500.0182 | 0.39218
|  |  | 1500.0186 | 0.40023
|  |  | 1500.0189 | 0.40842
|  |  | 1500.0192 | 0.41675
|  |  | 1500.0195 | 0.42523
|  |  | 1500.0198 | 0.43384
|  |  | 1500.0202 | 0.44261
|  |  | 1500.0205 | 0.45152
|  |  | 1500.0208 | 0.46058
|  |  | 1500.0211 | 0.4698
|  |  | 1500.0214 | 0.47916
|  |  | 1500.0218 | 0.48869
|  |  | 1500.0221 | 0.49837
|  |  | 1500.0224 | 0.50821
|  |  | 1500.0227 | 0.51821
|  |  | 1500.023 | 0.52837
|  |  | 1500.0234 | 0.5387
|  |  | 1500.0237 | 0.5492
|  |  | 1500.024 | 0.55986
|  |  | 1500.0243 | 0.5707
|  |  | 1500.0246 | 0.58171
|  |  | 1500.025 | 0.59289
|  |  | 1500.0253 | 0.60425
|  |  | 1500.0256 | 0.61578
|  |  | 1500.0259 | 0.6275
|  |  | 1500.0262 | 0.6394
|  |  | 1500.0266 | 0.65148
|  |  | 1500.0269 | 0.66375
|  |  | 1500.0272 | 0.6762
|  |  | 1500.0275 | 0.68885
|  |  | 1500.0278 | 0.70168
|  |  | 1500.0282 | 0.71472
|  |  | 1500.0285 | 0.72794
|  |  | 1500.0288 | 0.74136
|  |  | 1500.0291 | 0.75499
|  |  | 1500.0294 | 0.76881
|  |  | 1500.0298 | 0.78284
|  |  | 1500.0301 | 0.79707
|  |  | 1500.0304 | 0.81151
|  |  | 1500.0307 | 0.82615
|  |  | 1500.031 | 0.84101
|  |  | 1500.0314 | 0.85608
|  |  | 1500.0317 | 0.87136
|  |  | 1500.032 | 0.88686
|  |  | 1500.0323 | 0.90258
|  |  | 1500.0326 | 0.91852
|  |  | 1500.033 | 0.93468
|  |  | 1500.0333 | 0.95106
|  |  | 1500.0336 | 0.96767
|  |  | 1500.0339 | 0.98451
|  |  | 1500.0342 | 1.0016
|  |  | 1500.0346 | 1.0189
|  |  | 1500.0349 | 1.0364
|  |  | 1500.0352 | 1.0542
|  |  | 1500.0355 | 1.0721
|  |  | 1500.0358 | 1.0904
|  |  | 1500.0362 | 1.1089
|  |  | 1500.0365 | 1.1276
|  |  | 1500.0368 | 1.1465
|  |  | 1500.0371 | 1.1657
|  |  | 1500.0374 | 1.1852
|  |  | 1500.0378 | 1.2049
|  |  | 1500.0381 | 1.2248
|  |  | 1500.0384 | 1.245
|  |  | 1500.0387 | 1.2654
|  |  | 1500.039 | 1.2861
|  |  | 1500.0394 | 1.3071
|  |  | 1500.0397 | 1.3283
|  |  | 1500.04 | 1.3498
|  |  | 1500.0403 | 1.3715
|  |  | 1500.0406 | 1.3935
|  |  | 1500.041 | 1.4157
|  |  | 1500.0413 | 1.4382
|  |  | 1500.0416 | 1.461
|  |  | 1500.0419 | 1.4841
|  |  | 1500.0422 | 1.5074
|  |  | 1500.0426 | 1.531
|  |  | 1500.0429 | 1.5548
|  |  | 1500.0432 | 1.5789
|  |  | 1500.0435 | 1.6033
|  |  | 1500.0438 | 1.628
|  |  | 1500.0442 | 1.6529
|  |  | 1500.0445 | 1.6781
|  |  | 1500.0448 | 1.7036
|  |  | 1500.0451 | 1.7293
|  |  | 1500.0454 | 1.7554
|  |  | 1500.0458 | 1.7817
|  |  | 1500.0461 | 1.8083
|  |  | 1500.0464 | 1.8352
|  |  | 1500.0467 | 1.8623
|  |  | 1500.047 | 1.8898
|  |  | 1500.0474 | 1.9175
|  |  | 1500.0477 | 1.9455
|  |  | 1500.048 | 1.9738
|  |  | 1500.0483 | 2.0023
|  |  | 1500.0486 | 2.0312
|  |  | 1500.049 | 2.0603
|  |  | 1500.0493 | 2.0897
|  |  | 1500.0496 | 2.1194
|  |  | 1500.0499 | 2.1494
|  |  | 1500.0502 | 2.1797
|  |  | 1500.0506 | 2.2102
|  |  | 1500.0509 | 2.2411
|  |  | 1500.0512 | 2.2722
|  |  | 1500.0515 | 2.3036
|  |  | 1500.0518 | 2.3353
|  |  | 1500.0522 | 2.3673
|  |  | 1500.0525 | 2.3995
|  |  | 1500.0528 | 2.4321
|  |  | 1500.0531 | 2.4649
|  |  | 1500.0534 | 2.498
|  |  | 1500.0538 | 2.5314
|  |  | 1500.0541 | 2.5651
|  |  | 1500.0544 | 2.599
|  |  | 1500.0547 | 2.6333
|  |  | 1500.055 | 2.6678
|  |  | 1500.0554 | 2.7026
|  |  | 1500.0557 | 2.7377
|  |  | 1500.056 | 2.773
|  |  | 1500.0563 | 2.8087
|  |  | 1500.0566 | 2.8446
|  |  | 1500.057 | 2.8807
|  |  | 1500.0573 | 2.9172
|  |  | 1500.0576 | 2.9539
|  |  | 1500.0579 | 2.9909
|  |  | 1500.0582 | 3.0282
|  |  | 1500.0586 | 3.0657
|  |  | 1500.0589 | 3.1035
|  |  | 1500.0592 | 3.1416
|  |  | 1500.0595 | 3.1799
|  |  | 1500.0598 | 3.2185
|  |  | 1500.0602 | 3.2574
|  |  | 1500.0605 | 3.2965
|  |  | 1500.0608 | 3.3359
|  |  | 1500.0611 | 3.3755
|  |  | 1500.0614 | 3.4154
|  |  | 1500.0618 | 3.4555
|  |  | 1500.0621 | 3.4959
|  |  | 1500.0624 | 3.5365
|  |  | 1500.0627 | 3.5774
|  |  | 1500.063 | 3.6185
|  |  | 1500.0634 | 3.6598
|  |  | 1500.0637 | 3.7014
|  |  | 1500.064 | 3.7432
|  |  | 1500.0643 | 3.7852
|  |  | 1500.0646 | 3.8275
|  |  | 1500.065 | 3.87
|  |  | 1500.0653 | 3.9127
|  |  | 1500.0656 | 3.9556
|  |  | 1500.0659 | 3.9988
|  |  | 1500.0662 | 4.0421
|  |  | 1500.0666 | 4.0857
|  |  | 1500.0669 | 4.1295
|  |  | 1500.0672 | 4.1734
|  |  | 1500.0675 | 4.2176
|  |  | 1500.0678 | 4.262
|  |  | 1500.0682 | 4.3065
|  |  | 1500.0685 | 4.3513
|  |  | 1500.0688 | 4.3962
|  |  | 1500.0691 | 4.4413
|  |  | 1500.0694 | 4.4866
|  |  | 1500.0698 | 4.532
|  |  | 1500.0701 | 4.5776
|  |  | 1500.0704 | 4.6234
|  |  | 1500.0707 | 4.6694
|  |  | 1500.071 | 4.7155
|  |  | 1500.0714 | 4.7617
|  |  | 1500.0717 | 4.8081
|  |  | 1500.072 | 4.8547
|  |  | 1500.0723 | 4.9013
|  |  | 1500.0726 | 4.9481
|  |  | 1500.073 | 4.9951
|  |  | 1500.0733 | 5.0421
|  |  | 1500.0736 | 5.0893
|  |  | 1500.0739 | 5.1366
|  |  | 1500.0742 | 5.184
|  |  | 1500.0746 | 5.2315
|  |  | 1500.0749 | 5.279
|  |  | 1500.0752 | 5.3267
|  |  | 1500.0755 | 5.3745
|  |  | 1500.0758 | 5.4223
|  |  | 1500.0762 | 5.4703
|  |  | 1500.0765 | 5.5183
|  |  | 1500.0768 | 5.5663
|  |  | 1500.0771 | 5.6145
|  |  | 1500.0774 | 5.6626
|  |  | 1500.0778 | 5.7109
|  |  | 1500.0781 | 5.7591
|  |  | 1500.0784 | 5.8074
|  |  | 1500.0787 | 5.8558
|  |  | 1500.079 | 5.9041
|  |  | 1500.0794 | 5.9525
|  |  | 1500.0797 | 6.0009
|  |  | 1500.08 | 6.0493
|  |  | 1500.0803 | 6.0977
|  |  | 1500.0806 | 6.146
|  |  | 1500.081 | 6.1944
|  |  | 1500.0813 | 6.2428
|  |  | 1500.0816 | 6.2911
|  |  | 1500.0819 | 6.3394
|  |  | 1500.0822 | 6.3877
|  |  | 1500.0826 | 6.4359
|  |  | 1500.0829 | 6.484
|  |  | 1500.0832 | 6.5322
|  |  | 1500.0835 | 6.5802
|  |  | 1500.0838 | 6.6282
|  |  | 1500.0842 | 6.6761
|  |  | 1500.0845 | 6.7239
|  |  | 1500.0848 | 6.7716
|  |  | 1500.0851 | 6.8192
|  |  | 1500.0854 | 6.8667
|  |  | 1500.0858 | 6.9141
|  |  | 1500.0861 | 6.9614
|  |  | 1500.0864 | 7.0086
|  |  | 1500.0867 | 7.0556
|  |  | 1500.087 | 7.1025
|  |  | 1500.0874 | 7.1493
|  |  | 1500.0877 | 7.1959
|  |  | 1500.088 | 7.2423
|  |  | 1500.0883 | 7.2886
|  |  | 1500.0886 | 7.3346
|  |  | 1500.089 | 7.3806
|  |  | 1500.0893 | 7.4263
|  |  | 1500.0896 | 7.4718
|  |  | 1500.0899 | 7.5171
|  |  | 1500.0902 | 7.5623
|  |  | 1500.0906 | 7.6072
|  |  | 1500.0909 | 7.6518
|  |  | 1500.0912 | 7.6963
|  |  | 1500.0915 | 7.7405
|  |  | 1500.0918 | 7.7845
|  |  | 1500.0922 | 7.8282
|  |  | 1500.0925 | 7.8716
|  |  | 1500.0928 | 7.9148
|  |  | 1500.0931 | 7.9577
|  |  | 1500.0934 | 8.0004
|  |  | 1500.0938 | 8.0427
|  |  | 1500.0941 | 8.0848
|  |  | 1500.0944 | 8.1266
|  |  | 1500.0947 | 8.168
|  |  | 1500.095 | 8.2091
|  |  | 1500.0954 | 8.25
|  |  | 1500.0957 | 8.2905
|  |  | 1500.096 | 8.3306
|  |  | 1500.0963 | 8.3704
|  |  | 1500.0966 | 8.4099
|  |  | 1500.097 | 8.449
|  |  | 1500.0973 | 8.4878
|  |  | 1500.0976 | 8.5261
|  |  | 1500.0979 | 8.5642
|  |  | 1500.0982 | 8.6018
|  |  | 1500.0986 | 8.639
|  |  | 1500.0989 | 8.6759
|  |  | 1500.0992 | 8.7123
|  |  | 1500.0995 | 8.7484
|  |  | 1500.0998 | 8.784
|  |  | 1500.1002 | 8.8192
|  |  | 1500.1005 | 8.854
|  |  | 1500.1008 | 8.8883
|  |  | 1500.1011 | 8.9222
|  |  | 1500.1014 | 8.9557
|  |  | 1500.1018 | 8.9887
|  |  | 1500.1021 | 9.0213
|  |  | 1500.1024 | 9.0534
|  |  | 1500.1027 | 9.085
|  |  | 1500.103 | 9.1162
|  |  | 1500.1034 | 9.1468
|  |  | 1500.1037 | 9.177
|  |  | 1500.104 | 9.2068
|  |  | 1500.1043 | 9.236
|  |  | 1500.1046 | 9.2647
|  |  | 1500.105 | 9.2929
|  |  | 1500.1053 | 9.3206
|  |  | 1500.1056 | 9.3478
|  |  | 1500.1059 | 9.3744
|  |  | 1500.1062 | 9.4006
|  |  | 1500.1066 | 9.4262
|  |  | 1500.1069 | 9.4512
|  |  | 1500.1072 | 9.4758
|  |  | 1500.1075 | 9.4997
|  |  | 1500.1078 | 9.5232
|  |  | 1500.1082 | 9.5461
|  |  | 1500.1085 | 9.5684
|  |  | 1500.1088 | 9.5902
|  |  | 1500.1091 | 9.6114
|  |  | 1500.1094 | 9.632
|  |  | 1500.1098 | 9.652
|  |  | 1500.1101 | 9.6715
|  |  | 1500.1104 | 9.6904
|  |  | 1500.1107 | 9.7087
|  |  | 1500.111 | 9.7265
|  |  | 1500.1114 | 9.7436
|  |  | 1500.1117 | 9.7601
|  |  | 1500.112 | 9.7761
|  |  | 1500.1123 | 9.7914
|  |  | 1500.1126 | 9.8061
|  |  | 1500.113 | 9.8203
|  |  | 1500.1133 | 9.8338
|  |  | 1500.1136 | 9.8467
|  |  | 1500.1139 | 9.859
|  |  | 1500.1142 | 9.8707
|  |  | 1500.1146 | 9.8817
|  |  | 1500.1149 | 9.8922
|  |  | 1500.1152 | 9.902
|  |  | 1500.1155 | 9.9112
|  |  | 1500.1158 | 9.9198
|  |  | 1500.1162 | 9.9277
|  |  | 1500.1165 | 9.935
|  |  | 1500.1168 | 9.9417
|  |  | 1500.1171 | 9.9477
|  |  | 1500.1174 | 9.9532
|  |  | 1500.1178 | 9.9579
|  |  | 1500.1181 | 9.9621
|  |  | 1500.1184 | 9.9656
|  |  | 1500.1187 | 9.9685
|  |  | 1500.119 | 9.9707
|  |  | 1500.1194 | 9.9723
|  |  | 1500.1197 | 9.9732
|  |  | 1500.12 | 9.9736
|  |  | 1500.1203 | 9.9732
|  |  | 1500.1206 | 9.9723
|  |  | 1500.121 | 9.9707
|  |  | 1500.1213 | 9.9685
|  |  | 1500.1216 | 9.9656
|  |  | 1500.1219 | 9.9621
|  |  | 1500.1222 | 9.9579
|  |  | 1500.1226 | 9.9532
|  |  | 1500.1229 | 9.9477
|  |  | 1500.1232 | 9.9417
|  |  | 1500.1235 | 9.935
|  |  | 1500.1238 | 9.9277
|  |  | 1500.1242 | 9.9198
|  |  | 1500.1245 | 9.9112
|  |  | 1500.1248 | 9.902
|  |  | 1500.1251 | 9.8922
|  |  | 1500.1254 | 9.8817
|  |  | 1500.1258 | 9.8707
|  |  | 1500.1261 | 9.859
|  |  | 1500.1264 | 9.8467
|  |  | 1500.1267 | 9.8338
|  |  | 1500.127 | 9.8203
|  |  | 1500.1274 | 9.8061
|  |  | 1500.1277 | 9.7914
|  |  | 1500.128 | 9.7761
|  |  | 1500.1283 | 9.7601
|  |  | 1500.1286 | 9.7436
|  |  | 1500.129 | 9.7265
|  |  | 1500.1293 | 9.7087
|  |  | 1500.1296 | 9.6904
|  |  | 1500.1299 | 9.6715
|  |  | 1500.1302 | 9.652
|  |  | 1500.1306 | 9.632
|  |  | 1500.1309 | 9.6114
|  |  | 1500.1312 | 9.5902
|  |  | 1500.1315 | 9.5684
|  |  | 1500.1318 | 9.5461
|  |  | 1500.1322 | 9.5232
|  |  | 1500.1325 | 9.4997
|  |  | 1500.1328 | 9.4758
|  |  | 1500.1331 | 9.4512
|  |  | 1500.1334 | 9.4262
|  |  | 1500.1338 | 9.4006
|  |  | 1500.1341 | 9.3744
|  |  | 1500.1344 | 9.3478
|  |  | 1500.1347 | 9.3206
|  |  | 1500.135 | 9.2929
|  |  | 1500.1354 | 9.2647
|  |  | 1500.1357 | 9.236
|  |  | 1500.136 | 9.2068
|  |  | 1500.1363 | 9.177
|  |  | 1500.1366 | 9.1468
|  |  | 1500.137 | 9.1162
|  |  | 1500.1373 | 9.085
|  |  | 1500.1376 | 9.0534
|  |  | 1500.1379 | 9.0213
|  |  | 1500.1382 | 8.9887
|  |  | 1500.1386 | 8.9557
|  |  | 1500.1389 | 8.9222
|  |  | 1500.1392 | 8.8883
|  |  | 1500.1395 | 8.854
|  |  | 1500.1398 | 8.8192
|  |  | 1500.1402 | 8.784
|  |  | 1500.1405 | 8.7484
|  |  | 1500.1408 | 8.7123
|  |  | 1500.1411 | 8.6759
|  |  | 1500.1414 | 8.639
|  |  | 1500.1418 | 8.6018
|  |  | 1500.1421 | 8.5642
|  |  | 1500.1424 | 8.5261
|  |  | 1500.1427 | 8.4878
|  |  | 1500.143 | 8.449
|  |  | 1500.1434 | 8.4099
|  |  | 1500.1437 | 8.3704
|  |  | 1500.144 | 8.3306
|  |  | 1500.1443 | 8.2905
|  |  | 1500.1446 | 8.25
|  |  | 1500.145 | 8.2091
|  |  | 1500.1453 | 8.168
|  |  | 1500.1456 | 8.1266
|  |  | 1500.1459 | 8.0848
|  |  | 1500.1462 | 8.0427
|  |  | 1500.1466 | 8.0004
|  |  | 1500.1469 | 7.9577
|  |  | 1500.1472 | 7.9148
|  |  | 1500.1475 | 7.8716
|  |  | 1500.1478 | 7.8282
|  |  | 1500.1482 | 7.7845
|  |  | 1500.1485 | 7.7405
|  |  | 1500.1488 | 7.6963
|  |  | 1500.1491 | 7.6518
|  |  | 1500.1494 | 7.6072
|  |  | 1500.1498 | 7.5623
|  |  | 1500.1501 | 7.5171
|  |  | 1500.1504 | 7.4718
|  |  | 1500.1507 | 7.4263
|  |  | 1500.151 | 7.3806
|  |  | 1500.1514 | 7.3346
|  |  | 1500.1517 | 7.2886
|  |  | 1500.152 | 7.2423
|  |  | 1500.1523 | 7.1959
|  |  | 1500.1526 | 7.1493
|  |  | 1500.153 | 7.1025
|  |  | 1500.1533 | 7.0556
|  |  | 1500.1536 | 7.0086
|  |  | 1500.1539 | 6.9614
|  |  | 1500.1542 | 6.9141
|  |  | 1500.1546 | 6.8667
|  |  | 1500.1549 | 6.8192
|  |  | 1500.1552 | 6.7716
|  |  | 1500.1555 | 6.7239
|  |  | 1500.1558 | 6.6761
|  |  | 1500.1562 | 6.6282
|  |  | 1500.1565 | 6.5802
|  |  | 1500.1568 | 6.5322
|  |  | 1500.1571 | 6.484
|  |  | 1500.1574 | 6.4359
|  |  | 1500.1578 | 6.3877
|  |  | 1500.1581 | 6.3394
|  |  | 1500.1584 | 6.2911
|  |  | 1500.1587 | 6.2428
|  |  | 1500.159 | 6.1944
|  |  | 1500.1594 | 6.146
|  |  | 1500.1597 | 6.0977
|  |  | 1500.16 | 6.0493
|  |  | 1500.1603 | 6.0009
|  |  | 1500.1606 | 5.9525
|  |  | 1500.161 | 5.9041
|  |  | 1500.1613 | 5.8558
|  |  | 1500.1616 | 5.8074
|  |  | 1500.1619 | 5.7591
|  |  | 1500.1622 | 5.7109
|  |  | 1500.1626 | 5.6626
|  |  | 1500.1629 | 5.6145
|  |  | 1500.1632 | 5.5663
|  |  | 1500.1635 | 5.5183
|  |  | 1500.1638 | 5.4703
|  |  | 1500.1642 | 5.4223
|  |  | 1500.1645 | 5.3745
|  |  | 1500.1648 | 5.3267
|  |  | 1500.1651 | 5.279
|  |  | 1500.1654 | 5.2315
|  |  | 1500.1658 | 5.184
|  |  | 1500.1661 | 5.1366
|  |  | 1500.1664 | 5.0893
|  |  | 1500.1667 | 5.0421
|  |  | 1500.167 | 4.9951
|  |  | 1500.1674 | 4.9481
|  |  | 1500.1677 | 4.9013
|  |  | 1500.168 | 4.8547
|  |  | 1500.1683 | 4.8081
|  |  | 1500.1686 | 4.7617
|  |  | 1500.169 | 4.7155
|  |  | 1500.1693 | 4.6694
|  |  | 1500.1696 | 4.6234
|  |  | 1500.1699 | 4.5776
|  |  | 1500.1702 | 4.532
|  |  | 1500.1706 | 4.4866
|  |  | 1500.1709 | 4.4413
|  |  | 1500.1712 | 4.3962
|  |  | 1500.1715 | 4.3513
|  |  | 1500.1718 | 4.3065
|  |  | 1500.1722 | 4.262
|  |  | 1500.1725 | 4.2176
|  |  | 1500.1728 | 4.1734
|  |  | 1500.1731 | 4.1295
|  |  | 1500.1734 | 4.0857
|  |  | 1500.1738 | 4.0421
|  |  | 1500.1741 | 3.9988
|  |  | 1500.1744 | 3.9556
|  |  | 1500.1747 | 3.9127
|  |  | 1500.175 | 3.87
|  |  | 1500.1754 | 3.8275
|  |  | 1500.1757 | 3.7852
|  |  | 1500.176 | 3.7432
|  |  | 1500.1763 | 3.7014
|  |  | 1500.1766 | 3.6598
|  |  | 1500.177 | 3.6185
|  |  | 1500.1773 | 3.5774
|  |  | 1500.1776 | 3.5365
|  |  | 1500.1779 | 3.4959
|  |  | 1500.1782 | 3.4555
|  |  | 1500.1786 | 3.4154
|  |  | 1500.1789 | 3.3755
|  |  | 1500.1792 | 3.3359
|  |  | 1500.1795 | 3.2965
|  |  | 1500.1798 | 3.2574
|  |  | 1500.1802 | 3.2185
|  |  | 1500.1805 | 3.1799
|  |  | 1500.1808 | 3.1416
|  |  | 1500.1811 | 3.1035
|  |  | 1500.1814 | 3.0657
|  |  | 1500.1818 | 3.0282
|  |  | 1500.1821 | 2.9909
|  |  | 1500.1824 | 2.9539
|  |  | 1500.1827 | 2.9172
|  |  | 1500.183 | 2.8807
|  |  | 1500.1834 | 2.8446
|  |  | 1500.1837 | 2.8087
|  |  | 1500.184 | 2.773
|  |  | 1500.1843 | 2.7377
|  |  | 1500.1846 | 2.7026
|  |  | 1500.185 | 2.6678
|  |  | 1500.1853 | 2.6333
|  |  | 1500.1856 | 2.599
|  |  | 1500.1859 | 2.5651
|  |  | 1500.1862 | 2.5314
|  |  | 1500.1866 | 2.498
|  |  | 1500.1869 | 2.4649
|  |  | 1500.1872 | 2.4321
|  |  | 1500.1875 | 2.3995
|  |  | 1500.1878 | 2.3673
|  |  | 1500.1882 | 2.3353
|  |  | 1500.1885 | 2.3036
|  |  | 1500.1888 | 2.2722
|  |  | 1500.1891 | 2.2411
|  |  | 1500.1894 | 2.2102
|  |  | 1500.1898 | 2.1797
|  |  | 1500.1901 | 2.1494
|  |  | 1500.1904 | 2.1194
|  |  | 1500.1907 | 2.0897
|  |  | 1500.191 | 2.0603
|  |  | 1500.1914 | 2.0312
|  |  | 1500.1917 | 2.0023
|  |  | 1500.192 | 1.9738
|  |  | 1500.1923 | 1.9455
|  |  | 1500.1926 | 1.9175
|  |  | 1500.193 | 1.8898
|  |  | 1500.1933 | 1.8623
|  |  | 1500.1936 | 1.8352
|  |  | 1500.1939 | 1.8083
|  |  | 1500.1942 | 1.7817
|  |  | 1500.1946 | 1.7554
|  |  | 1500.1949 | 1.7293
|  |  | 1500.1952 | 1.7036
|  |  | 1500.1955 | 1.6781
|  |  | 1500.1958 | 1.6529
|  |  | 1500.1962 | 1.628
|  |  | 1500.1965 | 1.6033
|  |  | 1500.1968 | 1.5789
|  |  | 1500.1971 | 1.5548
|  |  | 1500.1974 | 1.531
|  |  | 1500.1978 | 1.5074
|  |  | 1500.1981 | 1.4841
|  |  | 1500.1984 | 1.461
|  |  | 1500.1987 | 1.4382
|  |  | 1500.199 | 1.4157
|  |  | 1500.1994 | 1.3935
|  |  | 1500.1997 | 1.3715
|  |  | 1500.2 | 1.3498
|  |  | 1500.2003 | 1.3283
|  |  | 1500.2006 | 1.3071
|  |  | 1500.201 | 1.2861
|  |  | 1500.2013 | 1.2654
|  |  | 1500.2016 | 1.245
|  |  | 1500.2019 | 1.2248
|  |  | 1500.2022 | 1.2049
|  |  | 1500.2026 | 1.1852
|  |  | 1500.2029 | 1.1657
|  |  | 1500.2032 | 1.1465
|  |  | 1500.2035 | 1.1276
|  |  | 1500.2038 | 1.1089
|  |  | 1500.2042 | 1.0904
|  |  | 1500.2045 | 1.0721
|  |  | 1500.2048 | 1.0542
|  |  | 1500.2051 | 1.0364
|  |  | 1500.2054 | 1.0189
|  |  | 1500.2058 | 1.0016
|  |  | 1500.2061 | 0.98451
|  |  | 1500.2064 | 0.96767
|  |  | 1500.2067 | 0.95106
|  |  | 1500.207 | 0.93468
|  |  | 1500.2074 | 0.91852
|  |  | 1500.2077 | 0.90258
|  |  | 1500.208 | 0.88686
|  |  | 1500.2083 | 0.87136
|  |  | 1500.2086 | 0.85608
|  |  | 1500.209 | 0.84101
|  |  | 1500.2093 | 0.82615
|  |  | 1500.2096 | 0.81151
|  |  | 1500.2099 | 0.79707
|  |  | 1500.2102 | 0.78284
|  |  | 1500.2106 | 0.76881
|  |  | 1500.2109 | 0.75499
|  |  | 1500.2112 | 0.74136
|  |  | 1500.2115 | 0.72794
|  |  | 1500.2118 | 0.71472
|  |  | 1500.2122 | 0.70168
|  |  | 1500.2125 | 0.68885
|  |  | 1500.2128 | 0.6762
|  |  | 1500.2131 | 0.66375
|  |  | 1500.2134 | 0.65148
|  |  | 1500.2138 | 0.6394
|  |  | 1500.2141 | 0.6275
|  |  | 1500.2144 | 0.61578
|  |  | 1500.2147 | 0.60425
|  |  | 1500.215 | 0.59289
|  |  | 1500.2154 | 0.58171
|  |  | 1500.2157 | 0.5707
|  |  | 1500.216 | 0.55986
|  |  | 1500.2163 | 0.5492
|  |  | 1500.2166 | 0.5387
|  |  | 1500.217 | 0.52837
|  |  | 1500.2173 | 0.51821
|  |  | 1500.2176 | 0.50821
|  |  | 1500.2179 | 0.49837
|  |  | 1500.2182 | 0.48869
|  |  | 1500.2186 | 0.47916
|  |  | 1500.2189 | 0.4698
|  |  | 1500.2192 | 0.46058
|  |  | 1500.2195 | 0.45152
|  |  | 1500.2198 | 0.44261
|  |  | 1500.2202 | 0.43384
|  |  | 1500.2205 | 0.42523
|  |  | 1500.2208 | 0.41675
|  |  | 1500.2211 | 0.40842
|  |  | 1500.2214 | 0.40023
|  |  | 1500.2218 | 0.39218
|  |  | 1500.2221 | 0.38427
|  |  | 1500.2224 | 0.37649
|  |  | 1500.2227 | 0.36885
|  |  | 1500.223 | 0.36133
|  |  | 1500.2234 | 0.35395
|  |  | 1500.2237 | 0.3467
|  |  | 1500.224 | 0.33957
|  |  | 1500.2243 | 0.33257
|  |  | 1500.2246 | 0.3257
|  |  | 1500.225 | 0.31894
|  |  | 1500.2253 | 0.31231
|  |  | 1500.2256 | 0.30579
|  |  | 1500.2259 | 0.29939
|  |  | 1500.2262 | 0.2931
|  |  | 1500.2266 | 0.28693
|  |  | 1500.2269 | 0.28087
|  |  | 1500.2272 | 0.27492
|  |  | 1500.2275 | 0.26908
|  |  | 1500.2278 | 0.26335
|  |  | 1500.2282 | 0.25772
|  |  | 1500.2285 | 0.2522
|  |  | 1500.2288 | 0.24678
|  |  | 1500.2291 | 0.24146
|  |  | 1500.2294 | 0.23624
|  |  | 1500.2298 | 0.23112
|  |  | 1500.2301 | 0.22609
|  |  | 1500.2304 | 0.22116
|  |  | 1500.2307 | 0.21632
|  |  | 1500.231 | 0.21158
|  |  | 1500.2314 | 0.20693
|  |  | 1500.2317 | 0.20236
|  |  | 1500.232 | 0.19789
|  |  | 1500.2323 | 0.1935
|  |  | 1500.2326 | 0.18919
|  |  | 1500.233 | 0.18497
|  |  | 1500.2333 | 0.18083
|  |  | 1500.2336 | 0.17678
|  |  | 1500.2339 | 0.1728
|  |  | 1500.2342 | 0.1689
|  |  | 1500.2346 | 0.16508
|  |  | 1500.2349 | 0.16134
|  |  | 1500.2352 | 0.15767
|  |  | 1500.2355 | 0.15407
|  |  | 1500.2358 | 0.15055
|  |  | 1500.2362 | 0.1471
|  |  | 1500.2365 | 0.14371
|  |  | 1500.2368 | 0.1404
|  |  | 1500.2371 | 0.13715
|  |  | 1500.2374 | 0.13397
|  |  | 1500.2378 | 0.13086
|  |  | 1500.2381 | 0.12781
|  |  | 1500.2384 | 0.12482
|  |  | 1500.2387 | 0.1219
|  |  | 1500.239 | 0.11903
|  |  | 1500.2394 | 0.11623
|  |  | 1500.2397 | 0.11348
|  |  | 1500.24 | 0.1108
|  |  | 1500.2403 | 0.10817
|  |  | 1500.2406 | 0.10559
|  |  | 1500.241 | 0.10307
|  |  | 1500.2413 | 0.1006
|  |  | 1500.2416 | 0.098189
|  |  | 1500.2419 | 0.095827
|  |  | 1500.2422 | 0.093515
|  |  | 1500.2426 | 0.091254
|  |  | 1500.2429 | 0.089041
|  |  | 1500.2432 | 0.086877
|  |  | 1500.2435 | 0.08476
|  |  | 1500.2438 | 0.082689
|  |  | 1500.2442 | 0.080663
|  |  | 1500.2445 | 0.078682
|  |  | 1500.2448 | 0.076745
|  |  | 1500.2451 | 0.074851
|  |  | 1500.2454 | 0.072999
|  |  | 1500.2458 | 0.071188
|  |  | 1500.2461 | 0.069418
|  |  | 1500.2464 | 0.067687
|  |  | 1500.2467 | 0.065995
|  |  | 1500.247 | 0.064341
|  |  | 1500.2474 | 0.062725
|  |  | 1500.2477 | 0.061146
|  |  | 1500.248 | 0.059602
|  |  | 1500.2483 | 0.058094
|  |  | 1500.2486 | 0.05662
|  |  | 1500.249 | 0.05518
|  |  | 1500.2493 | 0.053773
|  |  | 1500.2496 | 0.052399
|  |  | 1500.2499 | 0.051057
|  |  | 1500.2502 | 0.049746
|  |  | 1500.2506 | 0.048465
|  |  | 1500.2509 | 0.047214
|  |  | 1500.2512 | 0.045993
|  |  | 1500.2515 | 0.044801
|  |  | 1500.2518 | 0.043636
|  |  | 1500.2522 | 0.042499
|  |  | 1500.2525 | 0.041389
|  |  | 1500.2528 | 0.040306
|  |  | 1500.2531 | 0.039248
|  |  | 1500.2534 | 0.038216
|  |  | 1500.2538 | 0.037208
|  |  | 1500.2541 | 0.036225
|  |  | 1500.2544 | 0.035265
|  |  | 1500.2547 | 0.034329
|  |  | 1500.255 | 0.033415
|  |  | 1500.2554 | 0.032524
|  |  | 1500.2557 | 0.031654
|  |  | 1500.256 | 0.030805
|  |  | 1500.2563 | 0.029978
|  |  | 1500.2566 | 0.029171
|  |  | 1500.257 | 0.028383
|  |  | 1500.2573 | 0.027616
|  |  | 1500.2576 | 0.026867
|  |  | 1500.2579 | 0.026137
|  |  | 1500.2582 | 0.025425
|  |  | 1500.2586 | 0.024731
|  |  | 1500.2589 | 0.024054
|  |  | 1500.2592 | 0.023394
|  |  | 1500.2595 | 0.022751
|  |  | 1500.2598 | 0.022124
|  |  | 1500.2602 | 0.021514
|  |  | 1500.2605 | 0.020918
|  |  | 1500.2608 | 0.020338
|  |  | 1500.2611 | 0.019773
|  |  | 1500.2614 | 0.019222
|  |  | 1500.2618 | 0.018685
|  |  | 1500.2621 | 0.018162
|  |  | 1500.2624 | 0.017653
|  |  | 1500.2627 | 0.017157
|  |  | 1500.263 | 0.016673
|  |  | 1500.2634 | 0.016202
|  |  | 1500.2637 | 0.015744
|  |  | 1500.264 | 0.015298
|  |  | 1500.2643 | 0.014863
|  |  | 1500.2646 | 0.014439
|  |  | 1500.265 | 0.014027
|  |  | 1500.2653 | 0.013626
|  |  | 1500.2656 | 0.013235
|  |  | 1500.2659 | 0.012855
|  |  | 1500.2662 | 0.012485
|  |  | 1500.2666 | 0.012125
|  |  | 1500.2669 | 0.011774
|  |  | 1500.2672 | 0.011433
|  |  | 1500.2675 | 0.011101
|  |  | 1500.2678 | 0.010778
|  |  | 1500.2682 | 0.010463
|  |  | 1500.2685 | 0.010158
|  |  | 1500.2688 | 0.0098601
|  |  | 1500.2691 | 0.0095706
|  |  | 1500.2694 | 0.0092891
|  |  | 1500.2698 | 0.0090153
|  |  | 1500.2701 | 0.008749
|  |  | 1500.2704 | 0.00849
|  |  | 1500.2707 | 0.0082382
|  |  | 1500.271 | 0.0079933
|  |  | 1500.2714 | 0.0077552
|  |  | 1500.2717 | 0.0075237
|  |  | 1500.272 | 0.0072987
|  |  | 1500.2723 | 0.0070799
|  |  | 1500.2726 | 0.0068673
|  |  | 1500.273 | 0.0066606
|  |  | 1500.2733 | 0.0064597
|  |  | 1500.2736 | 0.0062645
|  |  | 1500.2739 | 0.0060747
|  |  | 1500.2742 | 0.0058904
|  |  | 1500.2746 | 0.0057113
|  |  | 1500.2749 | 0.0055373
|  |  | 1500.2752 | 0.0053682
|  |  | 1500.2755 | 0.005204
|  |  | 1500.2758 | 0.0050444
|  |  | 1500.2762 | 0.0048895
|  |  | 1500.2765 | 0.004739
|  |  | 1500.2768 | 0.0045928
|  |  | 1500.2771 | 0.0044509
|  |  | 1500.2774 | 0.004313
|  |  | 1500.2778 | 0.0041792
|  |  | 1500.2781 | 0.0040493
|  |  | 1500.2784 | 0.0039231
|  |  | 1500.2787 | 0.0038007
|  |  | 1500.279 | 0.0036818
|  |  | 1500.2794 | 0.0035664
|  |  | 1500.2797 | 0.0034544
|  |  | 1500.28 | 0.0033458
|  |  | 1500.2803 | 0.0032403
|  |  | 1500.2806 | 0.0031379
|  |  | 1500.281 | 0.0030386
|  |  | 1500.2813 | 0.0029423

## Guard Band Change MU
Risk Calculator
Upper Tolerance TU | 1500.2 | Conformance Probability Table | [8 empty] | W
Lower Tolerance TL | 1499.8 | k - Factor | Conformance Probability , Pc | Guard Band Multiplier, r Two Sided | [6 empty] | 0.01
Nominal Value | 1500 | [9 empty] | 0.05
Measurement Unc um | 0.04 | 0.083819 | 0.0668 | -0.75003 | [6 empty] | 0.1
Measured Value xm | 1500.12 | 0.20061 | 0.159 | -0.49929 | [6 empty] | 0.25
Tolerance T | 0.4 | 0.39682 | 0.3085 | -0.25005 | [6 empty] | 0.5
|  | 0.67449 | 0.5 | 0 | [6 empty] | 0.75
|  | 0.99982 | 0.6826 | 0.23749 | [6 empty] | 1
Probability of Conformance (pc) | 97.72% | 1.2816 | 0.8 | 0.42081 | [6 empty] | 1.25
Probability of NonConformance (1 - pc) | 2.275% | 1.4395 | 0.85 | 0.51822 | [6 empty] | 1.5
|  | 1.6449 | 0.9 | 0.64078 | [6 empty] | 1.75
Setting the Guard Band Upper and Lower pc |  | 1.96 | 0.95 | 0.82243 | [6 empty] | 2
Select Desired Conformance Probability | 97.72% | 2 | 0.9545 | 0.84507 | [6 empty] | 2.25
Risk if within GU & GL | 2.275% | 2.2776 | 0.97725 | 1 |  | AL GB Multiplier |  |  |  |  | 2.5
hU (GB Multiplier) | 1 | 2.3263 | 0.98 | 1.0269 | [6 empty] | 2.75
Guard Band Upper Gu | 1500.12 | 3 | 0.9973 | 1.3911 |  | 1 |  |  |  |  | 3
Guard Band Lower GL | 1499.88 | 3.2905 | 0.999 | 1.5451 | [6 empty] | 3.25
|  | 6.1094 | 1 | 2.9989 | [6 empty] | 3.5
Guard Band Based on MU Ratio | 1 | [9 empty] | 3.75
New MU Ratio 1:1 | 1 | [9 empty] | 4
um * (MU ratio) | 0.04 | [9 empty] | 4.25
Cm | 2.5 | [9 empty] | 4.5
4.75
5
5.25
5.5
5.75
6
[6 empty] | 0 |  |  |  | As-Found
[6 empty] | 0 | Nominal Value | LSL | USL | MV
[7 empty] | 1500 | 1499.8 | 1500.2 | 1500.12 | 0
[7 empty] | 1500 | 1499.8 | 1500.2 | 1500.12 | 9.9736
[8 empty] | LAL | UAL
[5 empty] | Increment | 0.00032 |  | 1499.88 | 1500.12 |  | 0
[8 empty] | 1499.88 | 1500.12 |  | 9.9736
Norm. Dist.
[5 empty] | x | y
[5 empty] | 1499.96 | 0.0033458
[5 empty] | 1499.9603 | 0.0034544
[5 empty] | 1499.9606 | 0.0035664
[5 empty] | 1499.961 | 0.0036818
[5 empty] | 1499.9613 | 0.0038007 | NORMAL DISTRIBUTION FUNCTION
[5 empty] | 1499.9616 | 0.0039231
[5 empty] | 1499.9619 | 0.0040493 | x = | x-axis value
[5 empty] | 1499.9622 | 0.0041792 | u = | Centerline of distribution
[5 empty] | 1499.9626 | 0.004313 | σ = | Std. Dev.
[5 empty] | 1499.9629 | 0.0044509
[5 empty] | 1499.9632 | 0.0045928
[5 empty] | 1499.9635 | 0.004739
[5 empty] | 1499.9638 | 0.0048895
[5 empty] | 1499.9642 | 0.0050444
[5 empty] | 1499.9645 | 0.005204
[5 empty] | 1499.9648 | 0.0053682
[5 empty] | 1499.9651 | 0.0055373
[5 empty] | 1499.9654 | 0.0057113
[5 empty] | 1499.9658 | 0.0058904
[5 empty] | 1499.9661 | 0.0060747
[5 empty] | 1499.9664 | 0.0062645
[5 empty] | 1499.9667 | 0.0064597
[5 empty] | 1499.967 | 0.0066606
[5 empty] | 1499.9674 | 0.0068673
[5 empty] | 1499.9677 | 0.0070799
[5 empty] | 1499.968 | 0.0072987
[5 empty] | 1499.9683 | 0.0075237
[5 empty] | 1499.9686 | 0.0077552
[5 empty] | 1499.969 | 0.0079933
[5 empty] | 1499.9693 | 0.0082382
[5 empty] | 1499.9696 | 0.00849
[5 empty] | 1499.9699 | 0.008749
[5 empty] | 1499.9702 | 0.0090153
[5 empty] | 1499.9706 | 0.0092891
[5 empty] | 1499.9709 | 0.0095706
[5 empty] | 1499.9712 | 0.0098601
[5 empty] | 1499.9715 | 0.010158
[5 empty] | 1499.9718 | 0.010463
[5 empty] | 1499.9722 | 0.010778
[5 empty] | 1499.9725 | 0.011101
[5 empty] | 1499.9728 | 0.011433
[5 empty] | 1499.9731 | 0.011774
[5 empty] | 1499.9734 | 0.012125
[5 empty] | 1499.9738 | 0.012485
[5 empty] | 1499.9741 | 0.012855
[5 empty] | 1499.9744 | 0.013235
[5 empty] | 1499.9747 | 0.013626
[5 empty] | 1499.975 | 0.014027
[5 empty] | 1499.9754 | 0.014439
[5 empty] | 1499.9757 | 0.014863
[5 empty] | 1499.976 | 0.015298
[5 empty] | 1499.9763 | 0.015744
[5 empty] | 1499.9766 | 0.016202
[5 empty] | 1499.977 | 0.016673
[5 empty] | 1499.9773 | 0.017157
[5 empty] | 1499.9776 | 0.017653
[5 empty] | 1499.9779 | 0.018162
[5 empty] | 1499.9782 | 0.018685
[5 empty] | 1499.9786 | 0.019222
[5 empty] | 1499.9789 | 0.019773
[5 empty] | 1499.9792 | 0.020338
[5 empty] | 1499.9795 | 0.020918
[5 empty] | 1499.9798 | 0.021514
[5 empty] | 1499.9802 | 0.022124
[5 empty] | 1499.9805 | 0.022751
[5 empty] | 1499.9808 | 0.023394
[5 empty] | 1499.9811 | 0.024054
[5 empty] | 1499.9814 | 0.024731
[5 empty] | 1499.9818 | 0.025425
[5 empty] | 1499.9821 | 0.026137
[5 empty] | 1499.9824 | 0.026867
[5 empty] | 1499.9827 | 0.027616
[5 empty] | 1499.983 | 0.028383
[5 empty] | 1499.9834 | 0.029171
[5 empty] | 1499.9837 | 0.029978
[5 empty] | 1499.984 | 0.030805
[5 empty] | 1499.9843 | 0.031654
[5 empty] | 1499.9846 | 0.032524
[5 empty] | 1499.985 | 0.033415
[5 empty] | 1499.9853 | 0.034329
[5 empty] | 1499.9856 | 0.035265
[5 empty] | 1499.9859 | 0.036225
[5 empty] | 1499.9862 | 0.037208
[5 empty] | 1499.9866 | 0.038216
[5 empty] | 1499.9869 | 0.039248
[5 empty] | 1499.9872 | 0.040306
[5 empty] | 1499.9875 | 0.041389
[5 empty] | 1499.9878 | 0.042499
[5 empty] | 1499.9882 | 0.043636
[5 empty] | 1499.9885 | 0.044801
[5 empty] | 1499.9888 | 0.045993
[5 empty] | 1499.9891 | 0.047214
[5 empty] | 1499.9894 | 0.048465
[5 empty] | 1499.9898 | 0.049746
[5 empty] | 1499.9901 | 0.051057
[5 empty] | 1499.9904 | 0.052399
[5 empty] | 1499.9907 | 0.053773
[5 empty] | 1499.991 | 0.05518
[5 empty] | 1499.9914 | 0.05662
[5 empty] | 1499.9917 | 0.058094
[5 empty] | 1499.992 | 0.059602
[5 empty] | 1499.9923 | 0.061146
[5 empty] | 1499.9926 | 0.062725
[5 empty] | 1499.993 | 0.064341
[5 empty] | 1499.9933 | 0.065995
[5 empty] | 1499.9936 | 0.067687
[5 empty] | 1499.9939 | 0.069418
[5 empty] | 1499.9942 | 0.071188
[5 empty] | 1499.9946 | 0.072999
[5 empty] | 1499.9949 | 0.074851
[5 empty] | 1499.9952 | 0.076745
[5 empty] | 1499.9955 | 0.078682
[5 empty] | 1499.9958 | 0.080663
[5 empty] | 1499.9962 | 0.082689
[5 empty] | 1499.9965 | 0.08476
[5 empty] | 1499.9968 | 0.086877
[5 empty] | 1499.9971 | 0.089041
[5 empty] | 1499.9974 | 0.091254
[5 empty] | 1499.9978 | 0.093515
[5 empty] | 1499.9981 | 0.095827
[5 empty] | 1499.9984 | 0.098189
[5 empty] | 1499.9987 | 0.1006
[5 empty] | 1499.999 | 0.10307
[5 empty] | 1499.9994 | 0.10559
[5 empty] | 1499.9997 | 0.10817
[5 empty] | 1500 | 0.1108
[5 empty] | 1500.0003 | 0.11348
[5 empty] | 1500.0006 | 0.11623
[5 empty] | 1500.001 | 0.11903
[5 empty] | 1500.0013 | 0.1219
[5 empty] | 1500.0016 | 0.12482
[5 empty] | 1500.0019 | 0.12781
[5 empty] | 1500.0022 | 0.13086
[5 empty] | 1500.0026 | 0.13397
[5 empty] | 1500.0029 | 0.13715
[5 empty] | 1500.0032 | 0.1404
[5 empty] | 1500.0035 | 0.14371
[5 empty] | 1500.0038 | 0.1471
[5 empty] | 1500.0042 | 0.15055
[5 empty] | 1500.0045 | 0.15407
[5 empty] | 1500.0048 | 0.15767
[5 empty] | 1500.0051 | 0.16134
[5 empty] | 1500.0054 | 0.16508
[5 empty] | 1500.0058 | 0.1689
[5 empty] | 1500.0061 | 0.1728
[5 empty] | 1500.0064 | 0.17678
[5 empty] | 1500.0067 | 0.18083
[5 empty] | 1500.007 | 0.18497
[5 empty] | 1500.0074 | 0.18919
[5 empty] | 1500.0077 | 0.1935
[5 empty] | 1500.008 | 0.19789
[5 empty] | 1500.0083 | 0.20236
[5 empty] | 1500.0086 | 0.20693
[5 empty] | 1500.009 | 0.21158
[5 empty] | 1500.0093 | 0.21632
[5 empty] | 1500.0096 | 0.22116
[5 empty] | 1500.0099 | 0.22609
[5 empty] | 1500.0102 | 0.23112
[5 empty] | 1500.0106 | 0.23624
[5 empty] | 1500.0109 | 0.24146
[5 empty] | 1500.0112 | 0.24678
[5 empty] | 1500.0115 | 0.2522
[5 empty] | 1500.0118 | 0.25772
[5 empty] | 1500.0122 | 0.26335
[5 empty] | 1500.0125 | 0.26908
[5 empty] | 1500.0128 | 0.27492
[5 empty] | 1500.0131 | 0.28087
[5 empty] | 1500.0134 | 0.28693
[5 empty] | 1500.0138 | 0.2931
[5 empty] | 1500.0141 | 0.29939
[5 empty] | 1500.0144 | 0.30579
[5 empty] | 1500.0147 | 0.31231
[5 empty] | 1500.015 | 0.31894
[5 empty] | 1500.0154 | 0.3257
[5 empty] | 1500.0157 | 0.33257
[5 empty] | 1500.016 | 0.33957
[5 empty] | 1500.0163 | 0.3467
[5 empty] | 1500.0166 | 0.35395
[5 empty] | 1500.017 | 0.36133
[5 empty] | 1500.0173 | 0.36885
[5 empty] | 1500.0176 | 0.37649
[5 empty] | 1500.0179 | 0.38427
[5 empty] | 1500.0182 | 0.39218
[5 empty] | 1500.0186 | 0.40023
[5 empty] | 1500.0189 | 0.40842
[5 empty] | 1500.0192 | 0.41675
[5 empty] | 1500.0195 | 0.42523
[5 empty] | 1500.0198 | 0.43384
[5 empty] | 1500.0202 | 0.44261
[5 empty] | 1500.0205 | 0.45152
[5 empty] | 1500.0208 | 0.46058
[5 empty] | 1500.0211 | 0.4698
[5 empty] | 1500.0214 | 0.47916
[5 empty] | 1500.0218 | 0.48869
[5 empty] | 1500.0221 | 0.49837
[5 empty] | 1500.0224 | 0.50821
[5 empty] | 1500.0227 | 0.51821
[5 empty] | 1500.023 | 0.52837
[5 empty] | 1500.0234 | 0.5387
[5 empty] | 1500.0237 | 0.5492
[5 empty] | 1500.024 | 0.55986
[5 empty] | 1500.0243 | 0.5707
[5 empty] | 1500.0246 | 0.58171
[5 empty] | 1500.025 | 0.59289
[5 empty] | 1500.0253 | 0.60425
[5 empty] | 1500.0256 | 0.61578
[5 empty] | 1500.0259 | 0.6275
[5 empty] | 1500.0262 | 0.6394
[5 empty] | 1500.0266 | 0.65148
[5 empty] | 1500.0269 | 0.66375
[5 empty] | 1500.0272 | 0.6762
[5 empty] | 1500.0275 | 0.68885
[5 empty] | 1500.0278 | 0.70168
[5 empty] | 1500.0282 | 0.71472
[5 empty] | 1500.0285 | 0.72794
[5 empty] | 1500.0288 | 0.74136
[5 empty] | 1500.0291 | 0.75499
[5 empty] | 1500.0294 | 0.76881
[5 empty] | 1500.0298 | 0.78284
[5 empty] | 1500.0301 | 0.79707
[5 empty] | 1500.0304 | 0.81151
[5 empty] | 1500.0307 | 0.82615
[5 empty] | 1500.031 | 0.84101
[5 empty] | 1500.0314 | 0.85608
[5 empty] | 1500.0317 | 0.87136
[5 empty] | 1500.032 | 0.88686
[5 empty] | 1500.0323 | 0.90258
[5 empty] | 1500.0326 | 0.91852
[5 empty] | 1500.033 | 0.93468
[5 empty] | 1500.0333 | 0.95106
[5 empty] | 1500.0336 | 0.96767
[5 empty] | 1500.0339 | 0.98451
[5 empty] | 1500.0342 | 1.0016
[5 empty] | 1500.0346 | 1.0189
[5 empty] | 1500.0349 | 1.0364
[5 empty] | 1500.0352 | 1.0542
[5 empty] | 1500.0355 | 1.0721
[5 empty] | 1500.0358 | 1.0904
[5 empty] | 1500.0362 | 1.1089
[5 empty] | 1500.0365 | 1.1276
[5 empty] | 1500.0368 | 1.1465
[5 empty] | 1500.0371 | 1.1657
[5 empty] | 1500.0374 | 1.1852
[5 empty] | 1500.0378 | 1.2049
[5 empty] | 1500.0381 | 1.2248
[5 empty] | 1500.0384 | 1.245
[5 empty] | 1500.0387 | 1.2654
[5 empty] | 1500.039 | 1.2861
[5 empty] | 1500.0394 | 1.3071
[5 empty] | 1500.0397 | 1.3283
[5 empty] | 1500.04 | 1.3498
[5 empty] | 1500.0403 | 1.3715
[5 empty] | 1500.0406 | 1.3935
[5 empty] | 1500.041 | 1.4157
[5 empty] | 1500.0413 | 1.4382
[5 empty] | 1500.0416 | 1.461
[5 empty] | 1500.0419 | 1.4841
[5 empty] | 1500.0422 | 1.5074
[5 empty] | 1500.0426 | 1.531
[5 empty] | 1500.0429 | 1.5548
[5 empty] | 1500.0432 | 1.5789
[5 empty] | 1500.0435 | 1.6033
[5 empty] | 1500.0438 | 1.628
[5 empty] | 1500.0442 | 1.6529
[5 empty] | 1500.0445 | 1.6781
[5 empty] | 1500.0448 | 1.7036
[5 empty] | 1500.0451 | 1.7293
[5 empty] | 1500.0454 | 1.7554
[5 empty] | 1500.0458 | 1.7817
[5 empty] | 1500.0461 | 1.8083
[5 empty] | 1500.0464 | 1.8352
[5 empty] | 1500.0467 | 1.8623
[5 empty] | 1500.047 | 1.8898
[5 empty] | 1500.0474 | 1.9175
[5 empty] | 1500.0477 | 1.9455
[5 empty] | 1500.048 | 1.9738
[5 empty] | 1500.0483 | 2.0023
[5 empty] | 1500.0486 | 2.0312
[5 empty] | 1500.049 | 2.0603
[5 empty] | 1500.0493 | 2.0897
[5 empty] | 1500.0496 | 2.1194
[5 empty] | 1500.0499 | 2.1494
[5 empty] | 1500.0502 | 2.1797
[5 empty] | 1500.0506 | 2.2102
[5 empty] | 1500.0509 | 2.2411
[5 empty] | 1500.0512 | 2.2722
[5 empty] | 1500.0515 | 2.3036
[5 empty] | 1500.0518 | 2.3353
[5 empty] | 1500.0522 | 2.3673
[5 empty] | 1500.0525 | 2.3995
[5 empty] | 1500.0528 | 2.4321
[5 empty] | 1500.0531 | 2.4649
[5 empty] | 1500.0534 | 2.498
[5 empty] | 1500.0538 | 2.5314
[5 empty] | 1500.0541 | 2.5651
[5 empty] | 1500.0544 | 2.599
[5 empty] | 1500.0547 | 2.6333
[5 empty] | 1500.055 | 2.6678
[5 empty] | 1500.0554 | 2.7026
[5 empty] | 1500.0557 | 2.7377
[5 empty] | 1500.056 | 2.773
[5 empty] | 1500.0563 | 2.8087
[5 empty] | 1500.0566 | 2.8446
[5 empty] | 1500.057 | 2.8807
[5 empty] | 1500.0573 | 2.9172
[5 empty] | 1500.0576 | 2.9539
[5 empty] | 1500.0579 | 2.9909
[5 empty] | 1500.0582 | 3.0282
[5 empty] | 1500.0586 | 3.0657
[5 empty] | 1500.0589 | 3.1035
[5 empty] | 1500.0592 | 3.1416
[5 empty] | 1500.0595 | 3.1799
[5 empty] | 1500.0598 | 3.2185
[5 empty] | 1500.0602 | 3.2574
[5 empty] | 1500.0605 | 3.2965
[5 empty] | 1500.0608 | 3.3359
[5 empty] | 1500.0611 | 3.3755
[5 empty] | 1500.0614 | 3.4154
[5 empty] | 1500.0618 | 3.4555
[5 empty] | 1500.0621 | 3.4959
[5 empty] | 1500.0624 | 3.5365
[5 empty] | 1500.0627 | 3.5774
[5 empty] | 1500.063 | 3.6185
[5 empty] | 1500.0634 | 3.6598
[5 empty] | 1500.0637 | 3.7014
[5 empty] | 1500.064 | 3.7432
[5 empty] | 1500.0643 | 3.7852
[5 empty] | 1500.0646 | 3.8275
[5 empty] | 1500.065 | 3.87
[5 empty] | 1500.0653 | 3.9127
[5 empty] | 1500.0656 | 3.9556
[5 empty] | 1500.0659 | 3.9988
[5 empty] | 1500.0662 | 4.0421
[5 empty] | 1500.0666 | 4.0857
[5 empty] | 1500.0669 | 4.1295
[5 empty] | 1500.0672 | 4.1734
[5 empty] | 1500.0675 | 4.2176
[5 empty] | 1500.0678 | 4.262
[5 empty] | 1500.0682 | 4.3065
[5 empty] | 1500.0685 | 4.3513
[5 empty] | 1500.0688 | 4.3962
[5 empty] | 1500.0691 | 4.4413
[5 empty] | 1500.0694 | 4.4866
[5 empty] | 1500.0698 | 4.532
[5 empty] | 1500.0701 | 4.5776
[5 empty] | 1500.0704 | 4.6234
[5 empty] | 1500.0707 | 4.6694
[5 empty] | 1500.071 | 4.7155
[5 empty] | 1500.0714 | 4.7617
[5 empty] | 1500.0717 | 4.8081
[5 empty] | 1500.072 | 4.8547
[5 empty] | 1500.0723 | 4.9013
[5 empty] | 1500.0726 | 4.9481
[5 empty] | 1500.073 | 4.9951
[5 empty] | 1500.0733 | 5.0421
[5 empty] | 1500.0736 | 5.0893
[5 empty] | 1500.0739 | 5.1366
[5 empty] | 1500.0742 | 5.184
[5 empty] | 1500.0746 | 5.2315
[5 empty] | 1500.0749 | 5.279
[5 empty] | 1500.0752 | 5.3267
[5 empty] | 1500.0755 | 5.3745
[5 empty] | 1500.0758 | 5.4223
[5 empty] | 1500.0762 | 5.4703
[5 empty] | 1500.0765 | 5.5183
[5 empty] | 1500.0768 | 5.5663
[5 empty] | 1500.0771 | 5.6145
[5 empty] | 1500.0774 | 5.6626
[5 empty] | 1500.0778 | 5.7109
[5 empty] | 1500.0781 | 5.7591
[5 empty] | 1500.0784 | 5.8074
[5 empty] | 1500.0787 | 5.8558
[5 empty] | 1500.079 | 5.9041
[5 empty] | 1500.0794 | 5.9525
[5 empty] | 1500.0797 | 6.0009
[5 empty] | 1500.08 | 6.0493
[5 empty] | 1500.0803 | 6.0977
[5 empty] | 1500.0806 | 6.146
[5 empty] | 1500.081 | 6.1944
[5 empty] | 1500.0813 | 6.2428
[5 empty] | 1500.0816 | 6.2911
[5 empty] | 1500.0819 | 6.3394
[5 empty] | 1500.0822 | 6.3877
[5 empty] | 1500.0826 | 6.4359
[5 empty] | 1500.0829 | 6.484
[5 empty] | 1500.0832 | 6.5322
[5 empty] | 1500.0835 | 6.5802
[5 empty] | 1500.0838 | 6.6282
[5 empty] | 1500.0842 | 6.6761
[5 empty] | 1500.0845 | 6.7239
[5 empty] | 1500.0848 | 6.7716
[5 empty] | 1500.0851 | 6.8192
[5 empty] | 1500.0854 | 6.8667
[5 empty] | 1500.0858 | 6.9141
[5 empty] | 1500.0861 | 6.9614
[5 empty] | 1500.0864 | 7.0086
[5 empty] | 1500.0867 | 7.0556
[5 empty] | 1500.087 | 7.1025
[5 empty] | 1500.0874 | 7.1493
[5 empty] | 1500.0877 | 7.1959
[5 empty] | 1500.088 | 7.2423
[5 empty] | 1500.0883 | 7.2886
[5 empty] | 1500.0886 | 7.3346
[5 empty] | 1500.089 | 7.3806
[5 empty] | 1500.0893 | 7.4263
[5 empty] | 1500.0896 | 7.4718
[5 empty] | 1500.0899 | 7.5171
[5 empty] | 1500.0902 | 7.5623
[5 empty] | 1500.0906 | 7.6072
[5 empty] | 1500.0909 | 7.6518
[5 empty] | 1500.0912 | 7.6963
[5 empty] | 1500.0915 | 7.7405
[5 empty] | 1500.0918 | 7.7845
[5 empty] | 1500.0922 | 7.8282
[5 empty] | 1500.0925 | 7.8716
[5 empty] | 1500.0928 | 7.9148
[5 empty] | 1500.0931 | 7.9577
[5 empty] | 1500.0934 | 8.0004
[5 empty] | 1500.0938 | 8.0427
[5 empty] | 1500.0941 | 8.0848
[5 empty] | 1500.0944 | 8.1266
[5 empty] | 1500.0947 | 8.168
[5 empty] | 1500.095 | 8.2091
[5 empty] | 1500.0954 | 8.25
[5 empty] | 1500.0957 | 8.2905
[5 empty] | 1500.096 | 8.3306
[5 empty] | 1500.0963 | 8.3704
[5 empty] | 1500.0966 | 8.4099
[5 empty] | 1500.097 | 8.449
[5 empty] | 1500.0973 | 8.4878
[5 empty] | 1500.0976 | 8.5261
[5 empty] | 1500.0979 | 8.5642
[5 empty] | 1500.0982 | 8.6018
[5 empty] | 1500.0986 | 8.639
[5 empty] | 1500.0989 | 8.6759
[5 empty] | 1500.0992 | 8.7123
[5 empty] | 1500.0995 | 8.7484
[5 empty] | 1500.0998 | 8.784
[5 empty] | 1500.1002 | 8.8192
[5 empty] | 1500.1005 | 8.854
[5 empty] | 1500.1008 | 8.8883
[5 empty] | 1500.1011 | 8.9222
[5 empty] | 1500.1014 | 8.9557
[5 empty] | 1500.1018 | 8.9887
[5 empty] | 1500.1021 | 9.0213
[5 empty] | 1500.1024 | 9.0534
[5 empty] | 1500.1027 | 9.085
[5 empty] | 1500.103 | 9.1162
[5 empty] | 1500.1034 | 9.1468
[5 empty] | 1500.1037 | 9.177
[5 empty] | 1500.104 | 9.2068
[5 empty] | 1500.1043 | 9.236
[5 empty] | 1500.1046 | 9.2647
[5 empty] | 1500.105 | 9.2929
[5 empty] | 1500.1053 | 9.3206
[5 empty] | 1500.1056 | 9.3478
[5 empty] | 1500.1059 | 9.3744
[5 empty] | 1500.1062 | 9.4006
[5 empty] | 1500.1066 | 9.4262
[5 empty] | 1500.1069 | 9.4512
[5 empty] | 1500.1072 | 9.4758
[5 empty] | 1500.1075 | 9.4997
[5 empty] | 1500.1078 | 9.5232
[5 empty] | 1500.1082 | 9.5461
[5 empty] | 1500.1085 | 9.5684
[5 empty] | 1500.1088 | 9.5902
[5 empty] | 1500.1091 | 9.6114
[5 empty] | 1500.1094 | 9.632
[5 empty] | 1500.1098 | 9.652
[5 empty] | 1500.1101 | 9.6715
[5 empty] | 1500.1104 | 9.6904
[5 empty] | 1500.1107 | 9.7087
[5 empty] | 1500.111 | 9.7265
[5 empty] | 1500.1114 | 9.7436
[5 empty] | 1500.1117 | 9.7601
[5 empty] | 1500.112 | 9.7761
[5 empty] | 1500.1123 | 9.7914
[5 empty] | 1500.1126 | 9.8061
[5 empty] | 1500.113 | 9.8203
[5 empty] | 1500.1133 | 9.8338
[5 empty] | 1500.1136 | 9.8467
[5 empty] | 1500.1139 | 9.859
[5 empty] | 1500.1142 | 9.8707
[5 empty] | 1500.1146 | 9.8817
[5 empty] | 1500.1149 | 9.8922
[5 empty] | 1500.1152 | 9.902
[5 empty] | 1500.1155 | 9.9112
[5 empty] | 1500.1158 | 9.9198
[5 empty] | 1500.1162 | 9.9277
[5 empty] | 1500.1165 | 9.935
[5 empty] | 1500.1168 | 9.9417
[5 empty] | 1500.1171 | 9.9477
[5 empty] | 1500.1174 | 9.9532
[5 empty] | 1500.1178 | 9.9579
[5 empty] | 1500.1181 | 9.9621
[5 empty] | 1500.1184 | 9.9656
[5 empty] | 1500.1187 | 9.9685
[5 empty] | 1500.119 | 9.9707
[5 empty] | 1500.1194 | 9.9723
[5 empty] | 1500.1197 | 9.9732
[5 empty] | 1500.12 | 9.9736
[5 empty] | 1500.1203 | 9.9732
[5 empty] | 1500.1206 | 9.9723
[5 empty] | 1500.121 | 9.9707
[5 empty] | 1500.1213 | 9.9685
[5 empty] | 1500.1216 | 9.9656
[5 empty] | 1500.1219 | 9.9621
[5 empty] | 1500.1222 | 9.9579
[5 empty] | 1500.1226 | 9.9532
[5 empty] | 1500.1229 | 9.9477
[5 empty] | 1500.1232 | 9.9417
[5 empty] | 1500.1235 | 9.935
[5 empty] | 1500.1238 | 9.9277
[5 empty] | 1500.1242 | 9.9198
[5 empty] | 1500.1245 | 9.9112
[5 empty] | 1500.1248 | 9.902
[5 empty] | 1500.1251 | 9.8922
[5 empty] | 1500.1254 | 9.8817
[5 empty] | 1500.1258 | 9.8707
[5 empty] | 1500.1261 | 9.859
[5 empty] | 1500.1264 | 9.8467
[5 empty] | 1500.1267 | 9.8338
[5 empty] | 1500.127 | 9.8203
[5 empty] | 1500.1274 | 9.8061
[5 empty] | 1500.1277 | 9.7914
[5 empty] | 1500.128 | 9.7761
[5 empty] | 1500.1283 | 9.7601
[5 empty] | 1500.1286 | 9.7436
[5 empty] | 1500.129 | 9.7265
[5 empty] | 1500.1293 | 9.7087
[5 empty] | 1500.1296 | 9.6904
[5 empty] | 1500.1299 | 9.6715
[5 empty] | 1500.1302 | 9.652
[5 empty] | 1500.1306 | 9.632
[5 empty] | 1500.1309 | 9.6114
[5 empty] | 1500.1312 | 9.5902
[5 empty] | 1500.1315 | 9.5684
[5 empty] | 1500.1318 | 9.5461
[5 empty] | 1500.1322 | 9.5232
[5 empty] | 1500.1325 | 9.4997
[5 empty] | 1500.1328 | 9.4758
[5 empty] | 1500.1331 | 9.4512
[5 empty] | 1500.1334 | 9.4262
[5 empty] | 1500.1338 | 9.4006
[5 empty] | 1500.1341 | 9.3744
[5 empty] | 1500.1344 | 9.3478
[5 empty] | 1500.1347 | 9.3206
[5 empty] | 1500.135 | 9.2929
[5 empty] | 1500.1354 | 9.2647
[5 empty] | 1500.1357 | 9.236
[5 empty] | 1500.136 | 9.2068
[5 empty] | 1500.1363 | 9.177
[5 empty] | 1500.1366 | 9.1468
[5 empty] | 1500.137 | 9.1162
[5 empty] | 1500.1373 | 9.085
[5 empty] | 1500.1376 | 9.0534
[5 empty] | 1500.1379 | 9.0213
[5 empty] | 1500.1382 | 8.9887
[5 empty] | 1500.1386 | 8.9557
[5 empty] | 1500.1389 | 8.9222
[5 empty] | 1500.1392 | 8.8883
[5 empty] | 1500.1395 | 8.854
[5 empty] | 1500.1398 | 8.8192
[5 empty] | 1500.1402 | 8.784
[5 empty] | 1500.1405 | 8.7484
[5 empty] | 1500.1408 | 8.7123
[5 empty] | 1500.1411 | 8.6759
[5 empty] | 1500.1414 | 8.639
[5 empty] | 1500.1418 | 8.6018
[5 empty] | 1500.1421 | 8.5642
[5 empty] | 1500.1424 | 8.5261
[5 empty] | 1500.1427 | 8.4878
[5 empty] | 1500.143 | 8.449
[5 empty] | 1500.1434 | 8.4099
[5 empty] | 1500.1437 | 8.3704
[5 empty] | 1500.144 | 8.3306
[5 empty] | 1500.1443 | 8.2905
[5 empty] | 1500.1446 | 8.25
[5 empty] | 1500.145 | 8.2091
[5 empty] | 1500.1453 | 8.168
[5 empty] | 1500.1456 | 8.1266
[5 empty] | 1500.1459 | 8.0848
[5 empty] | 1500.1462 | 8.0427
[5 empty] | 1500.1466 | 8.0004
[5 empty] | 1500.1469 | 7.9577
[5 empty] | 1500.1472 | 7.9148
[5 empty] | 1500.1475 | 7.8716
[5 empty] | 1500.1478 | 7.8282
[5 empty] | 1500.1482 | 7.7845
[5 empty] | 1500.1485 | 7.7405
[5 empty] | 1500.1488 | 7.6963
[5 empty] | 1500.1491 | 7.6518
[5 empty] | 1500.1494 | 7.6072
[5 empty] | 1500.1498 | 7.5623
[5 empty] | 1500.1501 | 7.5171
[5 empty] | 1500.1504 | 7.4718
[5 empty] | 1500.1507 | 7.4263
[5 empty] | 1500.151 | 7.3806
[5 empty] | 1500.1514 | 7.3346
[5 empty] | 1500.1517 | 7.2886
[5 empty] | 1500.152 | 7.2423
[5 empty] | 1500.1523 | 7.1959
[5 empty] | 1500.1526 | 7.1493
[5 empty] | 1500.153 | 7.1025
[5 empty] | 1500.1533 | 7.0556
[5 empty] | 1500.1536 | 7.0086
[5 empty] | 1500.1539 | 6.9614
[5 empty] | 1500.1542 | 6.9141
[5 empty] | 1500.1546 | 6.8667
[5 empty] | 1500.1549 | 6.8192
[5 empty] | 1500.1552 | 6.7716
[5 empty] | 1500.1555 | 6.7239
[5 empty] | 1500.1558 | 6.6761
[5 empty] | 1500.1562 | 6.6282
[5 empty] | 1500.1565 | 6.5802
[5 empty] | 1500.1568 | 6.5322
[5 empty] | 1500.1571 | 6.484
[5 empty] | 1500.1574 | 6.4359
[5 empty] | 1500.1578 | 6.3877
[5 empty] | 1500.1581 | 6.3394
[5 empty] | 1500.1584 | 6.2911
[5 empty] | 1500.1587 | 6.2428
[5 empty] | 1500.159 | 6.1944
[5 empty] | 1500.1594 | 6.146
[5 empty] | 1500.1597 | 6.0977
[5 empty] | 1500.16 | 6.0493
[5 empty] | 1500.1603 | 6.0009
[5 empty] | 1500.1606 | 5.9525
[5 empty] | 1500.161 | 5.9041
[5 empty] | 1500.1613 | 5.8558
[5 empty] | 1500.1616 | 5.8074
[5 empty] | 1500.1619 | 5.7591
[5 empty] | 1500.1622 | 5.7109
[5 empty] | 1500.1626 | 5.6626
[5 empty] | 1500.1629 | 5.6145
[5 empty] | 1500.1632 | 5.5663
[5 empty] | 1500.1635 | 5.5183
[5 empty] | 1500.1638 | 5.4703
[5 empty] | 1500.1642 | 5.4223
[5 empty] | 1500.1645 | 5.3745
[5 empty] | 1500.1648 | 5.3267
[5 empty] | 1500.1651 | 5.279
[5 empty] | 1500.1654 | 5.2315
[5 empty] | 1500.1658 | 5.184
[5 empty] | 1500.1661 | 5.1366
[5 empty] | 1500.1664 | 5.0893
[5 empty] | 1500.1667 | 5.0421
[5 empty] | 1500.167 | 4.9951
[5 empty] | 1500.1674 | 4.9481
[5 empty] | 1500.1677 | 4.9013
[5 empty] | 1500.168 | 4.8547
[5 empty] | 1500.1683 | 4.8081
[5 empty] | 1500.1686 | 4.7617
[5 empty] | 1500.169 | 4.7155
[5 empty] | 1500.1693 | 4.6694
[5 empty] | 1500.1696 | 4.6234
[5 empty] | 1500.1699 | 4.5776
[5 empty] | 1500.1702 | 4.532
[5 empty] | 1500.1706 | 4.4866
[5 empty] | 1500.1709 | 4.4413
[5 empty] | 1500.1712 | 4.3962
[5 empty] | 1500.1715 | 4.3513
[5 empty] | 1500.1718 | 4.3065
[5 empty] | 1500.1722 | 4.262
[5 empty] | 1500.1725 | 4.2176
[5 empty] | 1500.1728 | 4.1734
[5 empty] | 1500.1731 | 4.1295
[5 empty] | 1500.1734 | 4.0857
[5 empty] | 1500.1738 | 4.0421
[5 empty] | 1500.1741 | 3.9988
[5 empty] | 1500.1744 | 3.9556
[5 empty] | 1500.1747 | 3.9127
[5 empty] | 1500.175 | 3.87
[5 empty] | 1500.1754 | 3.8275
[5 empty] | 1500.1757 | 3.7852
[5 empty] | 1500.176 | 3.7432
[5 empty] | 1500.1763 | 3.7014
[5 empty] | 1500.1766 | 3.6598
[5 empty] | 1500.177 | 3.6185
[5 empty] | 1500.1773 | 3.5774
[5 empty] | 1500.1776 | 3.5365
[5 empty] | 1500.1779 | 3.4959
[5 empty] | 1500.1782 | 3.4555
[5 empty] | 1500.1786 | 3.4154
[5 empty] | 1500.1789 | 3.3755
[5 empty] | 1500.1792 | 3.3359
[5 empty] | 1500.1795 | 3.2965
[5 empty] | 1500.1798 | 3.2574
[5 empty] | 1500.1802 | 3.2185
[5 empty] | 1500.1805 | 3.1799
[5 empty] | 1500.1808 | 3.1416
[5 empty] | 1500.1811 | 3.1035
[5 empty] | 1500.1814 | 3.0657
[5 empty] | 1500.1818 | 3.0282
[5 empty] | 1500.1821 | 2.9909
[5 empty] | 1500.1824 | 2.9539
[5 empty] | 1500.1827 | 2.9172
[5 empty] | 1500.183 | 2.8807
[5 empty] | 1500.1834 | 2.8446
[5 empty] | 1500.1837 | 2.8087
[5 empty] | 1500.184 | 2.773
[5 empty] | 1500.1843 | 2.7377
[5 empty] | 1500.1846 | 2.7026
[5 empty] | 1500.185 | 2.6678
[5 empty] | 1500.1853 | 2.6333
[5 empty] | 1500.1856 | 2.599
[5 empty] | 1500.1859 | 2.5651
[5 empty] | 1500.1862 | 2.5314
[5 empty] | 1500.1866 | 2.498
[5 empty] | 1500.1869 | 2.4649
[5 empty] | 1500.1872 | 2.4321
[5 empty] | 1500.1875 | 2.3995
[5 empty] | 1500.1878 | 2.3673
[5 empty] | 1500.1882 | 2.3353
[5 empty] | 1500.1885 | 2.3036
[5 empty] | 1500.1888 | 2.2722
[5 empty] | 1500.1891 | 2.2411
[5 empty] | 1500.1894 | 2.2102
[5 empty] | 1500.1898 | 2.1797
[5 empty] | 1500.1901 | 2.1494
[5 empty] | 1500.1904 | 2.1194
[5 empty] | 1500.1907 | 2.0897
[5 empty] | 1500.191 | 2.0603
[5 empty] | 1500.1914 | 2.0312
[5 empty] | 1500.1917 | 2.0023
[5 empty] | 1500.192 | 1.9738
[5 empty] | 1500.1923 | 1.9455
[5 empty] | 1500.1926 | 1.9175
[5 empty] | 1500.193 | 1.8898
[5 empty] | 1500.1933 | 1.8623
[5 empty] | 1500.1936 | 1.8352
[5 empty] | 1500.1939 | 1.8083
[5 empty] | 1500.1942 | 1.7817
[5 empty] | 1500.1946 | 1.7554
[5 empty] | 1500.1949 | 1.7293
[5 empty] | 1500.1952 | 1.7036
[5 empty] | 1500.1955 | 1.6781
[5 empty] | 1500.1958 | 1.6529
[5 empty] | 1500.1962 | 1.628
[5 empty] | 1500.1965 | 1.6033
[5 empty] | 1500.1968 | 1.5789
[5 empty] | 1500.1971 | 1.5548
[5 empty] | 1500.1974 | 1.531
[5 empty] | 1500.1978 | 1.5074
[5 empty] | 1500.1981 | 1.4841
[5 empty] | 1500.1984 | 1.461
[5 empty] | 1500.1987 | 1.4382
[5 empty] | 1500.199 | 1.4157
[5 empty] | 1500.1994 | 1.3935
[5 empty] | 1500.1997 | 1.3715
[5 empty] | 1500.2 | 1.3498
[5 empty] | 1500.2003 | 1.3283
[5 empty] | 1500.2006 | 1.3071
[5 empty] | 1500.201 | 1.2861
[5 empty] | 1500.2013 | 1.2654
[5 empty] | 1500.2016 | 1.245
[5 empty] | 1500.2019 | 1.2248
[5 empty] | 1500.2022 | 1.2049
[5 empty] | 1500.2026 | 1.1852
[5 empty] | 1500.2029 | 1.1657
[5 empty] | 1500.2032 | 1.1465
[5 empty] | 1500.2035 | 1.1276
[5 empty] | 1500.2038 | 1.1089
[5 empty] | 1500.2042 | 1.0904
[5 empty] | 1500.2045 | 1.0721
[5 empty] | 1500.2048 | 1.0542
[5 empty] | 1500.2051 | 1.0364
[5 empty] | 1500.2054 | 1.0189
[5 empty] | 1500.2058 | 1.0016
[5 empty] | 1500.2061 | 0.98451
[5 empty] | 1500.2064 | 0.96767
[5 empty] | 1500.2067 | 0.95106
[5 empty] | 1500.207 | 0.93468
[5 empty] | 1500.2074 | 0.91852
[5 empty] | 1500.2077 | 0.90258
[5 empty] | 1500.208 | 0.88686
[5 empty] | 1500.2083 | 0.87136
[5 empty] | 1500.2086 | 0.85608
[5 empty] | 1500.209 | 0.84101
[5 empty] | 1500.2093 | 0.82615
[5 empty] | 1500.2096 | 0.81151
[5 empty] | 1500.2099 | 0.79707
[5 empty] | 1500.2102 | 0.78284
[5 empty] | 1500.2106 | 0.76881
[5 empty] | 1500.2109 | 0.75499
[5 empty] | 1500.2112 | 0.74136
[5 empty] | 1500.2115 | 0.72794
[5 empty] | 1500.2118 | 0.71472
[5 empty] | 1500.2122 | 0.70168
[5 empty] | 1500.2125 | 0.68885
[5 empty] | 1500.2128 | 0.6762
[5 empty] | 1500.2131 | 0.66375
[5 empty] | 1500.2134 | 0.65148
[5 empty] | 1500.2138 | 0.6394
[5 empty] | 1500.2141 | 0.6275
[5 empty] | 1500.2144 | 0.61578
[5 empty] | 1500.2147 | 0.60425
[5 empty] | 1500.215 | 0.59289
[5 empty] | 1500.2154 | 0.58171
[5 empty] | 1500.2157 | 0.5707
[5 empty] | 1500.216 | 0.55986
[5 empty] | 1500.2163 | 0.5492
[5 empty] | 1500.2166 | 0.5387
[5 empty] | 1500.217 | 0.52837
[5 empty] | 1500.2173 | 0.51821
[5 empty] | 1500.2176 | 0.50821
[5 empty] | 1500.2179 | 0.49837
[5 empty] | 1500.2182 | 0.48869
[5 empty] | 1500.2186 | 0.47916
[5 empty] | 1500.2189 | 0.4698
[5 empty] | 1500.2192 | 0.46058
[5 empty] | 1500.2195 | 0.45152
[5 empty] | 1500.2198 | 0.44261
[5 empty] | 1500.2202 | 0.43384
[5 empty] | 1500.2205 | 0.42523
[5 empty] | 1500.2208 | 0.41675
[5 empty] | 1500.2211 | 0.40842
[5 empty] | 1500.2214 | 0.40023
[5 empty] | 1500.2218 | 0.39218
[5 empty] | 1500.2221 | 0.38427
[5 empty] | 1500.2224 | 0.37649
[5 empty] | 1500.2227 | 0.36885
[5 empty] | 1500.223 | 0.36133
[5 empty] | 1500.2234 | 0.35395
[5 empty] | 1500.2237 | 0.3467
[5 empty] | 1500.224 | 0.33957
[5 empty] | 1500.2243 | 0.33257
[5 empty] | 1500.2246 | 0.3257
[5 empty] | 1500.225 | 0.31894
[5 empty] | 1500.2253 | 0.31231
[5 empty] | 1500.2256 | 0.30579
[5 empty] | 1500.2259 | 0.29939
[5 empty] | 1500.2262 | 0.2931
[5 empty] | 1500.2266 | 0.28693
[5 empty] | 1500.2269 | 0.28087
[5 empty] | 1500.2272 | 0.27492
[5 empty] | 1500.2275 | 0.26908
[5 empty] | 1500.2278 | 0.26335
[5 empty] | 1500.2282 | 0.25772
[5 empty] | 1500.2285 | 0.2522
[5 empty] | 1500.2288 | 0.24678
[5 empty] | 1500.2291 | 0.24146
[5 empty] | 1500.2294 | 0.23624
[5 empty] | 1500.2298 | 0.23112
[5 empty] | 1500.2301 | 0.22609
[5 empty] | 1500.2304 | 0.22116
[5 empty] | 1500.2307 | 0.21632
[5 empty] | 1500.231 | 0.21158
[5 empty] | 1500.2314 | 0.20693
[5 empty] | 1500.2317 | 0.20236
[5 empty] | 1500.232 | 0.19789
[5 empty] | 1500.2323 | 0.1935
[5 empty] | 1500.2326 | 0.18919
[5 empty] | 1500.233 | 0.18497
[5 empty] | 1500.2333 | 0.18083
[5 empty] | 1500.2336 | 0.17678
[5 empty] | 1500.2339 | 0.1728
[5 empty] | 1500.2342 | 0.1689
[5 empty] | 1500.2346 | 0.16508
[5 empty] | 1500.2349 | 0.16134
[5 empty] | 1500.2352 | 0.15767
[5 empty] | 1500.2355 | 0.15407
[5 empty] | 1500.2358 | 0.15055
[5 empty] | 1500.2362 | 0.1471
[5 empty] | 1500.2365 | 0.14371
[5 empty] | 1500.2368 | 0.1404
[5 empty] | 1500.2371 | 0.13715
[5 empty] | 1500.2374 | 0.13397
[5 empty] | 1500.2378 | 0.13086
[5 empty] | 1500.2381 | 0.12781
[5 empty] | 1500.2384 | 0.12482
[5 empty] | 1500.2387 | 0.1219
[5 empty] | 1500.239 | 0.11903
[5 empty] | 1500.2394 | 0.11623
[5 empty] | 1500.2397 | 0.11348
[5 empty] | 1500.24 | 0.1108
[5 empty] | 1500.2403 | 0.10817
[5 empty] | 1500.2406 | 0.10559
[5 empty] | 1500.241 | 0.10307
[5 empty] | 1500.2413 | 0.1006
[5 empty] | 1500.2416 | 0.098189
[5 empty] | 1500.2419 | 0.095827
[5 empty] | 1500.2422 | 0.093515
[5 empty] | 1500.2426 | 0.091254
[5 empty] | 1500.2429 | 0.089041
[5 empty] | 1500.2432 | 0.086877
[5 empty] | 1500.2435 | 0.08476
[5 empty] | 1500.2438 | 0.082689
[5 empty] | 1500.2442 | 0.080663
[5 empty] | 1500.2445 | 0.078682
[5 empty] | 1500.2448 | 0.076745
[5 empty] | 1500.2451 | 0.074851
[5 empty] | 1500.2454 | 0.072999
[5 empty] | 1500.2458 | 0.071188
[5 empty] | 1500.2461 | 0.069418
[5 empty] | 1500.2464 | 0.067687
[5 empty] | 1500.2467 | 0.065995
[5 empty] | 1500.247 | 0.064341
[5 empty] | 1500.2474 | 0.062725
[5 empty] | 1500.2477 | 0.061146
[5 empty] | 1500.248 | 0.059602
[5 empty] | 1500.2483 | 0.058094
[5 empty] | 1500.2486 | 0.05662
[5 empty] | 1500.249 | 0.05518
[5 empty] | 1500.2493 | 0.053773
[5 empty] | 1500.2496 | 0.052399
[5 empty] | 1500.2499 | 0.051057
[5 empty] | 1500.2502 | 0.049746
[5 empty] | 1500.2506 | 0.048465
[5 empty] | 1500.2509 | 0.047214
[5 empty] | 1500.2512 | 0.045993
[5 empty] | 1500.2515 | 0.044801
[5 empty] | 1500.2518 | 0.043636
[5 empty] | 1500.2522 | 0.042499
[5 empty] | 1500.2525 | 0.041389
[5 empty] | 1500.2528 | 0.040306
[5 empty] | 1500.2531 | 0.039248
[5 empty] | 1500.2534 | 0.038216
[5 empty] | 1500.2538 | 0.037208
[5 empty] | 1500.2541 | 0.036225
[5 empty] | 1500.2544 | 0.035265
[5 empty] | 1500.2547 | 0.034329
[5 empty] | 1500.255 | 0.033415
[5 empty] | 1500.2554 | 0.032524
[5 empty] | 1500.2557 | 0.031654
[5 empty] | 1500.256 | 0.030805
[5 empty] | 1500.2563 | 0.029978
[5 empty] | 1500.2566 | 0.029171
[5 empty] | 1500.257 | 0.028383
[5 empty] | 1500.2573 | 0.027616
[5 empty] | 1500.2576 | 0.026867
[5 empty] | 1500.2579 | 0.026137
[5 empty] | 1500.2582 | 0.025425
[5 empty] | 1500.2586 | 0.024731
[5 empty] | 1500.2589 | 0.024054
[5 empty] | 1500.2592 | 0.023394
[5 empty] | 1500.2595 | 0.022751
[5 empty] | 1500.2598 | 0.022124
[5 empty] | 1500.2602 | 0.021514
[5 empty] | 1500.2605 | 0.020918
[5 empty] | 1500.2608 | 0.020338
[5 empty] | 1500.2611 | 0.019773
[5 empty] | 1500.2614 | 0.019222
[5 empty] | 1500.2618 | 0.018685
[5 empty] | 1500.2621 | 0.018162
[5 empty] | 1500.2624 | 0.017653
[5 empty] | 1500.2627 | 0.017157
[5 empty] | 1500.263 | 0.016673
[5 empty] | 1500.2634 | 0.016202
[5 empty] | 1500.2637 | 0.015744
[5 empty] | 1500.264 | 0.015298
[5 empty] | 1500.2643 | 0.014863
[5 empty] | 1500.2646 | 0.014439
[5 empty] | 1500.265 | 0.014027
[5 empty] | 1500.2653 | 0.013626
[5 empty] | 1500.2656 | 0.013235
[5 empty] | 1500.2659 | 0.012855
[5 empty] | 1500.2662 | 0.012485
[5 empty] | 1500.2666 | 0.012125
[5 empty] | 1500.2669 | 0.011774
[5 empty] | 1500.2672 | 0.011433
[5 empty] | 1500.2675 | 0.011101
[5 empty] | 1500.2678 | 0.010778
[5 empty] | 1500.2682 | 0.010463
[5 empty] | 1500.2685 | 0.010158
[5 empty] | 1500.2688 | 0.0098601
[5 empty] | 1500.2691 | 0.0095706
[5 empty] | 1500.2694 | 0.0092891
[5 empty] | 1500.2698 | 0.0090153
[5 empty] | 1500.2701 | 0.008749
[5 empty] | 1500.2704 | 0.00849
[5 empty] | 1500.2707 | 0.0082382
[5 empty] | 1500.271 | 0.0079933
[5 empty] | 1500.2714 | 0.0077552
[5 empty] | 1500.2717 | 0.0075237
[5 empty] | 1500.272 | 0.0072987
[5 empty] | 1500.2723 | 0.0070799
[5 empty] | 1500.2726 | 0.0068673
[5 empty] | 1500.273 | 0.0066606
[5 empty] | 1500.2733 | 0.0064597
[5 empty] | 1500.2736 | 0.0062645
[5 empty] | 1500.2739 | 0.0060747
[5 empty] | 1500.2742 | 0.0058904
[5 empty] | 1500.2746 | 0.0057113
[5 empty] | 1500.2749 | 0.0055373
[5 empty] | 1500.2752 | 0.0053682
[5 empty] | 1500.2755 | 0.005204
[5 empty] | 1500.2758 | 0.0050444
[5 empty] | 1500.2762 | 0.0048895
[5 empty] | 1500.2765 | 0.004739
[5 empty] | 1500.2768 | 0.0045928
[5 empty] | 1500.2771 | 0.0044509
[5 empty] | 1500.2774 | 0.004313
[5 empty] | 1500.2778 | 0.0041792
[5 empty] | 1500.2781 | 0.0040493
[5 empty] | 1500.2784 | 0.0039231
[5 empty] | 1500.2787 | 0.0038007
[5 empty] | 1500.279 | 0.0036818
[5 empty] | 1500.2794 | 0.0035664
[5 empty] | 1500.2797 | 0.0034544
[5 empty] | 1500.28 | 0.0033458
[5 empty] | 1500.2803 | 0.0032403
[5 empty] | 1500.2806 | 0.0031379
[5 empty] | 1500.281 | 0.0030386
[5 empty] | 1500.2813 | 0.0029423