## Executive Summary
TRUE COST OF MEASUREMENT DECISIONS
Executive Summary Report | [7 empty] | 4
Morehouse Instrument Company — ISO/IEC 17025 Accredited — NVLAP 600259 — A2LA 1398.01 |  |  |  | Measurement Risk Scenario Selected |  |  | Simple Acceptance
KEY PARAMETERS
Tolerance | EOPR | Labor Rate | Rework Cost | Cals/Year | Price/Cal | Revenue/yr | Swap Req?
0.100% | 89% | $200/hr | $116.67 | 500 | $300 | $150,000 | No
RECOMMENDATION
Best Machine: | UCM Automated |  | Annual Profit: | 57005.7157 |  | Payback: | 1.8 years
Worst Available: | Deadweight |  | Annual Profit: | 29722.4697 |  | Spread: | 27283.246 | /yr
At 0.100% tolerance, 89% EOPR, 500 cals/yr at $200/hr: UCM Automated is the best machine at $57,006/yr profit, 1.8 years payback. Worst available: Deadweight at $29,722/yr profit.
Best option: UCM Automated at $57,006/yr profit, 1.8 years payback. Saves $27,283/yr vs worst (Deadweight).
Switching from Deadweight ($29,722 profit) to UCM Automated ($57,006 profit) saves $27,283/yr, payback in 1.8 years.
MACHINE COST & PROFITABILITY COMPARISON (Simple Acceptance Decision Rule) — Applied Force: 2,000 lbf
Machine | TUR | Labor/yr | Rework/yr | Cal Cost/yr | Swap/yr | Total Cost/yr | Profit/yr | Payback
BCM | 2 | 90000 | 3044.0993 | 2000 | 0 | 95044.0993 | 54955.9007 | 0.5 yr
PCM | 3.0303 | 90000 | 1786.3811 | 2000 | 0 | 93786.3811 | 56213.6189 | 0.6 yr
UCM Manual | 3.3333 | 100000 | 1590.479 | 2000 | 0 | 103590.479 | 46409.521 | 1.1 yr
UCM Automated | 5 | 90000 | 994.2843 | 2000 | 0 | 92994.2843 | 57005.7157 | 1.8 yr
Deadweight | 50 | 120000 | 277.5303 | 0 | 0 | 120277.5303 | 29722.4697 | 11.8 yr
Competitor 1 | 2 | 100000 | 3044.0993 | 2000 | 0 | 105044.0993 | 44955.9007 | 1.2 yr
Competitor 2 | 1.6667 | 100000 | 3632.0295 | 2000 | 0 | 105632.0295 | 44367.9705 | 0.9 yr
SAVINGS ANALYSIS — CURRENT vs. RECOMMENDED
| Worst Available | Best (Recommended) | Annual Savings
Machine | Deadweight | UCM Automated
Total Cost/yr | 120277.5303 | 92994.2843 | 27283.246
Rework Cost/yr | 277.5303 | 994.2843 | -716.754
Profit/yr | 29722.4697 | 57005.7157 | 27283.246
CUMULATIVE SAVINGS PROJECTIONS
| Investment Delta | Annual Savings | Payback | Year 1 | Year 3 | Year 5
Upgrade: Deadweight → UCM Automated | -250000 | 27283.246 | 0.0 yr | 27283.246 | 81849.738 | 136416.2299
Generated by the Morehouse True Cost Calculator. Figures are estimates based on user inputs. Decision rule: Simple Acceptance.
Report generated: May 5, 2026

## Graphs Summary
THE TRUE COST OF MEASUREMENT DECISIONS — EXECUTIVE SUMMARY
Morehouse Instrument Company — ISO/IEC 17025 Accredited — NVLAP 600259 — A2LA 1398.01
Tolerance: | 0.100% | EOPR: | 89% | Labor: | $200/hr | Cals/yr: | 500 | Swap: | No | Price/Cal: | $300
RECOMMENDATION
✔ UCM AUTOMATED — $57,006/yr PROFIT
At 0.100% tolerance, 89% EOPR, 500 cals/yr at $200/hr: 7 of 7 available machines are profitable (M6). Best: UCM Automated at $57,006/yr profit, 1.8-yr payback. Worst available: Deadweight at $29,722/yr profit.
Best option: UCM Automated at $57,006/yr profit, 1.8-yr payback. Saves $792/yr vs next best (PCM).
Switching from Deadweight ($29,722 profit) to UCM Automated ($57,006 profit) saves $27,283/yr, payback in 1.8 years.
TUR range: 1.7:1 to 50.0:1 (7 machines available). WARNING: At least one available machine below TUR 2:1.
Annual Rework by Decision Rule | [5 empty] | Annual Profit by Machine | Annual Profit by Machine
Total Annual Cost (Rework + Cal + Swap) | [5 empty] | Payback Period (Years)
Standard Switching: Cost by Loading Strategy
Breakeven Cals Chart Data | Global | Method 6 | Specific Risk
BCM | 231.0448 | 253.7997 | 389.0469
PCM | 311.3125 | 315.9331 | 371.276
UCM Manual | 538.6826 | 545.1834 | 649.5068
UCM Automated | 881.5454 | 884.4255 | 988.4482
Deadweight | 5887.8014 | 5887.8014 | 5938.4601
Competitor | 729.4168 | 1667.7256 | 0
Competitor 2 | 584.7456 | 4252.1472 | 0
Labor: Hrs/Cal editable on Range Analysis (col C). Includes setup, loading, and documentation. For C+T calibrations, roughly 1.7× single-mode time.
MACHINE COMPARISON AT A GLANCE
Machine | CMC (%) | TUR | PFR (M6) | Rework/yr | Cal/yr | Swap/yr | Total Cost | Revenue | Profit (M6) | Payback | Viable? | Hrs/Cal | Labor/yr
BCM | 0.05% | 2 | 10.4% | 3044.0993 | 2000 | 0 | 5044.0993 | 150000 | 54955.9007 | 0.5 yr | ⚠ 2.0 | 0.9 | 90000
PCM | 0.033% | 3.0303 | 4.472% | 1786.3811 | 2000 | 0 | 3786.3811 | 150000 | 56213.6189 | 0.6 yr | ⚠ 3.0 | 0.9 | 90000
UCM Manual | 0.03% | 3.3333 | 3.675% | 1590.479 | 2000 | 0 | 3590.479 | 150000 | 46409.521 | 1.1 yr | ⚠ 3.3 | 1 | 100000
UCM Automated | 0.02% | 5 | 1.704% | 994.2843 | 2000 | 0 | 2994.2843 | 150000 | 57005.7157 | 1.8 yr | ✓ 5.0 | 0.9 | 90000
Deadweight | 0.002% | 50 | 0.4758% | 277.5303 | 0 | 0 | 277.5303 | 150000 | 29722.4697 | 11.8 yr | ✓ 50.0 | 1.2 | 120000
Competitor 1 | 0.05% | 2 | 10.4% | 3044.0993 | 2000 | 0 | 5044.0993 | 150000 | 44955.9007 | 1.2 yr | ⚠ 2.0 | 1 | 100000
Competitor 2 | 0.06% | 1.6667 | 14.77% | 3632.0295 | 2000 | 0 | 5632.0295 | 150000 | 44367.9705 | 0.9 yr | ✗ 1.7 | 1 | 100000
Generated by the Morehouse True Cost Calculator. Figures are estimates based on user inputs.

## Instructions
The True Cost of Measurement Decisions
Calculator & Business Case Builder
Morehouse Instrument Company — mhforce.com — (717) 843-0081
Risk formulas validated against Sandia Primary Standards Lab bivariate normal integral (scipy dblquad).
COLOR KEY
Green text | Editable input — you enter these values
Black text | Calculated — do not edit (formula will break)
Green headers | Column headers and section titles
Gray italic text | Notes, guidance, and documentation
SHEET GUIDE
Dashboard | Start here. Enter cost inputs (labor rate, rework time, cals/yr), measurement inputs (tolerance, force), switching options, and profitability inputs (cal cost, price/cal, payback target). CMC per machine is a green input — changing it updates Range Analysis, Risk Calc, and all charts. Three charts: Annual Rework, Annual Profit, and Payback Period. Standard Switching summary and machine payback data at bottom.
Summary | Executive one-page overview. Dynamic recommendation paragraph, 5 charts (Rework, Profit, Total Cost, Payback, Standard Switching), and machine comparison table. All formula-driven — updates when Dashboard inputs change.
Range Analysis | Enter your actual force range (high/low), tolerance at each end, and CMC at each end per machine. CMC@High (col E) is wired to Dashboard — edit there. CMC@Low (col H) is a separate input on this sheet. Dual-point rework uses the WORST of PFR@High and PFR@Low. Includes the CMC Optimizer (Target TUR Reverse Calculator), Cell Sizing Calculator (auto cells/swaps), What-If Tolerance comparison, and CMC Sensitivity table with Savings vs BCM analysis.
Breakeven | Full business case: machine prices, upgrade deltas, profitability per machine (Revenue − Equipment Cost − Labor), payback periods. Includes Hrs/Cal per machine for labor cost modeling. Charts: Annual Profit and Payback Period.
Guard Band Analysis | Shows tolerance consumed by Specific Risk and Method 6 guard bands at 10 TUR levels. Chart: Tolerance Consumed by Specific Risk Guard Band vs M6 Guard Band. At TUR 4.6:1, Method 6 guard band drops to zero — full tolerance usable.
Standard Switching | Compares loading strategies (5%, 10%, 20%) with swap costs. Shows cells needed, swaps, swap cost, cal cost, rework, and total per loading level. Stacked bar chart: Annual Cost by Loading Strategy.
Risk Calc | The math engine. Bivariate normal integration for PFR, bisection solver for Global guard band, Method 6 deterministic formula, Specific Risk fixed multiplier. Single-point (Dashboard) and dual-point (Range Analysis) grids. CMC Sensitivity x-grid for the cost curve. Do not edit unless you understand the math.
Validation | Scipy comparison for all formulas. Used during development to verify accuracy.
INPUT REFERENCE
COST INPUTS |  | Cell
Hourly Labor Rate | Fully loaded labor cost including overhead. Default: $150/hr. | Dashboard B7
Rework Time | Minutes per false reject event (investigation, retest, documentation). Default: 45 min. | Dashboard B8
Calibrations per Year | Total calibrations performed annually. Drives all cost calculations. Default: 1,200. | Dashboard B10
EOPR | End-of-Period Reliability: fraction in tolerance at end of cal interval. Default: 89% (ANSI/NCSL Z540.3 threshold, NASA standard). | Dashboard B11
Show Simple Acceptance? | Yes/No toggle. Shows/hides the 4th decision rule (shared risk, no guard band) on charts. Default: No. | Dashboard B12
MEASUREMENT INPUTS |  | Cell
Device Tolerance (%) | Manufacturer's specification as a % of applied force. E.g., ±0.1% = 0.001. Flows to Range Analysis Tol@High and Risk Calc. | Dashboard B14
Applied Force (lbf) | The force point for the single-point comparison on the Dashboard. | Dashboard B15
What-If Tolerance (%) | A second tolerance for side-by-side TUR comparison. E.g., 0.025% for reference-grade. | Range Analysis B21
SWITCHING OPTIONS |  | Cell
Standard Switching Required? | Yes/No dropdown. "Yes" = lab swaps reference cells mid-run (auto-calculated swap cost). "No" = matched capacity references, zero swaps. | Dashboard B18
Swap Time per Change (min) | Time to swap one reference cell. Default: 30 min. | Dashboard B19
PROFITABILITY INPUTS |  | Cell
Cal Cost per Cell ($/yr) | Annual calibration cost per reference load cell. Default: $2,000. | Dashboard B21
Price per Calibration ($) | What you charge customers per calibration. Drives revenue and profitability. | Dashboard B22
Target Payback (years) | Desired years to recoup machine investment. | Dashboard B23
MACHINE INPUTS |  | Cell
CMC (%) per machine | Calibration and Measurement Capability at capacity. Green inputs on Dashboard rows 25–31. Flows to Range Analysis CMC@High. | Dashboard B25:B31
CMC@Low (%) per machine | CMC at the low force point. Separate from Dashboard CMC. Must be set on Range Analysis for accurate dual-point results. | Range Analysis H24:H30
Max Capacity (lbf) per machine | Maximum force capacity per machine. Machines exceeding Applied Force are marked N/A and zeroed on all charts. | Dashboard T25:T31
Hrs/Cal per machine | Average hours per calibration per machine. Includes setup, loading, documentation. Drives labor cost. | Breakeven O31:O37
Target TUR for Cell Sizing | TUR each reference cell must maintain across the range. Drives auto cell/swap count. Lower = fewer cells, more rework. | Range Analysis B13
DECISION RULES EXPLAINED
Global (PFA ≤ 2%) | Guard band width set so average PFA across the population is ≤ 2%. Uses bisection solver (10 iterations). Most permissive — widest acceptance zone.
Method 6 (Dobbert) | Guard band from the Dobbert formula: M = max(0, 1.04 − e^(0.38·ln(TUR) − 0.54)). At TUR ≥ 4.6:1, guard band = 0. ANSI/NCSL Z540.3 default.
Specific Risk | Guard band = 0.98 × U. Most conservative — controls PFA to ≤ 2.5% per side at each individual measurement. Highest rework cost.
Simple Acceptance (Shared Risk) | No guard band (GB = 0). Full tolerance as acceptance limit. Lowest PFR but highest PFA — customer bears all risk. Up to 50% PFA at low TUR. Toggle: Dashboard B12. Default: hidden.
HOW TO USE THE CMC OPTIMIZER
1. Go to Range Analysis, rows 14-20 (CMC OPTIMIZER section).
2. Enter your desired TUR at the low force point in B17. The calculator shows the Required CMC and U.
3. It also shows the TUR at the high force point with the same CMC — this tells you how much headroom you have.
4. The CMC Sensitivity table (rows 52-68) shows annual cost at 14 CMC levels.
5. "Savings vs BCM" (col M) shows cumulative savings from upgrading beyond 0.04% CMC.
6. "Savings per 0.01% CMC" (col N) shows the rate of return — when this drops, further improvement yields diminishing returns.
7. The Cell Sizing Calculator (rows 75+) shows how many reference cells each machine needs at the Target TUR.
CELL SIZING & STANDARD SWITCHING
The calculator auto-sizes reference cells using a CMC curve model estimated from CMC@High and CMC@Low.
It calculates the force point where each machine's CMC crosses the Required CMC (= Tolerance / Target TUR).
From this "swap force," it determines cells needed to cover the full range.
When "Standard Switching Required?" = Yes, Swaps = Cells − 1. When No, Swaps = 0 (matched capacity).
Change Target TUR on Range Analysis B13 to explore: lower TUR = fewer cells but more rework.
EOPR BASELINE
Default: 89%. This is the ANSI/NCSL Z540.3 threshold (PFA ≤ 2% at TUR 4.6:1) and the NASA standard.
EOPR should be based on: (1) OEM stated reliability at the cal interval, or (2) historical drift data when worse than OEM spec.
The commonly cited 85% originated from a U.S. Navy Corona fleet-wide assumption across all M&TE and does not reflect controlled force calibration equipment.
TECHNICAL NOTES
• PFA and PFR are unconditional probabilities computed via numerical integration of the bivariate normal (Sandia PSL method).
• Integration uses 40-point trapezoidal rule for PFR (accuracy < 0.05% vs scipy at TUR ≥ 1.5) and 30-point midpoint for PFA.
• EOPR determines the process standard deviation: σ_product = Tolerance / NORMINV((1+EOPR)/2).
• All four decision rules use the same EOPR and the same integral — only the guard band formula differs. Simple Acceptance uses GB = 0.
• Dashboard inputs (Tolerance, CMC, EOPR, labor, cals) flow to Range Analysis and Risk Calc. CMC@Low and Tol@Low are separate RA inputs for the dual-point analysis.
• Cell sizing uses a 1/F + offset CMC model estimated from the two data points (CMC@High and CMC@Low).
• No Excel LAMBDA functions are used. Compatible with Excel 2016+ and LibreOffice (use NORMSINV, not NORM.S.INV).
• Validated against Sandia PSL PFA Risk Calculator (SAND2025-10736O) and scipy.integrate.dblquad.

## Dashboard
The True Cost of Measurement Decisions
Calculator from Morehouse Instrument Company
Risk calculated via Sandia PSL bivariate normal integral (no lookup tables).
▶ START HERE — Enter values in GREEN cells below, then check Range Analysis for any Machine Pricing Adjustment + Summary tab for report |  |  |  | 💡 Hover any green input cell for a description, or see the Instructions tab for full guidance.
COST INPUTS
Hourly Labor Rate ($/hr) | 200
Rework Time per Event (minutes) | 35
Rework Cost per Event ($) | 116.6667
Calibrations per Year | 500
EOPR (In-Tolerance Probability) | 89%
Show Simple Acceptance? | Yes
MEASUREMENT INPUTS
Device Tolerance (%) | 0.1%
Applied Force (lbf) | 2000
Tolerance in Force Units (lbf) | 2
SWITCHING OPTIONS
Standard Switching Required? | No
Swap Time per Change (min) | 30
PROFITABILITY INPUTS
Cal Cost per Cell ($/yr) | 2000
Price per Calibration ($) | 300
Target Payback (years) | 3
Machine | CMC (%) | U (lbf) | TUR | Price ($) | CMC Low (%) | Hrs/Cal | GB Global | PFA Glob | PFR Glob | RW Glob | GB M6 | PFA M6 | PFR M6 | RW M6 | GB Spec | PFA Spec | PFR Spec | RW Spec | FR/yr Glob | FR/yr M6 | FR/yr Spec | Max Cap (lbf) | Available? | PFR Simple | RW Simple
BCM | 0.05% | 1 | 2 | 25000 | 0.06% | 0.9 | 0.075605 | 1.997% | 6.375% | 3718.743 | 0.28165 | 1.135% | 10.4% | 6066.4961 | 0.98 | 0.06484% | 33.98% | 19822.3575 | 31.8749 | 51.9985 | 169.9059 | 10000 | ✓ | 5.218% | 3044.0993
PCM | 0.033% | 0.66 | 3.0303 | 35000 | 0.03% | 0.9 | 0 | 1.77% | 3.062% | 1786.3811 | 0.10028 | 1.199% | 4.472% | 2608.5204 | 0.6468 | 0.0439% | 18.63% | 10865.2552 | 15.3118 | 22.3587 | 93.1308 | 2000 | ✓ | 3.062% | 1786.3811
UCM Manual | 0.03% | 0.6 | 3.3333 | 50000 | 0.03% | 1 | 0 | 1.64% | 2.727% | 1590.479 | 0.071508 | 1.212% | 3.675% | 2143.8717 | 0.588 | 0.03977% | 16.3% | 9509.2584 | 13.6327 | 18.376 | 81.5079 | 1000000 | ✓ | 2.727% | 1590.479
UCM Automated | 0.02% | 0.4 | 5 | 100000 | 0.025% | 0.9 | 0 | 1.155% | 1.704% | 994.2843 | 0 | 1.155% | 1.704% | 994.2843 | 0.392 | 0.02465% | 9.485% | 5533.1425 | 8.5224 | 8.5224 | 47.4269 | 1000000 | ✓ | 1.704% | 994.2843
Deadweight | 0.002% | 0.04 | 50 | 350000 | 0.002% | 1.2 | 0 | 4.028e-05% | 0.4758% | 277.5303 | 0 | 4.028e-05% | 0.4758% | 277.5303 | 0.0392 | 1.159e-09% | 0.9104% | 531.0811 | 2.3788 | 2.3788 | 4.5521 | 1000000 | ✓ | 0.4758% | 277.5303
Competitor 1 | 0.05% | 1 | 2 | 55000 | 0.12% | 1 | 0.075605 | 1.997% | 6.375% | 3718.743 | 0.28165 | 1.135% | 10.4% | 6066.4961 | 0.98 | 0.06484% | 33.98% | 19822.3575 | 31.8749 | 51.9985 | 169.9059 | 1000000 | ✓ | 5.218% | 3044.0993
Competitor 2 | 0.06% | 1.2 | 1.6667 | 40000 | 0.15% | 1 | 0 | 0% | 6.226% | 3632.0295 | 0.39889 | 0% | 14.77% | 8613.9616 | 1.176 | 0% | 44.33% | 25856.9044 | 31.1317 | 73.834 | 221.6306 | 1000000 | ✓ | 6.226% | 3632.0295
All PFA/PFR computed via Sandia PSL bivariate normal integral (trapezoidal rule, no lookup tables).
Global: guard band set so PFA ≤ target (bisection solver). M6: Dobbert formula. Specific: h=0.98×U.
EOPR determines σ_product. All three models use same EOPR and same integral—only guard band width differs.
CHART DATA
Machine | Global (PFA≤2%) | Method 6 | Specific Risk |  |  |  | Simple Accept
BCM | 3718.743 | 6066.4961 | 19822.3575 |  |  |  | 3044.0993
PCM | 1786.3811 | 2608.5204 | 10865.2552 |  |  |  | 1786.3811
UCM Manual | 1590.479 | 2143.8717 | 9509.2584 |  |  |  | 1590.479
UCM Automated | 994.2843 | 994.2843 | 5533.1425 |  |  |  | 994.2843
Deadweight | 277.5303 | 277.5303 | 531.0811 |  |  |  | 277.5303
Competitor 1 | 3718.743 | 6066.4961 | 19822.3575 |  |  |  | 3044.0993
Competitor 2 | 3632.0295 | 8613.9616 | 25856.9044 |  |  |  | 3632.0295
MACHINE PAYBACK DATA (capped at 20yr for display)
Machine | Payback (Global) | Payback (M6) | Payback (Specific) |  |  |  | Payback (Simple)
BCM | 0.46209 | 0.5076 | 0.77809 |  |  |  | 0.45491
PCM | 0.62262 | 0.63187 | 0.74255 |  |  |  | 0.62262
UCM Manual | 1.0774 | 1.0904 | 1.299 |  |  |  | 1.0774
UCM Automated | 1.7631 | 1.7689 | 1.9769 |  |  |  | 1.7542
Deadweight | 11.7756 | 11.7756 | 11.8769 |  |  |  | 11.7756
Competitor | 1.4588 | 3.3355 | 0 |  |  |  | 1.2234
Competitor 2 | 1.1695 | 8.5043 | 0 |  |  |  | 0.90155
ANNUAL PROFIT DATA
Machine | Profit (Global) | Profit (M6) | Profit (Specific) |  |  |  | Profit (Simple)
BCM | 54102.0515 | 49251.4383 | 32129.8014 |  |  |  | 54955.9007
PCM | 56213.6189 | 55391.4796 | 47134.7448 |  |  |  | 56213.6189
UCM Manual | 46409.521 | 45856.1283 | 38490.7416 |  |  |  | 46409.521
UCM Automated | 56718.575 | 56533.8752 | 50584.3388 |  |  |  | 57005.7157
Deadweight | 29722.4697 | 29722.4697 | 29468.9189 |  |  |  | 29722.4697
Competitor | 37701.3519 | 16489.5233 | 0 |  |  |  | 44955.9007
Competitor 2 | 34202.9096 | 4703.5061 | 0 |  |  |  | 44367.9705
STANDARD SWITCHING SUMMARY
Min Loading | Cells | Swaps/Run | Swap $/Cal |  |  |  | Swap $/yr | Cal $/yr | Total M6/yr
5% (UPC) | 1 | 0 | 0 |  |  |  | 0 | 2000 | 4662.8333
10% (PC) | 2 | 1 | 100 |  |  |  | 0 | 4000 | 5466.1248
20% (Commercial) | 3 | 2 | 200 |  |  |  | 0 | 6000 | 6994.2843
Deadweight | 0 | 0 | 0 |  |  |  | 0 | 0 | 0
RECOMMENDATION HELPERS (hidden from print)
Machine | Available? | Profit M6 | Payback M6 |  |  |  | Profit Rank | Worst Helper | TUR Helper | TUR Max Help
BCM | TRUE | 54955.9007 | 0.45491 |  |  |  | 54955.9007 | 54955.9007 | 2 | 2
PCM | TRUE | 56213.6189 | 0.62262 |  |  |  | 56213.6189 | 56213.6189 | 3.0303 | 3.0303
UCM Manual | TRUE | 46409.521 | 1.0774 |  |  |  | 46409.521 | 46409.521 | 3.3333 | 3.3333
UCM Automated | TRUE | 57005.7157 | 1.7542 |  |  |  | 57005.7157 | 57005.7157 | 5 | 5
Deadweight | TRUE | 29722.4697 | 11.7756 |  |  |  | 29722.4697 | 29722.4697 | 50 | 50
Competitor 1 | TRUE | 44955.9007 | 1.2234 |  |  |  | 44955.9007 | 44955.9007 | 2 | 2
Competitor 2 | TRUE | 44367.9705 | 0.90155 |  |  |  | 44367.9705 | 44367.9705 | 1.6667 | 1.6667
Available count | 7
Profitable (avail) | 7
Best profit | 57005.7157 | UCM Automated | 1.7542
Worst profit (avail) | 29722.4697 | Deadweight
Min TUR (avail) | 1.6667
Max TUR (avail) | 50
REQUIRED $/CAL TO HIT 3-YEAR PAYBACK
Machine | Machine Price | Current $/Cal | Global (PFA ≤ 2%) | Method 6 (Dobbert) | Specific Risk (PFA ≤ 2%) | Simple Acceptance | Capacity (lbf)
BCM | 25000 | 300 | 208.4626 | 218.1638 | 252.4071 | 206.7549 | 10000
PCM | 35000 | 300 | 210.9061 | 212.5504 | 229.0638 | 210.9061 | 2000
UCM Manual | 50000 | 300 | 240.5143 | 241.6211 | 256.3519 | 240.5143 | 1000000
UCM Automated | 100000 | 300 | 253.2295 | 253.5989 | 265.498 | 252.6552 | 1000000
Deadweight | 350000 | 300 | 473.8884 | 473.8884 | 474.3955 | 473.8884 | 1000000
Competitor 1 | 55000 | 300 | 261.264 | 303.6876 | 344.5 | 246.7549 | 1000000
Competitor 2 | 40000 | 300 | 258.2608 | 317.2597 | 334.5 | 237.9307 | 1000000
Note on A12: henry zumbrun: Simple Acceptence may not be applicable as in most cases it does not account for measurement uncertainty.
Note on A18: henry zumbrun: If you intend to use multiple load cells, check yes

## Range Analysis
Range Analysis — Full Cost at High and Low Force Points
Enter your operating range, CMC at each end, swaps per machine. Risk via Sandia bivariate integral.
RANGE INPUTS |  |  |  | DERIVED
High Force (lbf) | 2000 |  |  | Tol @ High (lbf) | 2
Low Force (lbf) | 500 |  |  | Tol @ Low (lbf) | 0.5
Tolerance @ High (%) | 0.1% |  |  | σ_p @ High | 1.2514
Tolerance @ Low (%) | 0.1% |  |  | σ_p @ Low | 0.31285
EOPR | 89%
Calibrations per Year | 500
Rework Cost per Event ($) | 116.6667
Swap Time per Change (min) | 30
Target TUR for Cell Sizing | 4 | Forward sizing: pick the TUR each reference cell must hold across the range. Required CMC = Tolerance (0.100 %) ÷ Target TUR. Lower TUR = fewer cells but more rework risk.
CMC OPTIMIZER
TARGET TUR REVERSE CALCULATOR
Desired TUR at Low Force | 4 | To achieve your target TUR at the low force point,
Required CMC (%) | 0.025% | you need a reference standard with this CMC or better.
Required U at Low (lbf) | 0.125
TUR at High Force (same CMC) | 4
What-If Tolerance (%) | 0.025% | Enter a second tolerance to compare (e.g., 0.025% for reference-grade)
MACHINE COMPARISON
Machine | Price ($) | Hrs/Cal | Cells | Swaps/Run | CMC @ High (%) | U @ High | TUR @ High | CMC @ Low (%) | U @ Low | TUR @ Low | PFR Glob @ Worst | PFR M6 @ Worst | PFR Spec @ Worst | RW Glob | RW M6 | RW Spec | Cal $/yr | Swap $/yr | Total Glob | Total M6 | Total Spec | What-If Tol (lbf) | TUR @ What-If Hi | TUR @ What-If Lo | RW M6 @ What-If | Total M6 @ What-If | CMC Model a | CMC Model b | Req CMC (%) | Swap Force (lbf) | Auto Cells | Auto Swaps
BCM | 25000 | 0.9 | 1 | 0 | 0.0005 | 1 | 2 | 0.06% | 0.3 | 1.6667 | 6.682% | 15% | 44.35% | 3897.9485 | 8748.5617 | 25870.1986 | 2000 | 0 | 5897.9485 | 10748.5617 | 27870.1986 | 0.5 | 0.5 | 0.41667 | N/A — TUR<1 | Not Viable | 0.066667 | 0.00046667 | 0.025% | 2000 | 1 | 0
PCM | 35000 | 0.9 | 1 | 0 | 0.00033 | 0.66 | 3.0303 | 0.03% | 0.15 | 3.3333 | 3.062% | 4.472% | 18.63% | 1786.3811 | 2608.5204 | 10865.2552 | 2000 | 0 | 3786.3811 | 4608.5204 | 12865.2552 | 0.5 | 0.75758 | 0.83333 | N/A — TUR<1 | Not Viable | -0.02 | 0.00034 | 0.025% | 2000 | 1 | 0
UCM Manual | 50000 | 1 | 1 | 0 | 0.0003 | 0.6 | 3.3333 | 0.03% | 0.15 | 3.3333 | 2.727% | 3.675% | 16.3% | 1590.479 | 2143.8717 | 9509.2584 | 2000 | 0 | 3590.479 | 4143.8717 | 11509.2584 | 0.5 | 0.83333 | 0.83333 | N/A — TUR<1 | Not Viable | 0 | 0.0003 | 0.025% | 2000 | 1 | 0
UCM Automated | 100000 | 0.9 | 1 | 0 | 0.0002 | 0.4 | 5 | 0.025% | 0.125 | 4 | 2.197% | 2.513% | 12.71% | 1281.425 | 1466.1248 | 7415.6612 | 2000 | 0 | 3281.425 | 3466.1248 | 9415.6612 | 0.5 | 1.25 | 1 | Very High | See Calculator | 0.033333 | 0.00018333 | 0.025% | 500 | 1 | 0
Deadweight | 350000 | 1.2 | 0 | 0 | 2e-05 | 0.04 | 50 | 0.002% | 0.01 | 50 | 0.4758% | 0.4758% | 0.9104% | 277.5303 | 277.5303 | 531.0811 | 0 | 0 | 277.5303 | 277.5303 | 531.0811 | 0.5 | 12.5 | 12.5 | Low | See Calculator | 0 | 0 | 0 | 0 | 0 | 0
Competitor | 55000 | 1 | 1 | 0 | 0.0005 | 1 | 2 | 0.12% | 0.6 | 0.83333 | 17.65% | 54.02% | 89% | 10298.6481 | 31510.4767 | 51916.6667 | 2000 | 0 | 12298.6481 | 33510.4767 | 53916.6667 | 0.5 | 0.5 | 0.20833 | N/A — TUR<1 | Not Viable | 0.46667 | 0.00026667 | 0.025% | 2000 | 1 | 0
Competitor 2 | 40000 | 1 | 1 | 0 | 0.0006 | 1.2 | 1.6667 | 0.15% | 0.75 | 0.66667 | 23.65% | 74.22% | 89% | 13797.0904 | 43296.4939 | 51916.6667 | 2000 | 0 | 15797.0904 | 45296.4939 | 53916.6667 | 0.5 | 0.41667 | 0.16667 | N/A — TUR<1 | Not Viable | 0.6 | 0.0003 | 0.025% | 2000 | 1 | 0
Note: CMC@High (col F) and Tolerance@High (0.100%) are wired to Dashboard. CMC@Low (col I), Hrs/Cal (col C), and Price (col B) are independent — set per machine.
CHART DATA
Machine | Global |  | Method 6 | Specific Risk | Simple Accept
BCM | 5897.9485 |  | 10748.5617 | 27870.1986 | 5044.0993
PCM | 3786.3811 |  | 4608.5204 | 12865.2552 | 3786.3811
UCM Manual | 3590.479 |  | 4143.8717 | 11509.2584 | 3590.479
UCM Automated | 3281.425 |  | 3466.1248 | 9415.6612 | 2994.2843
Deadweight | 277.5303 |  | 277.5303 | 531.0811 | 277.5303
Competitor | 12298.6481 |  | 33510.4767 | 53916.6667 | 5044.0993
Competitor 2 | 15797.0904 |  | 45296.4939 | 53916.6667 | 5632.0295
Notes:
PFR computed at both force points (2,000 and 500 lbf). Rework uses worst case (highest PFR). Blue cells are editable.
Swap cost = Swaps × (30 min / 60) × $200/hr × 500 cals/yr. Reset swaps to 0 if not switching standards.
CMC SENSITIVITY TABLE
CMC SENSITIVITY — Annual Cost vs. CMC at Low Force Point
Shows how total annual cost changes as CMC improves. Uses low force point (500 lbf, worst case). Cal cost = $2,000/cell.
CMC (%) | U (lbf) |  | TUR | PFR Global | PFR M6 | PFR Specific | Rework Global | Rework M6 | Rework Specific | Total Global* | Total M6* | Total Specific* | Savings vs BCM (M6) | Savings per 0.01% CMC
0.1% | 0.5 |  | 1 | 13.68% | 39.81% | 87.03% | 7979.6115 | 23220.9089 | 50766.6406 | 9979.6115 | 25220.9089 | 52766.6406 | 0 | —
0.08% | 0.4 |  | 1.25 | 9.963% | 26.36% | 66.21% | 5811.8305 | 15378.0063 | 38621.1314 | 7811.8305 | 17378.0063 | 40621.1314 | 0 | 3921.4513
0.06% | 0.3 |  | 1.6667 | 6.645% | 14.97% | 44.35% | 3876.0848 | 8734.2153 | 25868.0421 | 5876.0848 | 10734.2153 | 27868.0421 | 0 | 3321.8955
0.05% | 0.25 |  | 2 | 5.177% | 10.37% | 33.98% | 3019.9998 | 6048.9533 | 19819.8806 | 5019.9998 | 8048.9533 | 21819.8806 | 0 | 2685.262
0.04% | 0.2 |  | 2.5 | 3.851% | 6.583% | 24.52% | 2246.5387 | 3839.8325 | 14304.0723 | 4246.5387 | 5839.8325 | 16304.0723 | 0 | 2209.1209
0.035% | 0.175 |  | 2.8571 | 3.243% | 5% | 20.24% | 1891.4798 | 2916.6009 | 11806.4836 | 3891.4798 | 4916.6009 | 13806.4836 | 923.2316 | 1846.4632
0.03% | 0.15 |  | 3.3333 | 2.67% | 3.623% | 16.3% | 1557.4008 | 2113.4652 | 9505.5201 | 3557.4008 | 4113.4652 | 11505.5201 | 1726.3673 | 1606.2714
0.025% | 0.125 |  | 4 | 2.132% | 2.451% | 12.71% | 1243.8287 | 1429.6027 | 7411.307 | 3243.8287 | 3429.6027 | 9411.307 | 2410.2298 | 1367.7249
0.02% | 0.1 |  | 5 | 1.628% | 1.628% | 9.476% | 949.8619 | 949.8619 | 5527.8882 | 2949.8619 | 2949.8619 | 7527.8882 | 2889.9706 | 959.4816
0.015% | 0.075 |  | 6.6667 | 1.155% | 1.155% | 6.606% | 673.7478 | 673.7478 | 3853.4382 | 2673.7478 | 2673.7478 | 5853.4382 | 3166.0847 | 552.2283
0.01% | 0.05 |  | 10 | 0.7042% | 0.7042% | 4.082% | 410.7959 | 410.7959 | 2380.9001 | 2410.7959 | 2410.7959 | 4380.9001 | 3429.0366 | 525.9037
0.005% | 0.025 |  | 20 | 0.2184% | 0.2184% | 1.939% | 127.3824 | 127.3824 | 1131.0667 | 2127.3824 | 2127.3824 | 3131.0667 | 3712.45 | 566.8269
0.003% | 0.015 |  | 33.3333 | 0.01703% | 0.01703% | 1.005% | 9.9353 | 9.9353 | 586.4531 | 2009.9353 | 2009.9353 | 2586.4531 | 3829.8972 | 587.2359
0.002% | 0.01 |  | 50 | 0% | 0% | 0.2018% | 0 | 0 | 117.7238 | 2000 | 2000 | 2117.7238 | 3839.8325 | 99.3526
* Total = Rework + $2,000 cal (1 cell). Swap cost excluded to isolate measurement capability impact.
Savings vs BCM shows cumulative annual savings from upgrading beyond 0.04% CMC. When Savings per 0.01% CMC flattens, further improvement has diminishing returns.
CELL SIZING CALCULATOR
How many reference cells does each machine need to cover the full range at the Target TUR? Based on CMC curve interpolation from CMC@High and CMC@Low.
Machine | CMC@Hi |  | CMC@Lo | Req CMC | Swap Force | Cells | Swaps | Cal $/yr | Swap $/yr | Total $/yr
BCM | 0.05% |  | 0.06% | 0.025% | N/V | 1 | 0 | 2000 | 0 | 2000
PCM | 0.033% |  | 0.03% | 0.025% | N/V | 1 | 0 | 2000 | 0 | 2000
UCM Manual | 0.03% |  | 0.03% | 0.025% | N/V | 1 | 0 | 2000 | 0 | 2000
UCM Automated | 0.02% |  | 0.025% | 0.025% | 500 | 1 | 0 | 2000 | 0 | 2000
Deadweight | 0.002% |  | 0.002% | 0% | 0 | 0 | 0 | 0 | 0 | 0
Competitor | 0.05% |  | 0.12% | 0.025% | N/V | 1 | 0 | 2000 | 0 | 2000
Competitor 2 | 0.06% |  | 0.15% | 0.025% | N/V | 1 | 0 | 2000 | 0 | 2000
Target TUR (B13) controls how far down in a cell's range the CMC remains adequate. Lower TUR targets need fewer cells but accept more rework. The CMC model estimates CMC at any force from the two data points (CMC@High and CMC@Low) using a 1/F + offset curve. "N/V" means the machine cannot achieve the target TUR even at its best point.
When "Standard Switching Required?" = Yes on Dashboard, swaps are calculated. When No (matched capacity), cells are still needed but swaps = 0 — you buy matched reference cells for each sub-range.
EOPR SENSITIVITY — How Recommended CMC Changes with EOPR
Shows the optimal CMC at the low force point for different EOPR assumptions.
EOPR | σ_product |  | Req CMC for TUR 4:1 | M6 Rework @ TUR 3 | M6 Rework @ TUR 4 | M6 Rework @ TUR 5
70% | 0.48242 |  | 5e-05% | 3012.8502 | 1720.9601 | 1176.4231
75% | 0.43465 |  | 5e-05% | 3036.415 | 1720.1777 | 1170.5175
80% | 0.39015 |  | 5e-05% | 2991.3398 | 1676.3941 | 1133.9191
85% | 0.34734 |  | 5e-05% | 2850.3373 | 1573.4862 | 1055.5358
90% | 0.30398 |  | 5e-05% | 2562.0808 | 1381.8741 | 915.3268
95% | 0.25511 |  | 5e-05% | 2000.2361 | 1031.4512 | 666.8752
At lower EOPR, more items are near the tolerance edge → PFR increases → rework goes up → the optimal CMC shifts lower (need better equipment).
REFERENCE CELL COVERAGE CALCULATOR
How many reference cells do you need to cover your range at your target TUR?
INPUTS
High Force (lbf) | 2000
Low Force (lbf) | 500
Target TUR | 4
Min Loading (%) | 5%
Device Tolerance (%) | 0.1%
RESULTS
Required CMC at Low Force (%) | 0.025%
Number of Cells Needed | 1
Swaps per Run | 0
Annual Swap Cost | 0
CELL CHAIN DETAIL
Cell # | Capacity (lbf) |  | Covers Down To | TUR at Bottom | Covers Low?
1 | 2000 |  | 500 | 4 | ✓ Yes
2
3
4
5
6
HOW THIS WORKS
Cell 1 starts at High Force and covers down to Capacity × Min Loading %.
If that doesn't reach Low Force, Cell 2 picks up where Cell 1 left off.
Each additional cell = one swap per calibration run.
At 5%: 10,000 → 500 → 25 = 2 cells, 1 swap. At 10%: 3 cells. At 20%: 4+ cells.
Better load cells (lower min loading %) = fewer cells = fewer swaps = lower cost.
EOPR SENSITIVITY — How Recommended CMC Changes with EOPR
Shows M6 rework at the low force point for different EOPR assumptions.
EOPR | σ_product |  | Req CMC for TUR 4:1 | M6 Rework @ TUR 3 | M6 Rework @ TUR 4 | M6 Rework @ TUR 5
70% | 0.48242 |  | 5e-05% | 3012.8502 | 1720.9601 | 1176.4231
75% | 0.43465 |  | 5e-05% | 3036.415 | 1720.1777 | 1170.5175
80% | 0.39015 |  | 5e-05% | 2991.3398 | 1676.3941 | 1133.9191
85% | 0.34734 |  | 5e-05% | 2850.3373 | 1573.4862 | 1055.5358
90% | 0.30398 |  | 5e-05% | 2562.0808 | 1381.8741 | 915.3268
95% | 0.25511 |  | 5e-05% | 2000.2361 | 1031.4512 | 666.8752
At lower EOPR, more items are near the tolerance edge → PFR increases → you need better equipment to achieve the same cost.
Note on C24: henry zumbrun: Avg hours per calibration for BCM. Include setup, loading, data collection, documentation.
Note on C25: henry zumbrun: Avg hours per calibration for PCM.
Note on C26: henry zumbrun: Avg hours per calibration for UCM Manual.
Note on C27: henry zumbrun: Avg hours per calibration for UCM Automated.
Note on C28: henry zumbrun: Avg hours per calibration for Deadweight.
Note on C29: henry zumbrun: Avg hours per calibration for Competitor 1.
Note on C30: henry zumbrun: Avg hours per calibration for Competitor 2.

## Breakeven
Breakeven Analysis — Machine Investment Payback
Prices from Range Analysis. Savings = difference in total annual cost. Toggle swap cost on Dashboard.
MACHINE ANNUAL COST SUMMARY
Machine | Price | Total Glob/yr | Total M6/yr | Total Spec/yr
BCM | 25000 | 5897.9485 | 10748.5617 | 27870.1986
PCM | 35000 | 3786.3811 | 4608.5204 | 12865.2552
UCM Manual | 50000 | 3590.479 | 4143.8717 | 11509.2584
UCM Automated | 100000 | 3281.425 | 3466.1248 | 9415.6612
Deadweight | 350000 | 277.5303 | 277.5303 | 531.0811
Competitor | 55000 | 12298.6481 | 33510.4767 | 53916.6667
Competitor 2 | 40000 | 15797.0904 | 45296.4939 | 53916.6667
UPGRADE DELTAS (auto-calculated from prices above)
Path | Price Delta | M6 Savings/yr | Payback (M6, yrs)
Competitor 1 → BCM | -30000 | 22761.915 | -1.318
Competitor 1 → UCM | -5000 | 29366.605 | -0.17026
Competitor 2 → UCM | 10000 | 41152.6222 | 0.243
BCM → UCM Manual | 25000 | 6604.69 | 3.7852
BCM → UCM Auto | 75000 | 7282.4369 | 10.2988
UCM Manual → Deadweight | 300000 | 3866.3414 | 77.5927
PROFITABILITY MODEL
Price per Calibration ($) | 300
Target Payback (years) | 3
Calibrations per Year | 500
Hourly Labor Rate ($/hr) | 200
PROFITABILITY PER MACHINE
Machine | Price | Revenue/yr | Cost Glob | Cost M6 | Cost Spec | Profit Glob | Profit M6 | Profit Spec | Payback Glob | Payback M6 | Payback Spec | Req Price (Glob) | Req Price (M6) | Hrs/Cal | Labor Cost/yr | Req Price (Spec) | Req $/Hr (M6) | Cost Simple | Profit Simple | Payback Simple
BCM | 25000 | 150000 | 5897.9485 | 10748.5617 | 27870.1986 | 54102.0515 | 49251.4383 | 32129.8014 | 0.46209 | 0.5076 | 0.77809 | 208.4626 | 218.1638 | 0.9 | 90000 | 252.4071 | 242.4042 | 5044.0993 | 54955.9007 | 0.45491
PCM | 35000 | 150000 | 3786.3811 | 4608.5204 | 12865.2552 | 56213.6189 | 55391.4796 | 47134.7448 | 0.62262 | 0.63187 | 0.74255 | 210.9061 | 212.5504 | 0.9 | 90000 | 229.0638 | 236.1671 | 3786.3811 | 56213.6189 | 0.62262
UCM Manual | 50000 | 150000 | 3590.479 | 4143.8717 | 11509.2584 | 46409.521 | 45856.1283 | 38490.7416 | 1.0774 | 1.0904 | 1.299 | 240.5143 | 241.6211 | 1 | 100000 | 256.3519 | 241.6211 | 3590.479 | 46409.521 | 1.0774
UCM Automated | 100000 | 150000 | 3281.425 | 3466.1248 | 9415.6612 | 56718.575 | 56533.8752 | 50584.3388 | 1.7631 | 1.7689 | 1.9769 | 253.2295 | 253.5989 | 0.9 | 90000 | 265.498 | 281.7766 | 2994.2843 | 57005.7157 | 1.7542
Deadweight | 350000 | 150000 | 277.5303 | 277.5303 | 531.0811 | 29722.4697 | 29722.4697 | 29468.9189 | 11.7756 | 11.7756 | 11.8769 | 473.8884 | 473.8884 | 1.2 | 120000 | 474.3955 | 394.907 | 277.5303 | 29722.4697 | 11.7756
Competitor | 55000 | 150000 | 12298.6481 | 33510.4767 | 53916.6667 | 37701.3519 | 16489.5233 | -3916.6667 | 1.4588 | 3.3355 | Never | 261.264 | 303.6876 | 1 | 100000 | 344.5 | 303.6876 | 5044.0993 | 44955.9007 | 1.2234
Competitor 2 | 40000 | 150000 | 15797.0904 | 45296.4939 | 53916.6667 | 34202.9096 | 4703.5061 | -3916.6667 | 1.1695 | 8.5043 | Never | 258.2608 | 317.2597 | 1 | 100000 | 334.5 | 317.2597 | 5632.0295 | 44367.9705 | 0.90155
Hrs/Cal pulled from Range Analysis tab. Edit there to change.
REQUIRED CALIBRATIONS FOR TARGET PAYBACK
Machine | Price | Req Cals (Glob) | Req Cals (M6) | Req Cals (Spec)
BCM | 25000 | 231.0448 | 253.7997 | 389.0469
PCM | 35000 | 311.3125 | 315.9331 | 371.276
UCM Manual | 50000 | 538.6826 | 545.1834 | 649.5068
UCM Automated | 100000 | 881.5454 | 884.4255 | 988.4482
Deadweight | 350000 | 5887.8014 | 5887.8014 | 5938.4601
Competitor | 55000 | 729.4168 | 1667.7256 | N/A
Competitor 2 | 40000 | 584.7456 | 4252.1472 | N/A
CHART DATA — ANNUAL PROFIT
Machine | Profit (Global) | Profit (M6) | Profit (Specific) | Profit (Simple)
BCM | 54102.0515 | 49251.4383 | 32129.8014 | 54955.9007
PCM | 56213.6189 | 55391.4796 | 47134.7448 | 56213.6189
UCM Manual | 46409.521 | 45856.1283 | 38490.7416 | 46409.521
UCM Automated | 56718.575 | 56533.8752 | 50584.3388 | 57005.7157
Deadweight | 29722.4697 | 29722.4697 | 29468.9189 | 29722.4697
Competitor | 37701.3519 | 16489.5233 | 0 | 44955.9007
Competitor 2 | 34202.9096 | 4703.5061 | 0 | 44367.9705
Machine | Payback (Global) | Payback (M6) | Payback (Specific) | Payback (Simple)
BCM | 0.46209 | 0.5076 | 0.77809 | 0.45491
PCM | 0.62262 | 0.63187 | 0.74255 | 0.62262
UCM Manual | 1.0774 | 1.0904 | 1.299 | 1.0774
UCM Automated | 1.7631 | 1.7689 | 1.9769 | 1.7542
Deadweight | 11.7756 | 11.7756 | 11.8769 | 11.7756
Competitor | 1.4588 | 3.3355 | 0 | 1.2234
Competitor 2 | 1.1695 | 8.5043 | 0 | 0.90155
PROFITABILITY vs TUR (Scatter Data)
Machine | TUR | Profit Global | TUR | Profit M6 | TUR | Profit Specific | TUR | Profit Simple
BCM | 2 | 54102.0515 | 2 | 49251.4383 | 2 | 32129.8014 | 2 | 54955.9007
PCM | 3.0303 | 56213.6189 | 3.0303 | 55391.4796 | 3.0303 | 47134.7448 | 3.0303 | 56213.6189
UCM Manual | 3.3333 | 46409.521 | 3.3333 | 45856.1283 | 3.3333 | 38490.7416 | 3.3333 | 46409.521
UCM Automated | 5 | 56718.575 | 5 | 56533.8752 | 5 | 50584.3388 | 5 | 57005.7157
Deadweight | 50 | 29722.4697 | 50 | 29722.4697 | 50 | 29468.9189 | 50 | 29722.4697
Competitor | 2 | 37701.3519 | 2 | 16489.5233 | 2 | -3916.6667 | 2 | 44955.9007
Competitor 2 | 1.6667 | 34202.9096 | 1.6667 | 4703.5061 | 1.6667 | -3916.6667 | 1.6667 | 44367.9705
$0 Line | 0 | 0
| 60 | 0
Each dot is a machine at its TUR. Points above the red $0 line are profitable; below are losses. The vertical spread at each TUR shows the cost of guard banding — Simple Accept (no guard band) is always the highest profit but carries the most customer risk.
Notes:
Revenue = Calibrations/yr × Price per Calibration.
Labor Cost = Hrs/Cal × Hourly Rate × Cals/yr. Edit Hrs/Cal per machine (green column) to reflect your actual calibration time.
Operating Cost = (Rework + Swap + Ref Cal) from Range Analysis. Does not include consumables or facility overhead.
Profit = Revenue − Labor − Operating Cost. Negative profit means the machine loses money at that price per cal.
Payback = Machine Price / Annual Profit. "Never" means the machine never pays for itself at current pricing.
Automation reduces Hrs/Cal directly, which is the single biggest lever on profitability (bigger than rework savings).

## Guard Band Analysis
Guard Band Tax Calculator
Shows tolerance consumed at various TUR levels
TUR | U (lbf) | GB Spec | Tol% Spec | Usable Spec | GB M6 | Tol% M6 | Usable M6 | GB Simple | Tol% Simple | Usable Simple
1.5 | 1.3333 | 1.3067 | 65.33% | 0.69333 | 0.48024 | 24.01% | 1.5198 | 0 | 0% | 2
2 | 1 | 0.98 | 49% | 1.02 | 0.28165 | 14.08% | 1.7184 | 0 | 0% | 2
2.5 | 0.8 | 0.784 | 39.2% | 1.216 | 0.17163 | 8.581% | 1.8284 | 0 | 0% | 2
3 | 0.66667 | 0.65333 | 32.67% | 1.3467 | 0.10355 | 5.177% | 1.8965 | 0 | 0% | 2
3.5 | 0.57143 | 0.56 | 28% | 1.44 | 0.058256 | 2.913% | 1.9417 | 0 | 0% | 2
4 | 0.5 | 0.49 | 24.5% | 1.51 | 0.026561 | 1.328% | 1.9734 | 0 | 0% | 2
4.6 | 0.43478 | 0.42609 | 21.3% | 1.5739 | 0 | 0% | 2 | 0 | 0% | 2
5 | 0.4 | 0.392 | 19.6% | 1.608 | 0 | 0% | 2 | 0 | 0% | 2
10 | 0.2 | 0.196 | 9.8% | 1.804 | 0 | 0% | 2 | 0 | 0% | 2
24 | 0.083333 | 0.081667 | 4.083% | 1.9183 | 0 | 0% | 2 | 0 | 0% | 2

## Standard Switching
Standard Switching Cost Analysis
Risk from Sandia bivariate integral
INPUTS
Swap Time (min) | 30
Cal Cost/Cell ($/yr) | 2000
Downtime (weeks) | 2.5
Min Load% | #Cells | Swaps | SwapCost/Cal | AnnSwapCost | AnnCalCost | DT(wks) | TUR @0.1% | PFR Spec | RW Spec | PFR M6 | RW M6 | Total Spec | Total M6 | Ship Events | PFR Glob | RW Glob | Total Glob
5% | 1 | 0 | 0 | 0 | 2000 | 2.5 | 3 | 18.89% | 11020.4502 | 4.565% | 2662.8333 | 13020.4502 | 4662.8333 | 1 | 3.101% | 1808.6326 | 3808.6326
10% | 2 | 1 | 100 | 0 | 4000 | 2.5 | 4 | 12.71% | 7415.6612 | 2.513% | 1466.1248 | 11415.6612 | 5466.1248 | 2 | 2.197% | 1281.425 | 5281.425
20% | 3 | 2 | 200 | 0 | 6000 | 2.5 | 5 | 9.485% | 5533.1425 | 1.704% | 994.2843 | 11533.1425 | 6994.2843 | 3 | 1.704% | 994.2843 | 6994.2843
Deadweight | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Notes:
PFR via Sandia bivariate integral. M6 always ≤ Specific at same TUR.
Shipping Risk = 1 event/cell/yr in proper Morehouse case with custom-cut foam.
CHART DATA
Strategy | Swap Cost | Cal Cost | Rework (M6) | Total (M6)
5% (UPC) | 0 | 2000 | 2662.8333 | 4662.8333
10% (PC) | 0 | 4000 | 1466.1248 | 5466.1248
20% (Commercial) | 0 | 6000 | 994.2843 | 6994.2843
Deadweight | 0 | 0 | 0 | 0

## Risk Calc
Sandia PSL Bivariate Normal Risk Calculator
Based on Sandia PSL PFA Risk Calculator (SAND2025-10736O). Uses ERF-based Drezner approximation.
SHARED INPUTS
Tolerance (± lbf) | 2
EOPR | 89%
σ_product | 1.2514
N (cals/yr) | 500
Rework cost ($) | 116.6667
BVN Constants
cx | -1.095
cy | -0.75651
Machine |  | BCM | PCM | UCM Manual | UCM Automated | Deadweight | Interface Machine |  | 5% Load (TUR 3) | 10% Load (TUR 4) | 20% Load (TUR 5)
CMC (%) |  | 0.05% | 0.033% | 0.03% | 0.02% | 0.002% | 0.05% |  | 0 | 0 | 0
U (lbf) |  | 1 | 0.66 | 0.6 | 0.4 | 0.04 | 1 |  | 0.66667 | 0.5 | 0.4
σ_t |  | 0.5 | 0.33 | 0.3 | 0.2 | 0.02 | 0.5 |  | 0.33333 | 0.25 | 0.2
TUR |  | 2 | 3.0303 | 3.3333 | 5 | 50 | 2 |  | 3 | 4 | 5
σ_total |  | 1.3476 | 1.2942 | 1.2869 | 1.2673 | 1.2516 | 1.3476 |  | 1.295 | 1.2761 | 1.2673
ρ |  | 0.92862 | 0.96694 | 0.97245 | 0.98747 | 0.99987 | 0.92862 |  | 0.96631 | 0.98062 | 0.98747
d = √(1-ρ²) |  | 0.37103 | 0.25499 | 0.23312 | 0.15782 | 0.01598 | 0.37103 |  | 0.25739 | 0.1959 | 0.15782
a_bvn = -ρ/d |  | -2.5028 | -3.7922 | -4.1714 | -6.2571 | -62.5707 | -2.5028 |  | -3.7542 | -5.0057 | -6.2571
GB Specific |  | 0.98 | 0.6468 | 0.588 | 0.392 | 0.0392 | 0.98 |  | 0.65333 | 0.49 | 0.392
Accept width (a_spec) |  | 1.02 | 1.3532 | 1.412 | 1.608 | 1.9608 | 1.02 |  | 1.3467 | 1.51 | 1.608
p1 = a/σ_t |  | 0.7569 | 1.0456 | 1.0972 | 1.2688 | 1.5667 | 0.7569 |  | 1.0399 | 1.1833 | 1.2688
p2 = -a/σ_t |  | -0.7569 | -1.0456 | -1.0972 | -1.2688 | -1.5667 | -0.7569 |  | -1.0399 | -1.1833 | -1.2688
q_neg = -t/σ_p |  | -1.5982 | -1.5982 | -1.5982 | -1.5982 | -1.5982 | -1.5982 |  | -1.5982 | -1.5982 | -1.5982
q_pos = t/σ_p |  | 1.5982 | 1.5982 | 1.5982 | 1.5982 | 1.5982 | 1.5982 |  | 1.5982 | 1.5982 | 1.5982
BVN1: b |  | 2.04 | 4.1006 | 4.7067 | 8.04 | 98.04 | 2.04 |  | 4.04 | 6.04 | 8.04
BVN1: aq+b |  | 6.04 | 10.1612 | 11.3733 | 18.04 | 198.04 | 6.04 |  | 10.04 | 14.04 | 18.04
BVN1: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN1: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN1: exp_arg1 |  | -0.22233 | -0.43983 | -0.48065 | -0.6186 | -0.85664 | -0.22233 |  | -0.43532 | -0.54948 | -0.6186
BVN1: exp_arg2 |  | 0.32814 | 0.094738 | 0.033952 | -0.21196 | -0.8054 | 0.32814 |  | 0.10112 | -0.080765 | -0.21196
BVN1: result
BVN2: b |  | -2.04 | -4.1006 | -4.7067 | -8.04 | -98.04 | -2.04 |  | -4.04 | -6.04 | -8.04
BVN2: aq+b |  | 1.96 | 1.96 | 1.96 | 1.96 | 1.96 | 1.96 |  | 1.96 | 1.96 | 1.96
BVN2: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN2: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN2: exp_arg1 |  | 0.32814 | 0.094738 | 0.033952 | -0.21196 | -0.8054 | 0.32814 |  | 0.10112 | -0.080765 | -0.21196
BVN2: exp_arg2 |  | -0.22233 | -0.43983 | -0.48065 | -0.6186 | -0.85664 | -0.22233 |  | -0.43532 | -0.54948 | -0.6186
BVN2: result
BVN3: b |  | -2.04 | -4.1006 | -4.7067 | -8.04 | -98.04 | -2.04 |  | -4.04 | -6.04 | -8.04
BVN3: aq+b |  | -6.04 | -10.1612 | -11.3733 | -18.04 | -198.04 | -6.04 |  | -10.04 | -14.04 | -18.04
BVN3: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN3: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN3: exp_arg1 |  | 0.32814 | 0.094738 | 0.033952 | -0.21196 | -0.8054 | 0.32814 |  | 0.10112 | -0.080765 | -0.21196
BVN3: exp_arg2 |  | -0.22233 | -0.43983 | -0.48065 | -0.6186 | -0.85664 | -0.22233 |  | -0.43532 | -0.54948 | -0.6186
BVN3: result
PFA (Specific) |  | 5.138% | 6.79% | 7.125% | 8.328% | 10.73% | 5.138% |  | 6.753% | 7.711% | 8.328%
PFR (Specific) |  | 33.98% | 18.63% | 16.3% | 9.485% | 0.9104% | 33.98% |  | 18.89% | 12.71% | 9.485%
M6 Multiplier |  | 0.28165 | 0.15193 | 0.11918 | 0 | 0 | 0.28165 |  | 0.15532 | 0.053121 | 0
GB M6 |  | 0.28165 | 0.10028 | 0.071508 | 0 | 0 | 0.28165 |  | 0.10355 | 0.026561 | 0
Accept width (a_m6) |  | 1.7184 | 1.8997 | 1.9285 | 2 | 2 | 1.7184 |  | 1.8965 | 1.9734 | 2
p1_m6 |  | 1.2751 | 1.4679 | 1.4986 | 1.5782 | 1.598 | 1.2751 |  | 1.4644 | 1.5464 | 1.5782
p2_m6 |  | -1.2751 | -1.4679 | -1.4986 | -1.5782 | -1.598 | -1.2751 |  | -1.4644 | -1.5464 | -1.5782
BVN4: b |  | 3.4367 | 5.7567 | 6.4283 | 10 | 100 | 3.4367 |  | 5.6894 | 7.8938 | 10
BVN4: aq+b |  | 7.4367 | 11.8173 | 13.095 | 20 | 200 | 7.4367 |  | 11.6894 | 15.8938 | 20
BVN4: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN4: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN4: exp_arg1 |  | -0.91495 | -1.0676 | -1.0867 | -1.105 | -0.90671 | -0.91495 |  | -1.0653 | -1.111 | -1.105
BVN4: exp_arg2 |  | 0.012403 | -0.31714 | -0.3839 | -0.59921 | -0.85444 | 0.012403 |  | -0.30985 | -0.49844 | -0.59921
BVN4: result
BVN5: b |  | -3.4367 | -5.7567 | -6.4283 | -10 | -100 | -3.4367 |  | -5.6894 | -7.8938 | -10
BVN5: aq+b |  | 0.56329 | 0.30387 | 0.23836 | 0 | 0 | 0.56329 |  | 0.31064 | 0.10624 | 0
BVN5: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN5: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN5: exp_arg1 |  | 0.012403 | -0.31714 | -0.3839 | -0.59921 | -0.85444 | 0.012403 |  | -0.30985 | -0.49844 | -0.59921
BVN5: exp_arg2 |  | -0.91495 | -1.0676 | -1.0867 | -1.105 | -0.90671 | -0.91495 |  | -1.0653 | -1.111 | -1.105
BVN5: result
BVN6: b |  | -3.4367 | -5.7567 | -6.4283 | -10 | -100 | -3.4367 |  | -5.6894 | -7.8938 | -10
BVN6: aq+b |  | -7.4367 | -11.8173 | -13.095 | -20 | -200 | -7.4367 |  | -11.6894 | -15.8938 | -20
BVN6: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN6: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN6: exp_arg1 |  | 0.012403 | -0.31714 | -0.3839 | -0.59921 | -0.85444 | 0.012403 |  | -0.30985 | -0.49844 | -0.59921
BVN6: exp_arg2 |  | -0.91495 | -1.0676 | -1.0867 | -1.105 | -0.90671 | -0.91495 |  | -1.0653 | -1.111 | -1.105
BVN6: result
PFA (Method 6) |  | 9.179% | 10.32% | 10.51% | 11% | 11% | 9.179% |  | 10.3% | 10.82% | 11%
PFR (Method 6) |  | 10.4% | 4.472% | 3.675% | 1.704% | 0.4758% | 10.4% |  | 4.565% | 2.513% | 1.704%
FR/yr (Specific) |  | 169.9059 | 93.1308 | 81.5079 | 47.4269 | 4.5521 | 169.9059 |  | 94.461 | 63.5628 | 47.4269
FR/yr (M6) |  | 51.9985 | 22.3587 | 18.376 | 8.5224 | 2.3788 | 51.9985 |  | 22.8243 | 12.5668 | 8.5224
Rework $/yr (Spec) |  | 19822.3575 | 10865.2552 | 9509.2584 | 5533.1425 | 531.0811 | 19822.3575 |  | 11020.4502 | 7415.6612 | 5533.1425
Rework $/yr (M6) |  | 6066.4961 | 2608.5204 | 2143.8717 | 994.2843 | 277.5303 | 6066.4961 |  | 2662.8333 | 1466.1248 | 994.2843
SIMPLE ACCEPTANCE (Shared Risk, GB = 0)
PFR (Simple) |  | 5.218% | 3.062% | 2.727% | 1.704% | 0.4758% | 5.218%
FR/yr (Simple) |  | 26.0923 | 15.3118 | 13.6327 | 8.5224 | 2.3788 | 26.0923
Rework $/yr (Simple) |  | 3044.0993 | 1786.3811 | 1590.479 | 994.2843 | 277.5303 | 3044.0993
NUMERICAL INTEGRATION (Trapezoidal Rule)
Integration points | 40 | [11 empty] | BCM Spec | BCM M6 | PCM Spec | PCM M6 | UCM Spec | UCM M6 | UCMAuto Spec | UCMAuto M6 | DW Spec | DW M6 | Interface Spec | Interface M6 | TUR3 Spec | TUR3 M6 | TUR4 Spec | TUR4 M6 | TUR5 Spec | TUR5 M6 | Global C | Global D | Global E | Global F | Global G | Global H | Global J | Global K | Global L
-2 | 0.10256 | [11 empty] | 0.0022221 | 0.025478 | 0.0022221 | 0.033834 | 0.0022221 | 0.036073 | 0.0022221 | 0.044446 | 0.0022221 | 0.044446 | 0.0022221 | 0.025478 | 0.0022221 | 0.033605 | 0.0022221 | 0.040686 | 0.0022221 | 0.044446 | 0.039104 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.039104 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446
-1.8974 | [12 empty] | 0.0040035 | 0.036369 | 0.0050046 | 0.050776 | 0.0053343 | 0.05466 | 0.0074658 | 0.070288 | 0.10092 | 0.10099 | 0.0040035 | 0.036369 | 0.0049725 | 0.050378 | 0.0061204 | 0.06256 | 0.0074658 | 0.070288 | 0.052668 | 0.062821 | 0.064008 | 0.070288 | 0.10099 | 0.052668 | 0.062701 | 0.066574 | 0.070288 | 0.058704 | 0.062821 | 0.064008 | 0.070288 | 0.10099 | 0.058704
-1.7949 | [12 empty] | 0.006907 | 0.050056 | 0.010301 | 0.071195 | 0.011504 | 0.076589 | 0.019952 | 0.096589 | 0.11397 | 0.11397 | 0.006907 | 0.050056 | 0.010186 | 0.070632 | 0.014503 | 0.086902 | 0.019952 | 0.096589 | 0.068635 | 0.083531 | 0.085815 | 0.096589 | 0.11397 | 0.068635 | 0.083298 | 0.0905 | 0.096589 | 0.075131 | 0.083531 | 0.085815 | 0.096589 | 0.11397 | 0.075131
-1.6923 | [12 empty] | 0.011419 | 0.066535 | 0.019429 | 0.093927 | 0.022366 | 0.10022 | 0.043015 | 0.11984 | 0.12776 | 0.12776 | 0.011419 | 0.066535 | 0.01915 | 0.09325 | 0.029759 | 0.1111 | 0.043015 | 0.11984 | 0.086716 | 0.10533 | 0.10827 | 0.11984 | 0.12776 | 0.086716 | 0.10502 | 0.11381 | 0.11984 | 0.093374 | 0.10533 | 0.10827 | 0.11984 | 0.12776 | 0.093374
-1.5897 | [12 empty] | 0.018102 | 0.085567 | 0.033679 | 0.11754 | 0.039372 | 0.12385 | 0.076302 | 0.1394 | 0.14226 | 0.14226 | 0.018102 | 0.085567 | 0.033136 | 0.11683 | 0.053328 | 0.13338 | 0.076302 | 0.1394 | 0.10646 | 0.12705 | 0.13006 | 0.1394 | 0.14226 | 0.10646 | 0.12672 | 0.13509 | 0.1394 | 0.11296 | 0.12705 | 0.13006 | 0.1394 | 0.14226 | 0.11296
-1.4872 | [12 empty] | 0.027543 | 0.10669 | 0.053868 | 0.14072 | 0.063102 | 0.14622 | 0.1144 | 0.15652 | 0.15734 | 0.15734 | 0.027543 | 0.10669 | 0.052973 | 0.14007 | 0.084391 | 0.15327 | 0.1144 | 0.15652 | 0.1273 | 0.14788 | 0.15046 | 0.15652 | 0.15734 | 0.1273 | 0.14759 | 0.15417 | 0.15652 | 0.13334 | 0.14788 | 0.15046 | 0.15652 | 0.15734 | 0.13334
-1.3846 | [12 empty] | 0.040262 | 0.12925 | 0.079871 | 0.16261 | 0.092712 | 0.16682 | 0.15003 | 0.17267 | 0.17285 | 0.17285 | 0.040262 | 0.12925 | 0.078592 | 0.16208 | 0.11961 | 0.17125 | 0.15003 | 0.17267 | 0.14862 | 0.16748 | 0.16937 | 0.17267 | 0.17285 | 0.14862 | 0.16725 | 0.17166 | 0.17267 | 0.15398 | 0.16748 | 0.16937 | 0.17267 | 0.17285 | 0.15398
-1.2821 | [12 empty] | 0.056607 | 0.15251 | 0.11041 | 0.18285 | 0.12592 | 0.18568 | 0.1789 | 0.18859 | 0.18862 | 0.18862 | 0.056607 | 0.15251 | 0.10881 | 0.18247 | 0.1545 | 0.18809 | 0.1789 | 0.18859 | 0.16987 | 0.18583 | 0.18705 | 0.18859 | 0.18862 | 0.16987 | 0.18568 | 0.18824 | 0.18859 | 0.17438 | 0.18583 | 0.18705 | 0.18859 | 0.18862 | 0.17438
-1.1795 | [12 empty] | 0.076645 | 0.17572 | 0.14326 | 0.20149 | 0.15965 | 0.20318 | 0.20117 | 0.20445 | 0.20446 | 0.20446 | 0.076645 | 0.17572 | 0.14149 | 0.20124 | 0.18543 | 0.20431 | 0.20117 | 0.20445 | 0.19053 | 0.20314 | 0.20382 | 0.20445 | 0.20446 | 0.19053 | 0.20304 | 0.20435 | 0.20445 | 0.19416 | 0.20314 | 0.20382 | 0.20445 | 0.20446 | 0.19416
-1.0769 | [12 empty] | 0.10009 | 0.19818 | 0.17584 | 0.21875 | 0.19108 | 0.21964 | 0.21927 | 0.22014 | 0.22014 | 0.22014 | 0.10009 | 0.19818 | 0.17409 | 0.2186 | 0.21098 | 0.2201 | 0.21927 | 0.22014 | 0.21022 | 0.21957 | 0.21991 | 0.22014 | 0.22014 | 0.21022 | 0.21952 | 0.22011 | 0.22014 | 0.213 | 0.21957 | 0.21991 | 0.22014 | 0.22014 | 0.213
-0.97436 | [12 empty] | 0.12627 | 0.21934 | 0.20589 | 0.23484 | 0.21841 | 0.23526 | 0.23525 | 0.23543 | 0.23543 | 0.23543 | 0.12627 | 0.21934 | 0.20435 | 0.23477 | 0.23165 | 0.23543 | 0.23525 | 0.23543 | 0.22867 | 0.23521 | 0.23536 | 0.23543 | 0.23543 | 0.22867 | 0.23519 | 0.23543 | 0.23543 | 0.2307 | 0.23521 | 0.23536 | 0.23543 | 0.23543 | 0.2307
-0.87179 | [12 empty] | 0.15418 | 0.2388 | 0.23202 | 0.24988 | 0.24113 | 0.25005 | 0.25008 | 0.25011 | 0.25011 | 0.25011 | 0.15418 | 0.2388 | 0.23081 | 0.24984 | 0.24877 | 0.2501 | 0.25008 | 0.25011 | 0.24569 | 0.25003 | 0.25008 | 0.25011 | 0.25011 | 0.24569 | 0.25002 | 0.2501 | 0.25011 | 0.2471 | 0.25003 | 0.25008 | 0.25011 | 0.25011 | 0.2471
-0.76923 | [12 empty] | 0.18258 | 0.2563 | 0.25378 | 0.26383 | 0.25967 | 0.2639 | 0.26391 | 0.26391 | 0.26391 | 0.26391 | 0.18258 | 0.2563 | 0.25293 | 0.26382 | 0.26351 | 0.26391 | 0.26391 | 0.26391 | 0.26116 | 0.26389 | 0.26391 | 0.26391 | 0.26391 | 0.26116 | 0.26388 | 0.26391 | 0.26391 | 0.26209 | 0.26389 | 0.26391 | 0.26391 | 0.26391 | 0.26209
-0.66667 | [12 empty] | 0.21016 | 0.27172 | 0.27143 | 0.27659 | 0.27482 | 0.27662 | 0.27662 | 0.27662 | 0.27662 | 0.27662 | 0.21016 | 0.27172 | 0.2709 | 0.27659 | 0.27652 | 0.27662 | 0.27662 | 0.27662 | 0.27498 | 0.27661 | 0.27662 | 0.27662 | 0.27662 | 0.27498 | 0.27661 | 0.27662 | 0.27662 | 0.27556 | 0.27661 | 0.27662 | 0.27662 | 0.27662 | 0.27556
-0.5641 | [12 empty] | 0.23567 | 0.28498 | 0.28558 | 0.28799 | 0.28732 | 0.28799 | 0.288 | 0.288 | 0.288 | 0.288 | 0.23567 | 0.28498 | 0.28528 | 0.28799 | 0.28797 | 0.288 | 0.288 | 0.288 | 0.28706 | 0.28799 | 0.288 | 0.288 | 0.288 | 0.28706 | 0.28799 | 0.288 | 0.288 | 0.28741 | 0.28799 | 0.288 | 0.288 | 0.288 | 0.28741
-0.46154 | [12 empty] | 0.25806 | 0.29605 | 0.29681 | 0.29783 | 0.2976 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.25806 | 0.29605 | 0.29665 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29732 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29732 | 0.29783 | 0.29783 | 0.29783 | 0.29752 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29752
-0.35897 | [12 empty] | 0.27658 | 0.30494 | 0.30555 | 0.30594 | 0.30587 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.27658 | 0.30494 | 0.30548 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30568 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30568 | 0.30594 | 0.30594 | 0.30594 | 0.30579 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30579
-0.25641 | [12 empty] | 0.29072 | 0.31162 | 0.31203 | 0.31217 | 0.31215 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.29072 | 0.31162 | 0.312 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31204 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31204 | 0.31217 | 0.31217 | 0.31217 | 0.31209 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31209
-0.15385 | [12 empty] | 0.30024 | 0.31609 | 0.31635 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.30024 | 0.31609 | 0.31634 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31633 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31633 | 0.31639 | 0.31639 | 0.31639 | 0.31636 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31636
-0.051282 | [12 empty] | 0.30501 | 0.31833 | 0.31851 | 0.31853 | 0.31852 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.30501 | 0.31833 | 0.31851 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.31848 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.31848 | 0.31853 | 0.31853 | 0.31853 | 0.3185 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.3185
0.051282 | [12 empty] | 0.30501 | 0.31833 | 0.31851 | 0.31853 | 0.31852 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.30501 | 0.31833 | 0.31851 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.31848 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.31848 | 0.31853 | 0.31853 | 0.31853 | 0.3185 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.3185
0.15385 | [12 empty] | 0.30024 | 0.31609 | 0.31635 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.30024 | 0.31609 | 0.31634 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31633 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31633 | 0.31639 | 0.31639 | 0.31639 | 0.31636 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31636
0.25641 | [12 empty] | 0.29072 | 0.31162 | 0.31203 | 0.31217 | 0.31215 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.29072 | 0.31162 | 0.312 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31204 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31204 | 0.31217 | 0.31217 | 0.31217 | 0.31209 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31209
0.35897 | [12 empty] | 0.27658 | 0.30494 | 0.30555 | 0.30594 | 0.30587 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.27658 | 0.30494 | 0.30548 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30568 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30568 | 0.30594 | 0.30594 | 0.30594 | 0.30579 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30579
0.46154 | [12 empty] | 0.25806 | 0.29605 | 0.29681 | 0.29783 | 0.2976 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.25806 | 0.29605 | 0.29665 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29732 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29732 | 0.29783 | 0.29783 | 0.29783 | 0.29752 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29752
0.5641 | [12 empty] | 0.23567 | 0.28498 | 0.28558 | 0.28799 | 0.28732 | 0.28799 | 0.288 | 0.288 | 0.288 | 0.288 | 0.23567 | 0.28498 | 0.28528 | 0.28799 | 0.28797 | 0.288 | 0.288 | 0.288 | 0.28706 | 0.28799 | 0.288 | 0.288 | 0.288 | 0.28706 | 0.28799 | 0.288 | 0.288 | 0.28741 | 0.28799 | 0.288 | 0.288 | 0.288 | 0.28741
0.66667 | [12 empty] | 0.21016 | 0.27172 | 0.27143 | 0.27659 | 0.27482 | 0.27662 | 0.27662 | 0.27662 | 0.27662 | 0.27662 | 0.21016 | 0.27172 | 0.2709 | 0.27659 | 0.27652 | 0.27662 | 0.27662 | 0.27662 | 0.27498 | 0.27661 | 0.27662 | 0.27662 | 0.27662 | 0.27498 | 0.27661 | 0.27662 | 0.27662 | 0.27556 | 0.27661 | 0.27662 | 0.27662 | 0.27662 | 0.27556
0.76923 | [12 empty] | 0.18258 | 0.2563 | 0.25378 | 0.26383 | 0.25967 | 0.2639 | 0.26391 | 0.26391 | 0.26391 | 0.26391 | 0.18258 | 0.2563 | 0.25293 | 0.26382 | 0.26351 | 0.26391 | 0.26391 | 0.26391 | 0.26116 | 0.26389 | 0.26391 | 0.26391 | 0.26391 | 0.26116 | 0.26388 | 0.26391 | 0.26391 | 0.26209 | 0.26389 | 0.26391 | 0.26391 | 0.26391 | 0.26209
0.87179 | [12 empty] | 0.15418 | 0.2388 | 0.23202 | 0.24988 | 0.24113 | 0.25005 | 0.25008 | 0.25011 | 0.25011 | 0.25011 | 0.15418 | 0.2388 | 0.23081 | 0.24984 | 0.24877 | 0.2501 | 0.25008 | 0.25011 | 0.24569 | 0.25003 | 0.25008 | 0.25011 | 0.25011 | 0.24569 | 0.25002 | 0.2501 | 0.25011 | 0.2471 | 0.25003 | 0.25008 | 0.25011 | 0.25011 | 0.2471
0.97436 | [12 empty] | 0.12627 | 0.21934 | 0.20589 | 0.23484 | 0.21841 | 0.23526 | 0.23525 | 0.23543 | 0.23543 | 0.23543 | 0.12627 | 0.21934 | 0.20435 | 0.23477 | 0.23165 | 0.23543 | 0.23525 | 0.23543 | 0.22867 | 0.23521 | 0.23536 | 0.23543 | 0.23543 | 0.22867 | 0.23519 | 0.23543 | 0.23543 | 0.2307 | 0.23521 | 0.23536 | 0.23543 | 0.23543 | 0.2307
1.0769 | [12 empty] | 0.10009 | 0.19818 | 0.17584 | 0.21875 | 0.19108 | 0.21964 | 0.21927 | 0.22014 | 0.22014 | 0.22014 | 0.10009 | 0.19818 | 0.17409 | 0.2186 | 0.21098 | 0.2201 | 0.21927 | 0.22014 | 0.21022 | 0.21957 | 0.21991 | 0.22014 | 0.22014 | 0.21022 | 0.21952 | 0.22011 | 0.22014 | 0.213 | 0.21957 | 0.21991 | 0.22014 | 0.22014 | 0.213
1.1795 | [12 empty] | 0.076645 | 0.17572 | 0.14326 | 0.20149 | 0.15965 | 0.20318 | 0.20117 | 0.20445 | 0.20446 | 0.20446 | 0.076645 | 0.17572 | 0.14149 | 0.20124 | 0.18543 | 0.20431 | 0.20117 | 0.20445 | 0.19053 | 0.20314 | 0.20382 | 0.20445 | 0.20446 | 0.19053 | 0.20304 | 0.20435 | 0.20445 | 0.19416 | 0.20314 | 0.20382 | 0.20445 | 0.20446 | 0.19416
1.2821 | [12 empty] | 0.056607 | 0.15251 | 0.11041 | 0.18285 | 0.12592 | 0.18568 | 0.1789 | 0.18859 | 0.18862 | 0.18862 | 0.056607 | 0.15251 | 0.10881 | 0.18247 | 0.1545 | 0.18809 | 0.1789 | 0.18859 | 0.16987 | 0.18583 | 0.18705 | 0.18859 | 0.18862 | 0.16987 | 0.18568 | 0.18824 | 0.18859 | 0.17438 | 0.18583 | 0.18705 | 0.18859 | 0.18862 | 0.17438
1.3846 | [12 empty] | 0.040262 | 0.12925 | 0.079871 | 0.16261 | 0.092712 | 0.16682 | 0.15003 | 0.17267 | 0.17285 | 0.17285 | 0.040262 | 0.12925 | 0.078592 | 0.16208 | 0.11961 | 0.17125 | 0.15003 | 0.17267 | 0.14862 | 0.16748 | 0.16937 | 0.17267 | 0.17285 | 0.14862 | 0.16725 | 0.17166 | 0.17267 | 0.15398 | 0.16748 | 0.16937 | 0.17267 | 0.17285 | 0.15398
1.4872 | [12 empty] | 0.027543 | 0.10669 | 0.053868 | 0.14072 | 0.063102 | 0.14622 | 0.1144 | 0.15652 | 0.15734 | 0.15734 | 0.027543 | 0.10669 | 0.052973 | 0.14007 | 0.084391 | 0.15327 | 0.1144 | 0.15652 | 0.1273 | 0.14788 | 0.15046 | 0.15652 | 0.15734 | 0.1273 | 0.14759 | 0.15417 | 0.15652 | 0.13334 | 0.14788 | 0.15046 | 0.15652 | 0.15734 | 0.13334
1.5897 | [12 empty] | 0.018102 | 0.085567 | 0.033679 | 0.11754 | 0.039372 | 0.12385 | 0.076302 | 0.1394 | 0.14226 | 0.14226 | 0.018102 | 0.085567 | 0.033136 | 0.11683 | 0.053328 | 0.13338 | 0.076302 | 0.1394 | 0.10646 | 0.12705 | 0.13006 | 0.1394 | 0.14226 | 0.10646 | 0.12672 | 0.13509 | 0.1394 | 0.11296 | 0.12705 | 0.13006 | 0.1394 | 0.14226 | 0.11296
1.6923 | [12 empty] | 0.011419 | 0.066535 | 0.019429 | 0.093927 | 0.022366 | 0.10022 | 0.043015 | 0.11984 | 0.12776 | 0.12776 | 0.011419 | 0.066535 | 0.01915 | 0.09325 | 0.029759 | 0.1111 | 0.043015 | 0.11984 | 0.086716 | 0.10533 | 0.10827 | 0.11984 | 0.12776 | 0.086716 | 0.10502 | 0.11381 | 0.11984 | 0.093374 | 0.10533 | 0.10827 | 0.11984 | 0.12776 | 0.093374
1.7949 | [12 empty] | 0.006907 | 0.050056 | 0.010301 | 0.071195 | 0.011504 | 0.076589 | 0.019952 | 0.096589 | 0.11397 | 0.11397 | 0.006907 | 0.050056 | 0.010186 | 0.070632 | 0.014503 | 0.086902 | 0.019952 | 0.096589 | 0.068635 | 0.083531 | 0.085815 | 0.096589 | 0.11397 | 0.068635 | 0.083298 | 0.0905 | 0.096589 | 0.075131 | 0.083531 | 0.085815 | 0.096589 | 0.11397 | 0.075131
1.8974 | [12 empty] | 0.0040035 | 0.036369 | 0.0050046 | 0.050776 | 0.0053343 | 0.05466 | 0.0074658 | 0.070288 | 0.10092 | 0.10099 | 0.0040035 | 0.036369 | 0.0049725 | 0.050378 | 0.0061204 | 0.06256 | 0.0074658 | 0.070288 | 0.052668 | 0.062821 | 0.064008 | 0.070288 | 0.10099 | 0.052668 | 0.062701 | 0.066574 | 0.070288 | 0.058704 | 0.062821 | 0.064008 | 0.070288 | 0.10099 | 0.058704
2 | [12 empty] | 0.0022221 | 0.025478 | 0.0022221 | 0.033834 | 0.0022221 | 0.036073 | 0.0022221 | 0.044446 | 0.0022221 | 0.044446 | 0.0022221 | 0.025478 | 0.0022221 | 0.033605 | 0.0022221 | 0.040686 | 0.0022221 | 0.044446 | 0.039104 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.039104 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446
[13 empty] | 0.55019 | 0.786 | 0.70374 | 0.84528 | 0.72698 | 0.85325 | 0.79515 | 0.87296 | 0.8809 | 0.88524 | 0.55019 | 0.786 | 0.70108 | 0.84435 | 0.76287 | 0.86487 | 0.79515 | 0.87296 | 0.82625 | 0.85938 | 0.86273 | 0.87296 | 0.88524 | 0.82625 | 0.85899 | 0.86803 | 0.87296 | 0.83782 | 0.85938 | 0.86273 | 0.87296 | 0.88524 | 0.83782
2%
0.16686
| 2.0834 | [11 empty] | 0.0013328 | 0.01855 | 0.0010728 | 0.023032 | 0.0010052 | 0.02414 | 0.00069556 | 0.026972 | 3.4732e-11 | 1.207e-06 | 0.0013328 | 0.01855 | 0.0010798 | 0.022917 | 0.00086936 | 0.02631 | 0.00069556 | 0.026972
| 2.2503 | [11 empty] | 0.000439 | 0.0090952 | 0.00020758 | 0.0091175 | 0.00016462 | 0.0089699 | 4.1801e-05 | 0.0066707 | 5.587e-49 | 1.9775e-37 | 0.000439 | 0.0090952 | 0.00021239 | 0.0091292 | 9.7002e-05 | 0.0084856 | 4.1801e-05 | 0.0066707
| 2.4171 | [11 empty] | 0.00012838 | 0.0040042 | 3.1193e-05 | 0.0028851 | 1.9911e-05 | 0.0025507 | 1.2876e-06 | 0.00091333 | 7.6932e-117 | 3.2575e-98 | 0.00012838 | 0.0040042 | 3.2598e-05 | 0.0029191 | 7.0342e-06 | 0.001874 | 1.2876e-06 | 0.00091333
| 2.584 | [11 empty] | 3.3279e-05 | 0.001577 | 3.6251e-06 | 0.0007208 | 1.7695e-06 | 0.00054622 | 2.0062e-08 | 6.6191e-05 | 7.0996e-215 | 3.7053e-189 | 3.3279e-05 | 0.001577 | 3.8883e-06 | 0.00074022 | 3.2886e-07 | 0.00027601 | 2.0062e-08 | 6.6191e-05
| 2.7508 | [11 empty] | 7.6386e-06 | 0.0005539 | 3.2481e-07 | 0.00014094 | 1.151e-07 | 8.7115e-05 | 1.5681e-10 | 2.4742e-06 | 0 | 0 | 7.6386e-06 | 0.0005539 | 3.5933e-07 | 0.00014759 | 9.8577e-09 | 2.6656e-05 | 1.5681e-10 | 2.4742e-06
| 2.9177 | [11 empty] | 1.5508e-06 | 0.00017311 | 2.2384e-08 | 2.1432e-05 | 5.4645e-09 | 1.0268e-05 | 6.1141e-13 | 4.6968e-08 | 0 | 0 | 1.5508e-06 | 0.00017311 | 2.5668e-08 | 2.2998e-05 | 1.8871e-10 | 1.6696e-06 | 6.1141e-13 | 4.6968e-08
| 3.0846 | [11 empty] | 2.7828e-07 | 4.8046e-05 | 1.1843e-09 | 2.523e-06 | 1.8895e-10 | 8.8962e-07 | 1.1845e-15 | 4.4839e-10 | 0 | 0 | 2.7828e-07 | 4.8046e-05 | 1.4148e-09 | 2.7878e-06 | 2.3006e-12 | 6.7329e-08 | 1.1845e-15 | 4.4839e-10
| 3.2514 | [11 empty] | 4.4104e-08 | 1.1824e-05 | 4.8043e-11 | 2.2917e-07 | 4.7504e-12 | 5.6432e-08 | 1.1371e-18 | 2.139e-12 | 0 | 0 | 4.4104e-08 | 1.1824e-05 | 6.0085e-11 | 2.6204e-07 | 1.7824e-14 | 1.7391e-09 | 1.1371e-18 | 2.139e-12
| 3.4183 | [11 empty] | 6.17e-09 | 2.5771e-06 | 1.4925e-12 | 1.602e-08 | 8.6729e-14 | 2.6133e-09 | 5.3975e-22 | 5.0763e-15 | 0 | 0 | 6.17e-09 | 2.5771e-06 | 1.9641e-12 | 1.9051e-08 | 8.7609e-17 | 2.8667e-11 | 5.3975e-22 | 5.0763e-15
| 3.5851 | [11 empty] | 7.6156e-10 | 4.9692e-07 | 3.5474e-14 | 8.6021e-10 | 1.1487e-15 | 8.8157e-11 | 1.2649e-25 | 5.9747e-18 | 0 | 0 | 7.6156e-10 | 4.9692e-07 | 4.9371e-14 | 1.0692e-09 | 2.7287e-19 | 3.0078e-13 | 1.2649e-25 | 5.9747e-18
| 3.752 | [11 empty] | 8.29e-11 | 8.4695e-08 | 6.4465e-16 | 3.5424e-11 | 1.1028e-17 | 2.1627e-12 | 1.4617e-29 | 3.4795e-21 | 0 | 0 | 8.29e-11 | 8.4695e-08 | 9.5364e-16 | 4.6254e-11 | 5.3799e-22 | 2.0047e-15 | 1.4617e-29 | 3.4795e-21
| 3.9188 | [11 empty] | 7.9558e-12 | 1.2751e-08 | 8.9514e-18 | 1.1175e-12 | 7.669e-20 | 3.8531e-14 | 8.3204e-34 | 1.0009e-24 | 0 | 0 | 7.9558e-12 | 1.2751e-08 | 1.4146e-17 | 1.5406e-12 | 6.7092e-25 | 8.4743e-18 | 8.3204e-34 | 1.0009e-24
| 4.0857 | [11 empty] | 6.7292e-13 | 1.6947e-09 | 9.4926e-20 | 2.6979e-14 | 3.8609e-22 | 4.9805e-16 | 2.3314e-38 | 1.4202e-28 | 0 | 0 | 6.7292e-13 | 1.6947e-09 | 1.6108e-19 | 3.9467e-14 | 5.2887e-28 | 2.2694e-20 | 2.3314e-38 | 1.4202e-28
| 4.2525 | 0.06484% | 0.0439% | 0.03977% | 0.02465% | 1.159e-09% | 0.06484% |  | 0.04435% | 0.03249% | 0.02465% |  | 5.0151e-14 | 1.9875e-10 | 7.6846e-22 | 4.9807e-16 | 1.4066e-24 | 4.6668e-18 | 3.2137e-43 | 9.9309e-33 | 0 | 0 | 5.0151e-14 | 1.9875e-10 | 1.4072e-21 | 7.7712e-16 | 2.6338e-31 | 3.8464e-23 | 3.2137e-43 | 9.9309e-33
| 4.4194 | 1.135% | 1.199% | 1.212% | 1.155% | 4.028e-05% | 1.135% |  | 1.197% | 1.234% | 1.155% |  | 3.2925e-15 | 2.0558e-11 | 4.7474e-24 | 7.0271e-18 | 3.7066e-27 | 3.1676e-20 | 2.1781e-48 | 3.419e-37 | 0 | 0 | 3.2925e-15 | 2.0558e-11 | 9.4296e-24 | 1.1753e-17 | 8.2828e-35 | 4.1229e-26 | 2.1781e-48 | 3.419e-37
| 4.5863 | [11 empty] | 1.9038e-16 | 1.8748e-12 | 2.2374e-26 | 7.5727e-20 | 7.0629e-30 | 1.5567e-22 | 7.2547e-54 | 5.7916e-42 | 0 | 0 | 1.9038e-16 | 1.8748e-12 | 4.8449e-26 | 1.3646e-19 | 1.6442e-38 | 2.7931e-29 | 7.2547e-54 | 5.7916e-42
| 4.7531 | 2.389% | 1.77% | 1.64% | 1.155% | 4.028e-05% | 2.389% |  | 1.783% | 1.409% | 1.155% |  | 9.6935e-18 | 1.5069e-13 | 8.0421e-29 | 6.2304e-22 | 9.7291e-33 | 5.5358e-25 | 1.1872e-59 | 4.8245e-47 | 0 | 0 | 9.6935e-18 | 1.5069e-13 | 1.9082e-28 | 1.2158e-21 | 2.0597e-42 | 1.1953e-32 | 1.1872e-59 | 4.8245e-47
| 4.92 | [11 empty] | 4.3455e-19 | 1.0673e-14 | 2.2041e-31 | 3.9121e-24 | 9.6856e-36 | 1.4241e-27 | 9.5409e-66 | 1.9754e-52 | 0 | 0 | 4.3455e-19 | 1.0673e-14 | 5.7596e-31 | 8.3092e-24 | 1.6277e-46 | 3.2297e-36 | 9.5409e-66 | 1.9754e-52
| 5.0868 | [11 empty] | 1.7149e-20 | 6.6594e-16 | 4.6054e-34 | 1.8741e-26 | 6.9673e-39 | 2.649e-30 | 3.7649e-72 | 3.9743e-58 | 0 | 0 | 1.7149e-20 | 6.6594e-16 | 1.3321e-33 | 4.3545e-26 | 8.113e-51 | 5.5085e-40 | 3.7649e-72 | 3.9743e-58
| 525.4% | [11 empty] | 5.9572e-22 | 3.6597e-17 | 7.3345e-37 | 6.8478e-29 | 3.6207e-42 | 3.5623e-33 | 7.2931e-79 | 3.9273e-64 | 0 | 0 | 5.9572e-22 | 3.6597e-17 | 2.3603e-36 | 1.7495e-28 | 2.5499e-55 | 5.9282e-44 | 7.2931e-79 | 3.9273e-64
| 5.4205 | [11 empty] | 1.8213e-23 | 1.7711e-18 | 8.9021e-40 | 1.908e-31 | 1.3591e-45 | 3.4622e-36 | 6.9336e-86 | 1.9057e-70 | 0 | 0 | 1.8213e-23 | 1.7711e-18 | 3.2034e-39 | 5.3869e-31 | 5.0529e-60 | 4.0246e-48 | 6.9336e-86 | 1.9057e-70
| 5.5874 | [11 empty] | 4.9001e-25 | 7.5464e-20 | 8.2332e-43 | 4.0531e-34 | 3.6844e-49 | 2.4314e-39 | 3.2347e-93 | 4.5398e-77 | 0 | 0 | 4.9001e-25 | 7.5464e-20 | 3.3298e-42 | 1.2711e-33 | 6.3117e-65 | 1.7232e-52 | 3.2347e-93 | 4.5398e-77
| 5.7542 | [11 empty] | 1.1601e-26 | 2.8307e-21 | 5.8016e-46 | 6.5628e-37 | 7.2127e-53 | 1.2336e-42 | 7.404e-101 | 5.3081e-84 | 0 | 0 | 1.1601e-26 | 2.8307e-21 | 2.6505e-45 | 2.2977e-36 | 4.9692e-70 | 4.6522e-57 | 7.404e-101 | 5.3081e-84
| 5.9211 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 2.4166e-28 | 9.3467e-23 | 3.1144e-49 | 8.0989e-40 | 1.0195e-56 | 4.5208e-46 | 8.3136e-109 | 3.0457e-91 | 0 | 0 | 2.4166e-28 | 9.3467e-23 | 1.6155e-48 | 3.1817e-39 | 2.4655e-75 | 7.9185e-62 | 8.3136e-109 | 3.0457e-91
| 6.0879 | 0.98 | 0.6468 | 0.588 | 0.392 | 0.0392 | 0.98 |  | 0.65333 | 0.49 | 0.392 |  | 4.4289e-30 | 2.7162e-24 | 1.2736e-52 | 7.6161e-43 | 1.0404e-60 | 1.1965e-49 | 4.5789e-117 | 8.5747e-99 | 0 | 0 | 4.4289e-30 | 2.7162e-24 | 7.539e-52 | 3.3745e-42 | 7.7084e-81 | 8.4958e-67 | 4.5789e-117 | 8.5747e-99
| 6.2548 | 0.49 | 0.3234 | 0.294 | 0.196 | 0.0196 | 0.49 |  | 0.32667 | 0.245 | 0.196 |  | 7.1409e-32 | 6.9465e-26 | 3.9668e-56 | 5.457e-46 | 7.664e-65 | 2.287e-53 | 1.2369e-125 | 1.1843e-106 | 0 | 0 | 7.1409e-32 | 6.9465e-26 | 2.6934e-55 | 2.7408e-45 | 1.5185e-86 | 5.745e-72 | 1.2369e-125 | 1.1843e-106
| 6.4217 | 0.0056682 | 0.0040111 | 0.0036778 | 0.0024496 | 3.44e-09 | 0.0056682 |  | 0.0040472 | 0.0030888 | 0.0024496 |  | 1.0129e-33 | 1.5633e-27 | 9.4103e-60 | 2.9788e-49 | 4.0754e-69 | 3.1561e-57 | 1.6386e-134 | 8.0237e-115 | 0 | 0 | 1.0129e-33 | 1.5633e-27 | 7.366e-59 | 1.7046e-48 | 1.8845e-92 | 2.4482e-77 | 1.6386e-134 | 8.0237e-115
| 6.5885 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 1.2637e-35 | 3.0956e-29 | 1.7001e-63 | 1.2387e-52 | 1.5642e-73 | 3.1445e-61 | 1.0644e-143 | 2.6662e-123 | 0 | 0 | 1.2637e-35 | 3.0956e-29 | 1.542e-62 | 8.1172e-52 | 1.4734e-98 | 6.5739e-83 | 1.0644e-143 | 2.6662e-123
| 6.7554 | 0.49 | 0.3234 | 0.294 | 0.196 | 0.0196 | 0.49 |  | 0.32667 | 0.245 | 0.196 |  | 1.3869e-37 | 5.393e-31 | 2.339e-67 | 3.9233e-56 | 4.333e-78 | 2.2617e-65 | 3.3904e-153 | 4.3451e-132 | 0 | 0 | 1.3869e-37 | 5.393e-31 | 2.4707e-66 | 2.9592e-55 | 7.2563e-105 | 1.1122e-88 | 3.3904e-153 | 4.3451e-132
| 6.9222 | 0.245 | 0.1617 | 0.147 | 0.098 | 0.0098 | 0.245 |  | 0.16333 | 0.1225 | 0.098 |  | 1.3387e-39 | 8.266e-33 | 2.4504e-71 | 9.4646e-60 | 8.6623e-83 | 1.1743e-69 | 5.2947e-163 | 3.4724e-141 | 0 | 0 | 1.3387e-39 | 8.266e-33 | 3.0299e-70 | 8.2586e-59 | 2.251e-111 | 1.1855e-94 | 5.2947e-163 | 3.4724e-141
|  | 0.01266 | 0.009161 | 0.0084467 | 0.0057954 | 4.1787e-08 | 0.01266 |  | 0.0092384 | 0.0071779 | 0.0057954
|  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0
|  | 0.245 | 0.1617 | 0.147 | 0.098 | 0.0098 | 0.245 |  | 0.16333 | 0.1225 | 0.098
|  | 0.1225 | 0.08085 | 0.0735 | 0.049 | 0.0049 | 0.1225 |  | 0.081667 | 0.06125 | 0.049
|  | 0.017741 | 0.012987 | 0.012007 | 0.0083505 | 1.3353e-07 | 0.017741 |  | 0.013093 | 0.01026 | 0.0083505
|  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0
|  | 0.1225 | 0.08085 | 0.0735 | 0.049 | 0.0049 | 0.1225 |  | 0.081667 | 0.06125 | 0.049
|  | 0.06125 | 0.040425 | 0.03675 | 0.0245 | 0.00245 | 0.06125 |  | 0.040833 | 0.030625 | 0.0245
|  | 0.020686 | 0.015233 | 0.014102 | 0.0098711 | 2.3358e-07 | 0.020686 |  | 0.015355 | 0.012083 | 0.0098711
|  | 0.06125 | 0 | 0 | 0 | 0 | 0.06125 |  | 0 | 0 | 0
|  | 0.1225 | 0.040425 | 0.03675 | 0.0245 | 0.00245 | 0.1225 |  | 0.040833 | 0.030625 | 0.0245
|  | 0.091875 | 0.020213 | 0.018375 | 0.01225 | 0.001225 | 0.091875 |  | 0.020417 | 0.015312 | 0.01225
|  | 0.019181 | 0.016437 | 0.015228 | 0.010693 | 3.0728e-07 | 0.019181 |  | 0.016568 | 0.013064 | 0.010693
|  | 0.06125 | 0 | 0 | 0 | 0 | 0.06125 |  | 0 | 0 | 0
|  | 0.091875 | 0.020213 | 0.018375 | 0.01225 | 0.001225 | 0.091875 |  | 0.020417 | 0.015312 | 0.01225
|  | 0.076563 | 0.010106 | 0.0091875 | 0.006125 | 0.0006125 | 0.076563 |  | 0.010208 | 0.0076562 | 0.006125
|  | 0.019926 | 0.01706 | 0.015809 | 0.011119 | 3.5197e-07 | 0.019926 |  | 0.017195 | 0.013572 | 0.011119
|  | 0.06125 | 0 | 0 | 0 | 0 | 0.06125 |  | 0 | 0 | 0
|  | 0.076563 | 0.010106 | 0.0091875 | 0.006125 | 0.0006125 | 0.076563 |  | 0.010208 | 0.0076562 | 0.006125
|  | 0.068906 | 0.0050531 | 0.0045937 | 0.0030625 | 0.00030625 | 0.068906 |  | 0.0051042 | 0.0038281 | 0.0030625
|  | 0.020304 | 0.017376 | 0.016105 | 0.011335 | 3.7656e-07 | 0.020304 |  | 0.017513 | 0.013831 | 0.011335
|  | 0.068906 | 0 | 0 | 0 | 0 | 0.068906 |  | 0 | 0 | 0
|  | 0.076563 | 0.0050531 | 0.0045937 | 0.0030625 | 0.00030625 | 0.076563 |  | 0.0051042 | 0.0038281 | 0.0030625
|  | 0.072734 | 0.0025266 | 0.0022969 | 0.0015313 | 0.00015312 | 0.072734 |  | 0.0025521 | 0.0019141 | 0.0015313
|  | 0.020114 | 0.017535 | 0.016254 | 0.011445 | 3.8946e-07 | 0.020114 |  | 0.017673 | 0.013961 | 0.011445
|  | 0.072734 | 0 | 0 | 0 | 0 | 0.072734 |  | 0 | 0 | 0
|  | 0.076563 | 0.0025266 | 0.0022969 | 0.0015313 | 0.00015312 | 0.076563 |  | 0.0025521 | 0.0019141 | 0.0015313
|  | 0.074648 | 0.0012633 | 0.0011484 | 0.00076563 | 7.6562e-05 | 0.074648 |  | 0.001276 | 0.00095703 | 0.00076563
|  | 0.02002 | 0.017615 | 0.016329 | 0.0115 | 3.9607e-07 | 0.02002 |  | 0.017754 | 0.014026 | 0.0115
|  | 0.074648 | 0 | 0 | 0 | 0 | 0.074648 |  | 0 | 0 | 0
|  | 0.076563 | 0.0012633 | 0.0011484 | 0.00076563 | 7.6562e-05 | 0.076563 |  | 0.001276 | 0.00095703 | 0.00076563
|  | 0.075605 | 0.00063164 | 0.00057422 | 0.00038281 | 3.8281e-05 | 0.075605 |  | 0.00063802 | 0.00047852 | 0.00038281
|  | 0.019973 | 0.017655 | 0.016366 | 0.011527 | 3.9941e-07 | 0.019973 |  | 0.017794 | 0.014059 | 0.011527
Solved GB (Global) |  | 0.075605 | 0 | 0 | 0 | 0 | 0.075605 |  | 0 | 0 | 0
Accept width (Global) |  | 1.9244 | 2 | 2 | 2 | 2 | 1.9244 |  | 2 | 2 | 2
PFA (Global) |  | 1.997% | 1.77% | 1.64% | 1.155% | 4.028e-05% | 1.997% |  | 1.783% | 1.409% | 1.155%
PFR (Global) |  | 6.375% | 3.062% | 2.727% | 1.704% | 0.4758% | 6.375% |  | 3.101% | 2.197% | 1.704%
FR/yr (Global) |  | 31.8749 | 15.3118 | 13.6327 | 8.5224 | 2.3788 | 31.8749 |  | 15.5026 | 10.9836 | 8.5224
Rework $/yr (Global) |  | 3718.743 | 1786.3811 | 1590.479 | 994.2843 | 277.5303 | 3718.743 |  | 1808.6326 | 1281.425 | 994.2843
RANGE ANALYSIS INTEGRATION
dx_hi | 0.10256
-2 |  | 0.0022221 | 0.0022221 | 0.0022221 | 0.0022221 | 0.0022221 | 0.0022221 | 0.002222 | 0.025478 | 0.033834 | 0.036073 | 0.044446 | 0.044446 | 0.025478 | 0.022497 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446
-1.8974 |  | 0.0040035 | 0.0050046 | 0.0053343 | 0.0074658 | 0.10092 | 0.0040035 | 0.0037165 | 0.036369 | 0.050776 | 0.05466 | 0.070288 | 0.10099 | 0.036369 | 0.031379 | 0.058704 | 0.062821 | 0.064008 | 0.070288 | 0.10099 | 0.058704 | 0.057351
-1.7949 |  | 0.006907 | 0.010301 | 0.011504 | 0.019952 | 0.11397 | 0.006907 | 0.0060192 | 0.050056 | 0.071195 | 0.076589 | 0.096589 | 0.11397 | 0.050056 | 0.042555 | 0.075131 | 0.083531 | 0.085815 | 0.096589 | 0.11397 | 0.075131 | 0.072234
-1.6923 |  | 0.011419 | 0.019429 | 0.022366 | 0.043015 | 0.12776 | 0.011419 | 0.0094428 | 0.066535 | 0.093927 | 0.10022 | 0.11984 | 0.12776 | 0.066535 | 0.056163 | 0.093374 | 0.10533 | 0.10827 | 0.11984 | 0.12776 | 0.093374 | 0.088917
-1.5897 |  | 0.018102 | 0.033679 | 0.039372 | 0.076302 | 0.14226 | 0.018102 | 0.014355 | 0.085567 | 0.11754 | 0.12385 | 0.1394 | 0.14226 | 0.085567 | 0.072204 | 0.11296 | 0.12705 | 0.13006 | 0.1394 | 0.14226 | 0.11296 | 0.10711
-1.4872 |  | 0.027543 | 0.053868 | 0.063102 | 0.1144 | 0.15734 | 0.027543 | 0.021155 | 0.10669 | 0.14072 | 0.14622 | 0.15652 | 0.15734 | 0.10669 | 0.090517 | 0.13334 | 0.14788 | 0.15046 | 0.15652 | 0.15734 | 0.13334 | 0.12644
-1.3846 |  | 0.040262 | 0.079871 | 0.092712 | 0.15003 | 0.17285 | 0.040262 | 0.030239 | 0.12925 | 0.16261 | 0.16682 | 0.17267 | 0.17285 | 0.12925 | 0.11078 | 0.15398 | 0.16748 | 0.16937 | 0.17267 | 0.17285 | 0.15398 | 0.14649
-1.2821 |  | 0.056607 | 0.11041 | 0.12592 | 0.1789 | 0.18862 | 0.056607 | 0.041947 | 0.15251 | 0.18285 | 0.18568 | 0.18859 | 0.18862 | 0.15251 | 0.13252 | 0.17438 | 0.18583 | 0.18705 | 0.18859 | 0.18862 | 0.17438 | 0.16679
-1.1795 |  | 0.076645 | 0.14326 | 0.15965 | 0.20117 | 0.20446 | 0.076645 | 0.056501 | 0.17572 | 0.20149 | 0.20318 | 0.20445 | 0.20446 | 0.17572 | 0.15516 | 0.19416 | 0.20314 | 0.20382 | 0.20445 | 0.20446 | 0.19416 | 0.18693
-1.0769 |  | 0.10009 | 0.17584 | 0.19108 | 0.21927 | 0.22014 | 0.10009 | 0.073948 | 0.19818 | 0.21875 | 0.21964 | 0.22014 | 0.22014 | 0.19818 | 0.17806 | 0.213 | 0.21957 | 0.21991 | 0.22014 | 0.22014 | 0.213 | 0.2065
-0.97436 |  | 0.12627 | 0.20589 | 0.21841 | 0.23525 | 0.23543 | 0.12627 | 0.094103 | 0.21934 | 0.23484 | 0.23526 | 0.23543 | 0.23543 | 0.21934 | 0.20056 | 0.2307 | 0.23521 | 0.23536 | 0.23543 | 0.23543 | 0.2307 | 0.22515
-0.87179 |  | 0.15418 | 0.23202 | 0.24113 | 0.25008 | 0.25011 | 0.15418 | 0.11652 | 0.2388 | 0.24988 | 0.25005 | 0.25011 | 0.25011 | 0.2388 | 0.22207 | 0.2471 | 0.25003 | 0.25008 | 0.25011 | 0.25011 | 0.2471 | 0.24259
-0.76923 |  | 0.18258 | 0.25378 | 0.25967 | 0.26391 | 0.26391 | 0.18258 | 0.14051 | 0.2563 | 0.26383 | 0.2639 | 0.26391 | 0.26391 | 0.2563 | 0.24205 | 0.26209 | 0.26389 | 0.26391 | 0.26391 | 0.26391 | 0.26209 | 0.2586
-0.66667 |  | 0.21016 | 0.27143 | 0.27482 | 0.27662 | 0.27662 | 0.21016 | 0.16512 | 0.27172 | 0.27659 | 0.27662 | 0.27662 | 0.27662 | 0.27172 | 0.26009 | 0.27556 | 0.27661 | 0.27662 | 0.27662 | 0.27662 | 0.27556 | 0.27298
-0.5641 |  | 0.23567 | 0.28558 | 0.28732 | 0.288 | 0.288 | 0.23567 | 0.18927 | 0.28498 | 0.28799 | 0.28799 | 0.288 | 0.288 | 0.28498 | 0.27587 | 0.28741 | 0.28799 | 0.288 | 0.288 | 0.288 | 0.28741 | 0.28559
-0.46154 |  | 0.25806 | 0.29681 | 0.2976 | 0.29783 | 0.29783 | 0.25806 | 0.21177 | 0.29605 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29605 | 0.28918 | 0.29752 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29752 | 0.29629
-0.35897 |  | 0.27658 | 0.30555 | 0.30587 | 0.30594 | 0.30594 | 0.27658 | 0.23145 | 0.30494 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30494 | 0.2999 | 0.30579 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30579 | 0.30498
-0.25641 |  | 0.29072 | 0.31203 | 0.31215 | 0.31217 | 0.31217 | 0.29072 | 0.24725 | 0.31162 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31162 | 0.30796 | 0.31209 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31209 | 0.31157
-0.15385 |  | 0.30024 | 0.31635 | 0.31639 | 0.31639 | 0.31639 | 0.30024 | 0.25831 | 0.31609 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31609 | 0.31334 | 0.31636 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31636 | 0.31601
-0.051282 |  | 0.30501 | 0.31851 | 0.31852 | 0.31853 | 0.31853 | 0.30501 | 0.26399 | 0.31833 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.31833 | 0.31603 | 0.3185 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.3185 | 0.31824
0.051282 |  | 0.30501 | 0.31851 | 0.31852 | 0.31853 | 0.31853 | 0.30501 | 0.26399 | 0.31833 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.31833 | 0.31603 | 0.3185 | 0.31853 | 0.31853 | 0.31853 | 0.31853 | 0.3185 | 0.31824
0.15385 |  | 0.30024 | 0.31635 | 0.31639 | 0.31639 | 0.31639 | 0.30024 | 0.25831 | 0.31609 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31609 | 0.31334 | 0.31636 | 0.31639 | 0.31639 | 0.31639 | 0.31639 | 0.31636 | 0.31601
0.25641 |  | 0.29072 | 0.31203 | 0.31215 | 0.31217 | 0.31217 | 0.29072 | 0.24725 | 0.31162 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31162 | 0.30796 | 0.31209 | 0.31217 | 0.31217 | 0.31217 | 0.31217 | 0.31209 | 0.31157
0.35897 |  | 0.27658 | 0.30555 | 0.30587 | 0.30594 | 0.30594 | 0.27658 | 0.23145 | 0.30494 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30494 | 0.2999 | 0.30579 | 0.30594 | 0.30594 | 0.30594 | 0.30594 | 0.30579 | 0.30498
0.46154 |  | 0.25806 | 0.29681 | 0.2976 | 0.29783 | 0.29783 | 0.25806 | 0.21177 | 0.29605 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29605 | 0.28918 | 0.29752 | 0.29783 | 0.29783 | 0.29783 | 0.29783 | 0.29752 | 0.29629
0.5641 |  | 0.23567 | 0.28558 | 0.28732 | 0.288 | 0.288 | 0.23567 | 0.18927 | 0.28498 | 0.28799 | 0.28799 | 0.288 | 0.288 | 0.28498 | 0.27587 | 0.28741 | 0.28799 | 0.288 | 0.288 | 0.288 | 0.28741 | 0.28559
0.66667 |  | 0.21016 | 0.27143 | 0.27482 | 0.27662 | 0.27662 | 0.21016 | 0.16512 | 0.27172 | 0.27659 | 0.27662 | 0.27662 | 0.27662 | 0.27172 | 0.26009 | 0.27556 | 0.27661 | 0.27662 | 0.27662 | 0.27662 | 0.27556 | 0.27298
0.76923 |  | 0.18258 | 0.25378 | 0.25967 | 0.26391 | 0.26391 | 0.18258 | 0.14051 | 0.2563 | 0.26383 | 0.2639 | 0.26391 | 0.26391 | 0.2563 | 0.24205 | 0.26209 | 0.26389 | 0.26391 | 0.26391 | 0.26391 | 0.26209 | 0.2586
0.87179 |  | 0.15418 | 0.23202 | 0.24113 | 0.25008 | 0.25011 | 0.15418 | 0.11652 | 0.2388 | 0.24988 | 0.25005 | 0.25011 | 0.25011 | 0.2388 | 0.22207 | 0.2471 | 0.25003 | 0.25008 | 0.25011 | 0.25011 | 0.2471 | 0.24259
0.97436 |  | 0.12627 | 0.20589 | 0.21841 | 0.23525 | 0.23543 | 0.12627 | 0.094103 | 0.21934 | 0.23484 | 0.23526 | 0.23543 | 0.23543 | 0.21934 | 0.20056 | 0.2307 | 0.23521 | 0.23536 | 0.23543 | 0.23543 | 0.2307 | 0.22515
1.0769 |  | 0.10009 | 0.17584 | 0.19108 | 0.21927 | 0.22014 | 0.10009 | 0.073948 | 0.19818 | 0.21875 | 0.21964 | 0.22014 | 0.22014 | 0.19818 | 0.17806 | 0.213 | 0.21957 | 0.21991 | 0.22014 | 0.22014 | 0.213 | 0.2065
1.1795 |  | 0.076645 | 0.14326 | 0.15965 | 0.20117 | 0.20446 | 0.076645 | 0.056501 | 0.17572 | 0.20149 | 0.20318 | 0.20445 | 0.20446 | 0.17572 | 0.15516 | 0.19416 | 0.20314 | 0.20382 | 0.20445 | 0.20446 | 0.19416 | 0.18693
1.2821 |  | 0.056607 | 0.11041 | 0.12592 | 0.1789 | 0.18862 | 0.056607 | 0.041947 | 0.15251 | 0.18285 | 0.18568 | 0.18859 | 0.18862 | 0.15251 | 0.13252 | 0.17438 | 0.18583 | 0.18705 | 0.18859 | 0.18862 | 0.17438 | 0.16679
1.3846 |  | 0.040262 | 0.079871 | 0.092712 | 0.15003 | 0.17285 | 0.040262 | 0.030239 | 0.12925 | 0.16261 | 0.16682 | 0.17267 | 0.17285 | 0.12925 | 0.11078 | 0.15398 | 0.16748 | 0.16937 | 0.17267 | 0.17285 | 0.15398 | 0.14649
1.4872 |  | 0.027543 | 0.053868 | 0.063102 | 0.1144 | 0.15734 | 0.027543 | 0.021155 | 0.10669 | 0.14072 | 0.14622 | 0.15652 | 0.15734 | 0.10669 | 0.090517 | 0.13334 | 0.14788 | 0.15046 | 0.15652 | 0.15734 | 0.13334 | 0.12644
1.5897 |  | 0.018102 | 0.033679 | 0.039372 | 0.076302 | 0.14226 | 0.018102 | 0.014355 | 0.085567 | 0.11754 | 0.12385 | 0.1394 | 0.14226 | 0.085567 | 0.072204 | 0.11296 | 0.12705 | 0.13006 | 0.1394 | 0.14226 | 0.11296 | 0.10711
1.6923 |  | 0.011419 | 0.019429 | 0.022366 | 0.043015 | 0.12776 | 0.011419 | 0.0094428 | 0.066535 | 0.093927 | 0.10022 | 0.11984 | 0.12776 | 0.066535 | 0.056163 | 0.093374 | 0.10533 | 0.10827 | 0.11984 | 0.12776 | 0.093374 | 0.088917
1.7949 |  | 0.006907 | 0.010301 | 0.011504 | 0.019952 | 0.11397 | 0.006907 | 0.0060192 | 0.050056 | 0.071195 | 0.076589 | 0.096589 | 0.11397 | 0.050056 | 0.042555 | 0.075131 | 0.083531 | 0.085815 | 0.096589 | 0.11397 | 0.075131 | 0.072234
1.8974 |  | 0.0040035 | 0.0050046 | 0.0053343 | 0.0074658 | 0.10092 | 0.0040035 | 0.0037165 | 0.036369 | 0.050776 | 0.05466 | 0.070288 | 0.10099 | 0.036369 | 0.031379 | 0.058704 | 0.062821 | 0.064008 | 0.070288 | 0.10099 | 0.058704 | 0.057351
2 |  | 0.0022221 | 0.0022221 | 0.0022221 | 0.0022221 | 0.0022221 | 0.0022221 | 0.002222 | 0.025478 | 0.033834 | 0.036073 | 0.044446 | 0.044446 | 0.025478 | 0.022497 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446 | 0.044446
|  | 0.55019 | 0.70374 | 0.72698 | 0.79515 | 0.8809 | 0.55019 | 0.44651 | 0.786 | 0.84528 | 0.85325 | 0.87296 | 0.88524 | 0.786 | 0.74002 | 0.83782 | 0.85938 | 0.86273 | 0.87296 | 0.88524 | 0.83782 | 0.82318
|  | 33.98% | 18.63% | 16.3% | 9.485% | 0.9104% | 33.98% | 44.35% | 10.4% | 4.472% | 3.675% | 1.704% | 0.4758% | 10.4% | 15% | 5.218% | 3.062% | 2.727% | 1.704% | 0.4758% | 5.218% | 6.682%
dx_lo | 0.025641
-0.5 |  | 0.0088881 | 0.0088885 | 0.0088885 | 0.0088885 | 0.0088885 | -0.021273 | -0.076408 | 0.08999 | 0.14429 | 0.14429 | 0.16274 | 0.17779 | 0.053639 | 0.029678 | 0.17779 | 0.17779 | 0.17779 | 0.17779 | 0.17779 | 0.17763 | 0.17642
-0.47436 |  | 0.014866 | 0.021337 | 0.021337 | 0.024482 | 0.40367 | -0.027659 | -0.093872 | 0.12551 | 0.21864 | 0.21864 | 0.25024 | 0.40398 | 0.06873 | 0.036782 | 0.2294 | 0.25603 | 0.25603 | 0.2663 | 0.40398 | 0.21551 | 0.21111
-0.44872 |  | 0.024077 | 0.046016 | 0.046016 | 0.058013 | 0.4559 | -0.03547 | -0.11408 | 0.17022 | 0.30636 | 0.30636 | 0.34761 | 0.4559 | 0.086917 | 0.045074 | 0.28894 | 0.34326 | 0.34326 | 0.362 | 0.4559 | 0.25853 | 0.25014
-0.42308 |  | 0.037771 | 0.089463 | 0.089463 | 0.11904 | 0.51104 | -0.044862 | -0.13714 | 0.22465 | 0.40088 | 0.40088 | 0.44441 | 0.51104 | 0.10848 | 0.054615 | 0.35567 | 0.43309 | 0.43309 | 0.45524 | 0.51104 | 0.3067 | 0.29352
-0.39744 |  | 0.057419 | 0.15749 | 0.15749 | 0.21331 | 0.56903 | -0.055962 | -0.16308 | 0.28882 | 0.49539 | 0.49539 | 0.53351 | 0.56903 | 0.13363 | 0.065433 | 0.42844 | 0.52025 | 0.52025 | 0.54035 | 0.56903 | 0.35985 | 0.34108
-0.37179 |  | 0.08462 | 0.25241 | 0.25241 | 0.33756 | 0.62935 | -0.06885 | -0.19184 | 0.36207 | 0.58489 | 0.58489 | 0.61306 | 0.62935 | 0.16246 | 0.077513 | 0.50577 | 0.60186 | 0.60186 | 0.61669 | 0.62935 | 0.41764 | 0.39255
-0.34615 |  | 0.12096 | 0.37085 | 0.37085 | 0.47846 | 0.69141 | -0.083543 | -0.22323 | 0.44311 | 0.66726 | 0.66726 | 0.68501 | 0.69141 | 0.19494 | 0.090791 | 0.58595 | 0.6775 | 0.6775 | 0.68663 | 0.69141 | 0.47954 | 0.44746
-0.32051 |  | 0.16779 | 0.50367 | 0.50367 | 0.61798 | 0.7545 | -0.099981 | -0.25695 | 0.53008 | 0.74274 | 0.74274 | 0.75236 | 0.7545 | 0.23088 | 0.10515 | 0.66718 | 0.7482 | 0.7482 | 0.75296 | 0.7545 | 0.54479 | 0.50519
-0.29487 |  | 0.226 | 0.6386 | 0.6386 | 0.74172 | 0.81784 | -0.11801 | -0.29258 | 0.62064 | 0.81271 | 0.81271 | 0.81722 | 0.81784 | 0.26988 | 0.12041 | 0.74772 | 0.81529 | 0.81529 | 0.81741 | 0.81784 | 0.61248 | 0.56496
-0.26923 |  | 0.29579 | 0.76431 | 0.76431 | 0.84391 | 0.88055 | -0.13738 | -0.32956 | 0.71223 | 0.87856 | 0.87856 | 0.88041 | 0.88055 | 0.31139 | 0.13634 | 0.82599 | 0.87963 | 0.87963 | 0.88046 | 0.88055 | 0.6815 | 0.62584
-0.24359 |  | 0.37641 | 0.87364 | 0.87364 | 0.9266 | 0.94173 | -0.15773 | -0.36721 | 0.80225 | 0.94104 | 0.94104 | 0.9417 | 0.94173 | 0.35463 | 0.15264 | 0.90059 | 0.94144 | 0.94144 | 0.94172 | 0.94173 | 0.75061 | 0.68676
-0.21795 |  | 0.4661 | 0.96453 | 0.96453 | 0.99508 | 1.0004 | -0.17862 | -0.40475 | 0.88827 | 1.0002 | 1.0002 | 1.0004 | 1.0004 | 0.39863 | 0.16897 | 0.97038 | 1.0003 | 1.0003 | 1.0004 | 1.0004 | 0.81843 | 0.74656
-0.19231 |  | 0.56204 | 1.0387 | 1.0387 | 1.054 | 1.0557 | -0.19949 | -0.44134 | 0.9682 | 1.0556 | 1.0556 | 1.0557 | 1.0557 | 0.4423 | 0.18495 | 1.0344 | 1.0556 | 1.0556 | 1.0557 | 1.0557 | 0.88354 | 0.80399
-0.16667 |  | 0.6605 | 1.0993 | 1.0993 | 1.1061 | 1.1065 | -0.21975 | -0.47605 | 1.0403 | 1.1065 | 1.1065 | 1.1065 | 1.1065 | 0.48441 | 0.20017 | 1.0919 | 1.1065 | 1.1065 | 1.1065 | 1.1065 | 0.9445 | 0.85779
-0.14103 |  | 0.75709 | 1.1493 | 1.1493 | 1.1519 | 1.152 | -0.23874 | -0.50796 | 1.1035 | 1.152 | 1.152 | 1.152 | 1.152 | 0.52367 | 0.2142 | 1.1424 | 1.152 | 1.152 | 1.152 | 1.152 | 0.99987 | 0.90672
-0.11538 |  | 0.84708 | 1.1904 | 1.1904 | 1.1913 | 1.1913 | -0.25582 | -0.53618 | 1.1567 | 1.1913 | 1.1913 | 1.1913 | 1.1913 | 0.55881 | 0.22665 | 1.1851 | 1.1913 | 1.1913 | 1.1913 | 1.1913 | 1.0483 | 0.94958
-0.089744 |  | 0.92581 | 1.2235 | 1.2235 | 1.2238 | 1.2238 | -0.27035 | -0.55988 | 1.1996 | 1.2238 | 1.2238 | 1.2238 | 1.2238 | 0.5886 | 0.23712 | 1.2199 | 1.2238 | 1.2238 | 1.2238 | 1.2238 | 1.0887 | 0.98529
-0.064103 |  | 0.98901 | 1.2486 | 1.2486 | 1.2487 | 1.2487 | -0.28179 | -0.57834 | 1.2318 | 1.2487 | 1.2487 | 1.2487 | 1.2487 | 0.61197 | 0.24529 | 1.2463 | 1.2487 | 1.2487 | 1.2487 | 1.2487 | 1.1199 | 1.0129
-0.038462 |  | 1.0332 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | -0.28969 | -0.59098 | 1.2534 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | 0.62807 | 0.2509 | 1.264 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | 1.1412 | 1.0318
-0.012821 |  | 1.056 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | -0.29372 | -0.5974 | 1.2641 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | 0.63628 | 0.25375 | 1.273 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | 1.1519 | 1.0414
0.012821 |  | 1.056 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | -0.29372 | -0.5974 | 1.2641 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | 0.63628 | 0.25375 | 1.273 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | 1.1519 | 1.0414
0.038462 |  | 1.0332 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | -0.28969 | -0.59098 | 1.2534 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | 0.62807 | 0.2509 | 1.264 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | 1.1412 | 1.0318
0.064103 |  | 0.98901 | 1.2486 | 1.2486 | 1.2487 | 1.2487 | -0.28179 | -0.57834 | 1.2318 | 1.2487 | 1.2487 | 1.2487 | 1.2487 | 0.61197 | 0.24529 | 1.2463 | 1.2487 | 1.2487 | 1.2487 | 1.2487 | 1.1199 | 1.0129
0.089744 |  | 0.92581 | 1.2235 | 1.2235 | 1.2238 | 1.2238 | -0.27035 | -0.55988 | 1.1996 | 1.2238 | 1.2238 | 1.2238 | 1.2238 | 0.5886 | 0.23712 | 1.2199 | 1.2238 | 1.2238 | 1.2238 | 1.2238 | 1.0887 | 0.98529
0.11538 |  | 0.84708 | 1.1904 | 1.1904 | 1.1913 | 1.1913 | -0.25582 | -0.53618 | 1.1567 | 1.1913 | 1.1913 | 1.1913 | 1.1913 | 0.55881 | 0.22665 | 1.1851 | 1.1913 | 1.1913 | 1.1913 | 1.1913 | 1.0483 | 0.94958
0.14103 |  | 0.75709 | 1.1493 | 1.1493 | 1.1519 | 1.152 | -0.23874 | -0.50796 | 1.1035 | 1.152 | 1.152 | 1.152 | 1.152 | 0.52367 | 0.2142 | 1.1424 | 1.152 | 1.152 | 1.152 | 1.152 | 0.99987 | 0.90672
0.16667 |  | 0.6605 | 1.0993 | 1.0993 | 1.1061 | 1.1065 | -0.21975 | -0.47605 | 1.0403 | 1.1065 | 1.1065 | 1.1065 | 1.1065 | 0.48441 | 0.20017 | 1.0919 | 1.1065 | 1.1065 | 1.1065 | 1.1065 | 0.9445 | 0.85779
0.19231 |  | 0.56204 | 1.0387 | 1.0387 | 1.054 | 1.0557 | -0.19949 | -0.44134 | 0.9682 | 1.0556 | 1.0556 | 1.0557 | 1.0557 | 0.4423 | 0.18495 | 1.0344 | 1.0556 | 1.0556 | 1.0557 | 1.0557 | 0.88354 | 0.80399
0.21795 |  | 0.4661 | 0.96453 | 0.96453 | 0.99508 | 1.0004 | -0.17862 | -0.40475 | 0.88827 | 1.0002 | 1.0002 | 1.0004 | 1.0004 | 0.39863 | 0.16897 | 0.97038 | 1.0003 | 1.0003 | 1.0004 | 1.0004 | 0.81843 | 0.74656
0.24359 |  | 0.37641 | 0.87364 | 0.87364 | 0.9266 | 0.94173 | -0.15773 | -0.36721 | 0.80225 | 0.94104 | 0.94104 | 0.9417 | 0.94173 | 0.35463 | 0.15264 | 0.90059 | 0.94144 | 0.94144 | 0.94172 | 0.94173 | 0.75061 | 0.68676
0.26923 |  | 0.29579 | 0.76431 | 0.76431 | 0.84391 | 0.88055 | -0.13738 | -0.32956 | 0.71223 | 0.87856 | 0.87856 | 0.88041 | 0.88055 | 0.31139 | 0.13634 | 0.82599 | 0.87963 | 0.87963 | 0.88046 | 0.88055 | 0.6815 | 0.62584
0.29487 |  | 0.226 | 0.6386 | 0.6386 | 0.74172 | 0.81784 | -0.11801 | -0.29258 | 0.62064 | 0.81271 | 0.81271 | 0.81722 | 0.81784 | 0.26988 | 0.12041 | 0.74772 | 0.81529 | 0.81529 | 0.81741 | 0.81784 | 0.61248 | 0.56496
0.32051 |  | 0.16779 | 0.50367 | 0.50367 | 0.61798 | 0.7545 | -0.099981 | -0.25695 | 0.53008 | 0.74274 | 0.74274 | 0.75236 | 0.7545 | 0.23088 | 0.10515 | 0.66718 | 0.7482 | 0.7482 | 0.75296 | 0.7545 | 0.54479 | 0.50519
0.34615 |  | 0.12096 | 0.37085 | 0.37085 | 0.47846 | 0.69141 | -0.083543 | -0.22323 | 0.44311 | 0.66726 | 0.66726 | 0.68501 | 0.69141 | 0.19494 | 0.090791 | 0.58595 | 0.6775 | 0.6775 | 0.68663 | 0.69141 | 0.47954 | 0.44746
0.37179 |  | 0.08462 | 0.25241 | 0.25241 | 0.33756 | 0.62935 | -0.06885 | -0.19184 | 0.36207 | 0.58489 | 0.58489 | 0.61306 | 0.62935 | 0.16246 | 0.077513 | 0.50577 | 0.60186 | 0.60186 | 0.61669 | 0.62935 | 0.41764 | 0.39255
0.39744 |  | 0.057419 | 0.15749 | 0.15749 | 0.21331 | 0.56903 | -0.055962 | -0.16308 | 0.28882 | 0.49539 | 0.49539 | 0.53351 | 0.56903 | 0.13363 | 0.065433 | 0.42844 | 0.52025 | 0.52025 | 0.54035 | 0.56903 | 0.35985 | 0.34108
0.42308 |  | 0.037771 | 0.089463 | 0.089463 | 0.11904 | 0.51104 | -0.044862 | -0.13714 | 0.22465 | 0.40088 | 0.40088 | 0.44441 | 0.51104 | 0.10848 | 0.054615 | 0.35567 | 0.43309 | 0.43309 | 0.45524 | 0.51104 | 0.3067 | 0.29352
0.44872 |  | 0.024077 | 0.046016 | 0.046016 | 0.058013 | 0.4559 | -0.03547 | -0.11408 | 0.17022 | 0.30636 | 0.30636 | 0.34761 | 0.4559 | 0.086917 | 0.045074 | 0.28894 | 0.34326 | 0.34326 | 0.362 | 0.4559 | 0.25853 | 0.25014
0.47436 |  | 0.014866 | 0.021337 | 0.021337 | 0.024482 | 0.40367 | -0.027659 | -0.093872 | 0.12551 | 0.21864 | 0.21864 | 0.25024 | 0.40398 | 0.06873 | 0.036782 | 0.2294 | 0.25603 | 0.25603 | 0.2663 | 0.40398 | 0.21551 | 0.21111
0.5 |  | 0.0088881 | 0.0088885 | 0.0088885 | 0.0088885 | 0.0088885 | -0.021273 | -0.076408 | 0.08999 | 0.14429 | 0.14429 | 0.16274 | 0.17779 | 0.053639 | 0.029678 | 0.17779 | 0.17779 | 0.17779 | 0.17779 | 0.17779 | 0.17763 | 0.17642
|  | 0.44651 | 0.72698 | 0.72698 | 0.76287 | 0.8809 | -0.15734 | -0.35388 | 0.74002 | 0.85325 | 0.85325 | 0.86487 | 0.88524 | 0.34982 | 0.14777 | 0.82318 | 0.86273 | 0.86273 | 0.86803 | 0.88524 | 0.71345 | 0.65348
|  | 44.35% | 16.3% | 16.3% | 12.71% | 0.9104% | 104.7% | 124.4% | 15% | 3.675% | 3.675% | 2.513% | 0.4758% | 54.02% | 74.22% | 6.682% | 2.727% | 2.727% | 2.197% | 0.4758% | 17.65% | 23.65%
dx_pfa_hi | 0.17261
2.0863
2.2589
2.4315
2.6041
2.7767
2.9493
3.122
3.2946
3.4672
3.6398
3.8124
3.985
4.1576
4.3302
4.5028
4.6754
4.848
5.0207
5.1933
5.3659
5.5385
5.7111
5.8837
6.0563
6.2289
6.4015
6.5741
6.7467
6.9193
7.092
dx_pfa_lo | 0.043152
0.52158
0.56473
0.60788
0.65103
0.69418
0.73734
0.78049
0.82364
0.86679
0.90995
0.9531
0.99625
1.0394
1.0826
1.1257
1.1689
1.212
1.2552
1.2983
1.3415
1.3846
1.4278
1.4709
1.5141
1.5572
1.6004
1.6435
1.6867
1.7298
1.773
CMC OPTIMIZER — Integration x-grid (midpoints at LOW force)
dx | 0.033333
-0.48333
-0.45
-0.41667
-0.38333
-0.35
-0.31667
-0.28333
-0.25
-0.21667
-0.18333
-0.15
-0.11667
-0.083333
-0.05
-0.016667
0.016667
0.05
0.083333
0.11667
0.15
0.18333
0.21667
0.25
0.28333
0.31667
0.35
0.38333
0.41667
0.45
0.48333

## Validation
Formula Validation & Proof Report
Morehouse True Cost Calculator v17.5
Tol=10.0 lbf ¦ EOPR=85% ¦ N=1,000 ¦ Force=10,000 lbf ¦ Rate=$200/hr ¦ 2026-05-05
Reference: scipy.integrate.quad · Sandia PSL PFA Risk Calculator (SAND2025-10736O)
TABLE 1: DASHBOARD PFR vs SCIPY (7×3 = 21 PFR checks + 21 rework checks)
Machine | CMC | TUR | Excel PFR_G | Scipy PFR_G | Δ |  | Excel PFR_M6 | Scipy PFR_M6 | Δ |  | Excel PFR_S | Scipy PFR_S | Δ
BCM | 0.04% | 2.5 | 5.216% | 4.169% | 1.047% | PASS | 7.061% | 7.039% | 0.02243% | PASS | 25.1% | 25.1% | 0.003057% | PASS
PCM | 0.033% | 3.0303 | 3.545% | 3.301% | 0.2436% | PASS | 4.84% | 4.812% | 0.02768% | PASS | 19.39% | 19.39% | 0.003601% | PASS
UCM Manual | 0.03% | 3.3333 | 2.981% | 2.948% | 0.03305% | PASS | 4.003% | 3.972% | 0.03063% | PASS | 17.11% | 17.1% | 0.00391% | PASS
UCM Auto | 0.02% | 5 | 1.893% | 1.848% | 0.04504% | PASS | 1.893% | 1.848% | 0.04504% | PASS | 10.24% | 10.23% | 0.005588% | PASS
Deadweight | 0.002% | 50 | 0.541% | 0.1647% | 0.3763% | PASS | 0.541% | 0.1647% | 0.3763% | PASS | 1.039% | 0.8233% | 0.2159% | PASS
Competitor 1 | 0.05% | 2 | 8.085% | 5.51% | 2.575% | PASS | 10.87% | 10.85% | 0.01746% | PASS | 34.05% | 34.05% | 0.002541% | PASS
Competitor 2 | 0.06% | 1.6667 | 6.465% | 6.967% | 0.5014% | PASS | 15.12% | 15.37% | 0.2503% | PASS | 43.68% | 43.7% | 0.02392% | PASS
TABLE 2: REWORK = PFR × N × $Cost (7 checks)
Machine | PFR (M6) | × Cals | × $Cost | Expected | Excel | Δ | Status
BCM | 7.061% | 1000 | 100 | 7061.174 | 7061.174 | 0 | PASS
PCM | 4.84% | 1000 | 100 | 4839.5276 | 4839.5276 | 0 | PASS
UCM Manual | 4.003% | 1000 | 100 | 4002.552 | 4002.552 | 0 | PASS
UCM Auto | 1.893% | 1000 | 100 | 1892.8336 | 1892.8336 | 0 | PASS
Deadweight | 0.541% | 1000 | 100 | 541.0064 | 541.0064 | 1.1369e-13 | PASS
Competitor 1 | 10.87% | 1000 | 100 | 10868.7975 | 10868.7975 | 0 | PASS
Competitor 2 | 15.12% | 1000 | 100 | 15121.1237 | 15121.1237 | 0 | PASS
TABLE 3: GUARD BANDS — Deterministic (20 checks)
TUR | U | Excel GB_Spec | Calc GB_Spec |  | Excel GB_M6 | Calc GB_M6 |  | Consumed
1.5 | 6.6667 | 6.5333 | 6.5333 | PASS | 2.4012 | 2.4012 | PASS | 65.33%
2 | 5 | 4.9 | 4.9 | PASS | 1.4082 | 1.4082 | PASS | 49%
2.5 | 4 | 3.92 | 3.92 | PASS | 0.85814 | 0.85814 | PASS | 39.2%
3 | 3.3333 | 3.2667 | 3.2667 | PASS | 0.51773 | 0.51773 | PASS | 32.67%
3.5 | 2.8571 | 2.8 | 2.8 | PASS | 0.29128 | 0.29128 | PASS | 28%
4 | 2.5 | 2.45 | 2.45 | PASS | 0.1328 | 0.1328 | PASS | 24.5%
4.6 | 2.1739 | 2.1304 | 2.1304 | PASS | 0 | 0 | PASS | 21.3%
5 | 2 | 1.96 | 1.96 | PASS | 0 | 0 | PASS | 19.6%
10 | 1 | 0.98 | 0.98 | PASS | 0 | 0 | PASS | 9.8%
24 | 0.41667 | 0.40833 | 0.40833 | PASS | 0 | 0 | PASS | 4.083%
TABLE 4: ORDERING — Global ≤ M6 ≤ Specific (7 checks)
Machine | RW Global | RW M6 | RW Specific | G≤M6 | M6≤S | Status
BCM | 5215.792 | 7061.174 | 25100.7166 | ✓ | ✓ | PASS
PCM | 3544.7645 | 4839.5276 | 19392.2608 | ✓ | ✓ | PASS
UCM Manual | 2980.6053 | 4002.552 | 17106.6567 | ✓ | ✓ | PASS
UCM Auto | 1892.8336 | 1892.8336 | 10238.5617 | ✓ | ✓ | PASS
Deadweight | 541.0064 | 541.0064 | 1039.1573 | ✓ | ✓ | PASS
Competitor 1 | 8085.3146 | 10868.7975 | 34053.0739 | ✓ | ✓ | PASS
Competitor 2 | 6465.3029 | 15121.1237 | 43680.5922 | ✓ | ✓ | PASS
TABLE 5: CROSS-SHEET REFERENCES (15 checks)
Reference | Description | Value | Expected | Status
Risk Calc B5 | Tol | 10 | 10 | PASS
Risk Calc B6 | EOPR | 0.85 | 0.85 | PASS
Risk Calc B7 | σ_p | 6.9467 | 6.9467 | PASS
Risk Calc B8 | N | 1000 | 1000 | PASS
Risk Calc B9 | RW$ | 100 | 100 | PASS
RA B9 | EOPR | 0.85 | 0.85 | PASS
RA B10 | N | 1000 | 1000 | PASS
RA B11 | RW$ | 100 | 100 | PASS
RA E5 | TolHi | 10 | 10 | PASS
RA E6 | TolLo | 0.025 | 0.025 | PASS
BRK B27 | N | 1000 | 1000 | PASS
BRK B28 | Rate | 200 | 200 | PASS
Dash B9 | RW$ | 100 | 100 | PASS
Dash B16 | Tol | 10 | 10 | PASS
BRK O31 | Hrs/Cal link | 0.9 | 0.9 | PASS
TABLE 6: BREAKEVEN — Rev($150/cal) − Labor($200/hr) − OpCost = Profit
Machine | Price | Hrs/Cal | Labor/yr | OpCost | Revenue | Profit M6 | Check | Payback | Req$/Cal | Req$/Hr | Status
BCM | 35000 | 1.5 | 300000 | 42867.7719 | 150000 | -192867.7719 | -192867.7719 | Never | 354.5344 | 236.3563 | PASS
PCM | 25000 | 1.5 | 300000 | 30263.5085 | 150000 | -180263.5085 | -180263.5085 | Never | 338.5968 | 225.7312 | PASS
UCM Manual | 50000 | 1 | 200000 | 12868.7975 | 150000 | -62868.7975 | -62868.7975 | Never | 229.5355 | 229.5355 | PASS
UCM Automated | 100000 | 0.5 | 100000 | 7436.1928 | 150000 | 42563.8072 | 42563.8072 | 2.3yr | 140.7695 | 281.5391 | PASS
Deadweight | 300000 | 1 | 200000 | 541.0064 | 150000 | -50541.0064 | -50541.0064 | Never | 300.541 | 300.541 | PASS
Competitor | 30000 | 1.5 | 300000 | 56122.264 | 150000 | -206122.264 | -206122.264 | Never | 366.1223 | 244.0815 | PASS
Competitor 2 | 40000 | 1.5 | 300000 | 75059.1029 | 150000 | -225059.1029 | -225059.1029 | Never | 388.3924 | 258.9283 | PASS
TABLE 7: STANDARD SWITCHING — Swap+Cal+RW=Total (3 checks)
MinLoad% | Swap$/yr | Cal$/yr | RW M6 | Expected | Excel | Δ | Status
5% | 0 | 2000 | 4936.835 | 6936.835 | 6936.835 | 0 | PASS
10% | 0 | 4000 | 2765.5605 | 6765.5605 | 6765.5605 | 0 | PASS
20% | 0 | 6000 | 1892.8336 | 7892.8336 | 7892.8336 | 0 | PASS
TABLE 8: CELL CHAIN — 10,000→25 lbf at 5% Loading
Check | Excel | Expected | Status
Cells needed | 2 | 2 | PASS
Swaps/run | 1 | 1 | PASS
Cell 1 | 10,000 lbf | → 500 lbf | No — need next cell | INFO
Cell 2 | 500 lbf | → 25 lbf | ✓ Yes | PASS
VALIDATION SUMMARY
Result | ✅ ZERO ERRORS — ALL 5,716 CHECKS PASSED
Total formulas | 5,588
Validation checks | 5716
Errors | 0
Warnings | 0
M6/Specific accuracy | < 0.05% vs scipy at TUR ≥ 1.5
Global accuracy | < 3% (40-pt trapezoidal)
Guard bands | Exact match — 10 TUR levels
Breakeven | Includes labor (Hrs/Cal on Range Analysis col C → Breakeven)
Known limitation | Deadweight TUR 50: Excel 0.54% vs scipy 0.16%
Reference | Sandia PSL PFA Risk Calculator (SAND2025-10736O)
INPUT CELL LEGEND
Editable input | Sage shading + green bold + thin green border
Calculated | No shading — formula cell, do not edit