## Executive Summary
TRUE COST OF MEASUREMENT DECISIONS
Executive Summary Report | [7 empty] | 1
Morehouse Instrument Company — ISO/IEC 17025 Accredited — NVLAP 600259 — A2LA 1398.01 |  |  |  | Measurement Risk Scenario Selected |  |  | Global (PFA ≤ 2%)
KEY PARAMETERS
Tolerance | EOPR | Labor Rate | Rework Cost | Cals/Year | Price/Cal | Revenue/yr | Swap Req?
0.100% | 89% | $200/hr | $116.67 | 2,000 | $500 | $1,000,000 | Yes
RECOMMENDATION
Best Machine: | UCM Automated |  | Annual Profit: | 632874.3 |  | Payback: | 0.2 years
Worst Available: | Competitor 2 |  | Annual Profit: | 340811.6384 |  | Spread: | 292062.6616 | /yr
At 0.100% tolerance, 89% EOPR, 2,000 cals/yr at $200/hr: UCM Automated is the best machine at $632,874/yr profit, 0.2 years payback. Worst available: Competitor 2 at $340,812/yr profit.
Best option: UCM Automated at $632,874/yr profit, 0.2 years payback. Saves $292,063/yr vs worst (Competitor 2).
Switching from Competitor 2 ($340,812 profit) to UCM Automated ($632,874 profit) saves $292,063/yr, payback in 0.2 years.
MACHINE COST & PROFITABILITY COMPARISON (Global (PFA ≤ 2%) Decision Rule) — Applied Force: 5,000 lbf
Machine | TUR | Labor/yr | Rework/yr | Cal Cost/yr | Swap/yr | Total Cost/yr | Profit/yr | Payback
BCM | 2 | 360000 | 14874.9718 | 4000 | 200000 | 579591.7939 | 420408.2061 | 0.1 yr
PCM | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
UCM Manual | 3.3333 | 400000 | 6361.9159 | 4000 | 200000 | 610361.9159 | 389638.0841 | 0.1 yr
UCM Automated | 5 | 360000 | 3977.1373 | 2000 | 0 | 367125.7 | 632874.3 | 0.2 yr
Deadweight | 50 | 480000 | 1110.1212 | 0 | 0 | 481110.1212 | 518889.8788 | 0.6 yr
Competitor 1 | 2 | 400000 | 14874.9718 | 4000 | 200000 | 645194.5924 | 354805.4076 | 0.2 yr
Competitor 2 | 1.6667 | 400000 | 14528.1178 | 4000 | 200000 | 659188.3616 | 340811.6384 | 0.1 yr
SAVINGS ANALYSIS — CURRENT vs. RECOMMENDED
| Worst Available | Best (Recommended) | Annual Savings
Machine | Competitor 2 | UCM Automated
Total Cost/yr | 659188.3616 | 367125.7 | 292062.6616
Rework Cost/yr | 14528.1178 | 3977.1373 | 10550.9805
Profit/yr | 340811.6384 | 632874.3 | 292062.6616
CUMULATIVE SAVINGS PROJECTIONS
| Investment Delta | Annual Savings | Payback | Year 1 | Year 3 | Year 5
Upgrade: Competitor 2 → UCM Automated | 60000 | 292062.6616 | 0.2 yr | 232062.6616 | 816187.9849 | 1400313.3082
Generated by the Morehouse True Cost Calculator. Figures are estimates based on user inputs. Decision rule: Global (PFA ≤ 2%).
Report generated: June 3, 2026

## Graphs Summary
THE TRUE COST OF MEASUREMENT DECISIONS — EXECUTIVE SUMMARY
Morehouse Instrument Company — ISO/IEC 17025 Accredited — NVLAP 600259 — A2LA 1398.01
Tolerance: | 0.100% | EOPR: | 89% | Labor: | $200/hr | Cals/yr: | 2,000 | Swap: | Yes | Price/Cal: | $500
RECOMMENDATION
✔ UCM AUTOMATED — $632,874/yr PROFIT
At 0.100% tolerance, 89% EOPR, 2,000 cals/yr at $200/hr: 6 of 6 available machines are profitable (M6). Best: UCM Automated at $632,874/yr profit, 0.2-yr payback. Worst available: Competitor 2 at $340,812/yr profit.
Best option: UCM Automated at $632,874/yr profit, 0.2-yr payback. Saves $113,984/yr vs next best (Deadweight).
Switching from Competitor 2 ($340,812 profit) to UCM Automated ($632,874 profit) saves $292,063/yr, payback in 0.2 years.
TUR range: 1.7:1 to 50.0:1 (6 machines available). WARNING: At least one available machine below TUR 2:1. PCM not available above 2,000 lbf. Swapping active — swap cost may dominate.
Annual Rework by Decision Rule | [5 empty] | Annual Profit by Machine | Annual Profit by Machine
Total Annual Cost (Rework + Cal + Swap) | [5 empty] | Payback Period (Years)
Standard Switching: Cost by Loading Strategy
Breakeven Cals Chart Data | Global | Method 6 | Specific Risk (PFA = 2.5 %)
BCM | 118.932 | 124.6865 | 150.3673
PCM | 163.2255 | 164.4869 | 178.3262
UCM Manual | 256.6484 | 258.1148 | 279.3585
UCM Automated | 316.0185 | 316.3879 | 328.7649
Deadweight | 1156.3147 | 1156.3147 | 1158.5792
Competitor | 310.0291 | 407.4707 | 584.0708
Competitor 2 | 234.7338 | 359.0438 | 424.7788
Labor: Hrs/Cal editable on Range Analysis (col C). Includes setup, loading, and documentation. For C+T calibrations, roughly 1.7× single-mode time.
MACHINE COMPARISON AT A GLANCE
Machine | CMC (%) | TUR | PFR (M6) | Rework/yr | Cal/yr | Swap/yr | Total Cost | Revenue | Profit (M6) | Payback | Viable? | Hrs/Cal | Labor/yr
BCM | 0.05% | 2 | 10.4% | 14874.9718 | 4000 | 200000 | 219591.7939 | 1000000 | 420408.2061 | 0.1 yr | ⚠ 2.0 | 0.9 | 360000
PCM | 0.033% | 3.0303 | 4.472% | 7145.5244 | 4000 | 200000 | 211145.5244 | 1000000 | 428854.4756 | 0.1 yr | N/A | 0.9 | 360000
UCM Manual | 0.03% | 3.3333 | 3.675% | 6361.9159 | 4000 | 200000 | 210361.9159 | 1000000 | 389638.0841 | 0.1 yr | ⚠ 3.3 | 1 | 400000
UCM Automated | 0.02% | 5 | 1.704% | 3977.1373 | 2000 | 0 | 7125.7 | 1000000 | 632874.3 | 0.2 yr | ✓ 5.0 | 0.9 | 360000
Deadweight | 0.002% | 50 | 0.4758% | 1110.1212 | 0 | 0 | 1110.1212 | 1000000 | 518889.8788 | 0.6 yr | ✓ 50.0 | 1.2 | 480000
Competitor 1 | 0.05% | 2 | 10.4% | 14874.9718 | 4000 | 200000 | 245194.5924 | 1000000 | 354805.4076 | 0.2 yr | ⚠ 2.0 | 1 | 400000
Competitor 2 | 0.06% | 1.6667 | 14.77% | 14528.1178 | 4000 | 200000 | 259188.3616 | 1000000 | 340811.6384 | 0.1 yr | ✗ 1.7 | 1 | 400000
Generated by the Morehouse True Cost Calculator. Figures are estimates based on user inputs.

## Instructions
The True Cost of Measurement Decisions
Calculator & Business Case Builder
Morehouse Instrument Company — mhforce.com — (717) 843-0081
Risk formulas validated against Sandia Primary Standards Lab bivariate normal integral (scipy dblquad).
COLOR KEY
Green text | Editable input — you enter these values
Black text | Calculated — do not edit (formula will break)
Green headers | Column headers and section titles
Gray italic text | Notes, guidance, and documentation
SHEET GUIDE
Dashboard | Start here. Enter cost inputs (labor rate, rework time, cals/yr), measurement inputs (tolerance, force), switching options, and profitability inputs (cal cost, price/cal, payback target). CMC per machine is a green input — changing it updates Range Analysis, Risk Calc, and all charts. Three charts: Annual Rework, Annual Profit, and Payback Period. Standard Switching summary and machine payback data at bottom.
Summary | Executive one-page overview. Dynamic recommendation paragraph, 5 charts (Rework, Profit, Total Cost, Payback, Standard Switching), and machine comparison table. All formula-driven — updates when Dashboard inputs change.
Range Analysis | Enter your actual force range (high/low), tolerance at each end, and CMC at each end per machine. CMC@High (col E) is wired to Dashboard — edit there. CMC@Low (col H) is a separate input on this sheet. Dual-point rework uses the WORST of PFR@High and PFR@Low. Includes the CMC Optimizer (Target TUR Reverse Calculator), Cell Sizing Calculator (auto cells/swaps), What-If Tolerance comparison, and CMC Sensitivity table with Savings vs BCM analysis.
Breakeven | Full business case: machine prices, upgrade deltas, profitability per machine (Revenue − Equipment Cost − Labor), payback periods. Includes Hrs/Cal per machine for labor cost modeling. Charts: Annual Profit and Payback Period.
Guard Band Analysis | Shows tolerance consumed by Specific Risk and Method 6 guard bands at 10 TUR levels. Chart: Tolerance Consumed by Specific Risk Guard Band vs M6 Guard Band. At TUR 4.6:1, Method 6 guard band drops to zero — full tolerance usable.
Standard Switching | Compares loading strategies (5%, 10%, 20%) with swap costs. Shows cells needed, swaps, swap cost, cal cost, rework, and total per loading level. Stacked bar chart: Annual Cost by Loading Strategy.
Risk Calc | The math engine. Bivariate normal integration for PFR, bisection solver for Global guard band, Method 6 deterministic formula, Specific Risk fixed multiplier. Single-point (Dashboard) and dual-point (Range Analysis) grids. CMC Sensitivity x-grid for the cost curve. Do not edit unless you understand the math.
Validation | Scipy comparison for all formulas. Used during development to verify accuracy.
INPUT REFERENCE
COST INPUTS |  | Cell
Hourly Labor Rate | Fully loaded labor cost including overhead. Default: $150/hr. | Dashboard B7
Rework Time | Minutes per false reject event (investigation, retest, documentation). Default: 45 min. | Dashboard B8
Calibrations per Year | Total calibrations performed annually. Drives all cost calculations. Default: 1,200. | Dashboard B10
EOPR | End-of-Period Reliability: fraction in tolerance at end of cal interval. Default: 89% (ANSI/NCSL Z540.3 threshold, NASA standard). | Dashboard B11
Show Simple Acceptance? | Yes/No toggle. Shows/hides the 4th decision rule (shared risk, no guard band) on charts. Default: No. | Dashboard B12
MEASUREMENT INPUTS |  | Cell
Device Tolerance (%) | Manufacturer's specification as a % of applied force. E.g., ±0.1% = 0.001. Flows to Range Analysis Tol@High and Risk Calc. | Dashboard B14
Applied Force (lbf) | The force point for the single-point comparison on the Dashboard. | Dashboard B15
What-If Tolerance (%) | A second tolerance for side-by-side TUR comparison. E.g., 0.025% for reference-grade. | Range Analysis B21
SWITCHING OPTIONS |  | Cell
Standard Switching Required? | Yes/No dropdown. "Yes" = lab swaps reference cells mid-run (auto-calculated swap cost). "No" = matched capacity references, zero swaps. | Dashboard B18
Swap Time per Change (min) | Time to swap one reference cell. Default: 30 min. | Dashboard B19
PROFITABILITY INPUTS |  | Cell
Cal Cost per Cell ($/yr) | Annual calibration cost per reference load cell. Default: $2,000. | Dashboard B21
Price per Calibration ($) | What you charge customers per calibration. Drives revenue and profitability. | Dashboard B22
Target Payback (years) | Desired years to recoup machine investment. | Dashboard B23
MACHINE INPUTS |  | Cell
CMC (%) per machine | Calibration and Measurement Capability at capacity. Green inputs on Dashboard rows 25–31. Flows to Range Analysis CMC@High. | Dashboard B25:B31
CMC@Low (%) per machine | CMC at the low force point. Separate from Dashboard CMC. Must be set on Range Analysis for accurate dual-point results. | Range Analysis H24:H30
Max Capacity (lbf) per machine | Maximum force capacity per machine. Machines exceeding Applied Force are marked N/A and zeroed on all charts. | Dashboard T25:T31
Hrs/Cal per machine | Average hours per calibration per machine. Includes setup, loading, documentation. Drives labor cost. | Breakeven O31:O37
Target TUR for Cell Sizing | TUR each reference cell must maintain across the range. Drives auto cell/swap count. Lower = fewer cells, more rework. | Range Analysis B13
DECISION RULES EXPLAINED
Global (PFA ≤ 2%) | Guard band width set so average PFA across the population is ≤ 2%. Uses bisection solver (10 iterations). Most permissive — widest acceptance zone.
Method 6 (Dobbert) | Guard band from the Dobbert formula: M = max(0, 1.04 − e^(0.38·ln(TUR) − 0.54)). At TUR ≥ 4.6:1, guard band = 0. ANSI/NCSL Z540.3 default.
Specific Risk | Guard band = 0.98 × U. Most conservative — controls PFA to ≤ 2.5% per side at each individual measurement. Highest rework cost.
Simple Acceptance (Shared Risk) | No guard band (GB = 0). Full tolerance as acceptance limit. Lowest PFR but highest PFA — customer bears all risk. Up to 50% PFA at low TUR. Toggle: Dashboard B12. Default: hidden.
HOW TO USE THE CMC OPTIMIZER
1. Go to Range Analysis, rows 14-20 (CMC OPTIMIZER section).
2. Enter your desired TUR at the low force point in B17. The calculator shows the Required CMC and U.
3. It also shows the TUR at the high force point with the same CMC — this tells you how much headroom you have.
4. The CMC Sensitivity table (rows 52-68) shows annual cost at 14 CMC levels.
5. "Savings vs BCM" (col M) shows cumulative savings from upgrading beyond 0.04% CMC.
6. "Savings per 0.01% CMC" (col N) shows the rate of return — when this drops, further improvement yields diminishing returns.
7. The Cell Sizing Calculator (rows 75+) shows how many reference cells each machine needs at the Target TUR.
CELL SIZING & STANDARD SWITCHING
The calculator auto-sizes reference cells using a CMC curve model estimated from CMC@High and CMC@Low.
It calculates the force point where each machine's CMC crosses the Required CMC (= Tolerance / Target TUR).
From this "swap force," it determines cells needed to cover the full range.
When "Standard Switching Required?" = Yes, Swaps = Cells − 1. When No, Swaps = 0 (matched capacity).
Change Target TUR on Range Analysis B13 to explore: lower TUR = fewer cells but more rework.
EOPR BASELINE
Default: 89%. This is the ANSI/NCSL Z540.3 threshold (PFA ≤ 2% at TUR 4.6:1) and the NASA standard.
EOPR should be based on: (1) OEM stated reliability at the cal interval, or (2) historical drift data when worse than OEM spec.
The commonly cited 85% originated from a U.S. Navy Corona fleet-wide assumption across all M&TE and does not reflect controlled force calibration equipment.
TECHNICAL NOTES
• PFA and PFR are unconditional probabilities computed via numerical integration of the bivariate normal (Sandia PSL method).
• Integration uses 40-point trapezoidal rule for PFR (accuracy < 0.05% vs scipy at TUR ≥ 1.5) and 30-point midpoint for PFA.
• EOPR determines the process standard deviation: σ_product = Tolerance / NORMINV((1+EOPR)/2).
• All four decision rules use the same EOPR and the same integral — only the guard band formula differs. Simple Acceptance uses GB = 0.
• Dashboard inputs (Tolerance, CMC, EOPR, labor, cals) flow to Range Analysis and Risk Calc. CMC@Low and Tol@Low are separate RA inputs for the dual-point analysis.
• Cell sizing uses a 1/F + offset CMC model estimated from the two data points (CMC@High and CMC@Low).
• No Excel LAMBDA functions are used. Compatible with Excel 2016+ and LibreOffice (use NORMSINV, not NORM.S.INV).
• Validated against Sandia PSL PFA Risk Calculator (SAND2025-10736O) and scipy.integrate.dblquad.

## Dashboard
The True Cost of Measurement Decisions
Calculator from Morehouse Instrument Company
Risk calculated via Sandia PSL bivariate normal integral (no lookup tables).
▶ START HERE — Enter values in GREEN cells below, then check Range Analysis for any Machine Pricing Adjustment + Summary tab for report |  |  |  | 💡 Hover any green input cell for a description, or see the Instructions tab for full guidance.
COST INPUTS
Hourly Labor Rate ($/hr) | 200
Rework Time per Event (min) | 35
Rework Cost per Event ($) | 116.6667
Calibrations per Year | 2000
EOPR (In-Tolerance Probability) | 89%
Show Simple Acceptance? | Yes
MEASUREMENT INPUTS
Device Tolerance (%) | 0.1%
Applied Force (lbf) | 5000
Tolerance in Force Units (lbf) | 5
SWITCHING OPTIONS
Standard Switching Required? | Yes
Swap Time per Change (min) | 30
PROFITABILITY INPUTS
Cal Cost per Cell ($/yr) | 2000
Price per Calibration ($) | 500
Target Payback (years) | 3
Machine | CMC (%) | U (lbf) | TUR | Price ($) | CMC Low (%) | Hrs/Cal | GB Global | PFA Glob | PFR Glob | RW Glob | GB M6 | PFA M6 | PFR M6 | RW M6 | GB Spec | PFA Global | PFR Spec | RW Spec | FR/yr Glob | FR/yr M6 | FR/yr Spec | Max Cap (lbf) | Available? | PFR Simple | RW Simple | PFA Specific (G8)
BCM | 0.05% | 2.5 | 2 | 25000 | 0.06% | 0.9 | 0.18901 | 1.997% | 6.375% | 14874.9718 | 0.70411 | 1.135% | 10.4% | 24265.9842 | 2.45 | 0.06773% | 33.98% | 79284.1758 | 127.4998 | 207.9942 | 679.5786 | 10000 | ✓ | 5.218% | 12176.3971 | 2.5%
PCM | 0.033% | 1.65 | 3.0303 | 35000 | 0.03% | 0.9 | 0 | 1.77% | 3.062% | 7145.5244 | 0.25069 | 1.199% | 4.472% | 10434.0816 | 1.617 | 0.04798% | 18.62% | 43453.6065 | 61.2474 | 89.435 | 372.4595 | 2000 | N/A — 2,000 lbf max | 3.062% | 7145.5244 | 2.5%
UCM Manual | 0.03% | 1.5 | 3.3333 | 50000 | 0.03% | 1 | 0 | 1.64% | 2.727% | 6361.9159 | 0.17877 | 1.212% | 3.675% | 8575.4868 | 1.47 | 0.04418% | 16.3% | 38028.9887 | 54.5307 | 73.5042 | 325.9628 | 1000000 | ✓ | 2.727% | 6361.9159 | 2.5%
UCM Automated | 0.02% | 1 | 5 | 100000 | 0.025% | 0.9 | 0 | 1.155% | 1.704% | 3977.1373 | 0 | 1.155% | 1.704% | 3977.1373 | 0.98 | 0.03074% | 9.48% | 22121.1057 | 34.0897 | 34.0897 | 189.6095 | 1000000 | ✓ | 1.704% | 3977.1373 | 2.5%
Deadweight | 0.002% | 0.1 | 50 | 300000 | 0.002% | 1.2 | 0 | 4.028e-05% | 0.4758% | 1110.1212 | 0 | 4.028e-05% | 0.4758% | 1110.1212 | 0.098 | 0.003325% | 0.7225% | 1685.7521 | 9.5153 | 9.5153 | 14.4493 | 1000000 | ✓ | 0.4758% | 1110.1212 | 2.5%
Competitor 1 | 0.05% | 2.5 | 2 | 55000 | 0.12% | 1 | 0.18901 | 1.997% | 6.375% | 14874.9718 | 0.70411 | 1.135% | 10.4% | 24265.9842 | 2.45 | 0.06773% | 33.98% | 79284.1758 | 127.4998 | 207.9942 | 679.5786 | 1000000 | ✓ | 5.218% | 12176.3971 | 2.5%
Competitor 2 | 0.06% | 3 | 1.6667 | 40000 | 0.15% | 1 | 0 | 0% | 6.226% | 14528.1178 | 0.99722 | 0% | 14.77% | 34455.8465 | 2.94 | 0% | 44.33% | 103427.6176 | 124.5267 | 295.3358 | 886.5224 | 1000000 | ✓ | 6.226% | 14528.1178 | 2.5%
All PFA/PFR computed via Sandia PSL bivariate normal integral (trapezoidal rule, no lookup tables).
Global: guard band set so PFA ≤ target (bisection solver). M6: Dobbert formula. Specific: h=0.98×U.
EOPR determines σ_product. All three models use same EOPR and same integral—only guard band width differs.
CHART DATA
Machine | Global (PFA≤2%) | Method 6 | Specific Risk (PFA = 2.5 %) |  |  |  | Simple Accept
BCM | 14874.9718 | 24265.9842 | 79284.1758 |  |  |  | 12176.3971
PCM (N/A) | 0 | 0 | 0 |  |  |  | 0
UCM Manual | 6361.9159 | 8575.4868 | 38028.9887 |  |  |  | 6361.9159
UCM Automated | 3977.1373 | 3977.1373 | 22121.1057 |  |  |  | 3977.1373
Deadweight | 1110.1212 | 1110.1212 | 1685.7521 |  |  |  | 1110.1212
Competitor 1 | 14874.9718 | 24265.9842 | 79284.1758 |  |  |  | 12176.3971
Competitor 2 | 14528.1178 | 34455.8465 | 103427.6176 |  |  |  | 14528.1178
MACHINE PAYBACK DATA (capped at 20yr for display)
Machine | Payback (Global) | Payback (M6) | Payback (Specific) |  |  |  | Payback (Simple)
BCM | 0.059466 | 0.062343 | 0.075184 |  |  |  | 0.058987
PCM (N/A) | 0 | 0 | 0 |  |  |  | 0
UCM Manual | 0.12832 | 0.12906 | 0.13968 |  |  |  | 0.12832
UCM Automated | 0.15801 | 0.15819 | 0.16438 |  |  |  | 0.15772
Deadweight | 0.57816 | 0.57816 | 0.57929 |  |  |  | 0.57816
Competitor | 0.15501 | 0.20374 | 0.29204 |  |  |  | 0.14329
Competitor 2 | 0.11737 | 0.17952 | 0.21239 |  |  |  | 0.10486
ANNUAL PROFIT DATA
Machine | Profit (Global) | Profit (M6) | Profit (Specific) |  |  |  | Profit (Simple)
BCM | 420408.2061 | 401005.7532 | 332519.2058 |  |  |  | 423823.6029
PCM (N/A) | 0 | 0 | 0 |  |  |  | 0
UCM Manual | 389638.0841 | 387424.5132 | 357962.9664 |  |  |  | 389638.0841
UCM Automated | 632874.3 | 632135.5008 | 608337.3553 |  |  |  | 634022.8627
Deadweight | 518889.8788 | 518889.8788 | 517875.6755 |  |  |  | 518889.8788
Competitor | 354805.4076 | 269958.0932 | 188333.3333 |  |  |  | 383823.6029
Competitor 2 | 340811.6384 | 222814.0243 | 188333.3333 |  |  |  | 381471.8822
STANDARD SWITCHING SUMMARY
Min Loading | Cells | Swaps/Run | Swap $/Cal |  |  |  | Swap $/yr | Cal $/yr | Total M6/yr
5% (UPC) | 1 | 0 | 0 |  |  |  | 0 | 2000 | 12651.333
10% (PC) | 2 | 1 | 100 |  |  |  | 200000 | 4000 | 209864.4992
20% (Commercial) | 3 | 2 | 200 |  |  |  | 400000 | 6000 | 409977.1373
Deadweight | 0 | 0 | 0 |  |  |  | 0 | 0 | 0
RECOMMENDATION HELPERS (hidden from print)
Machine | Available? | Profit M6 | Payback M6 |  |  |  | Profit Rank | Worst Helper | TUR Helper | TUR Max Help
BCM | TRUE | 420408.2061 | 0.059466 |  |  |  | 420408.2061 | 420408.2061 | 2 | 2
PCM | FALSE | [5 empty] | -9999999 | 9999999 | 9999 | 0
UCM Manual | TRUE | 389638.0841 | 0.12832 |  |  |  | 389638.0841 | 389638.0841 | 3.3333 | 3.3333
UCM Automated | TRUE | 632874.3 | 0.15801 |  |  |  | 632874.3 | 632874.3 | 5 | 5
Deadweight | TRUE | 518889.8788 | 0.57816 |  |  |  | 518889.8788 | 518889.8788 | 50 | 50
Competitor 1 | TRUE | 354805.4076 | 0.15501 |  |  |  | 354805.4076 | 354805.4076 | 2 | 2
Competitor 2 | TRUE | 340811.6384 | 0.11737 |  |  |  | 340811.6384 | 340811.6384 | 1.6667 | 1.6667
Available count | 6
Profitable (avail) | 6
Best profit | 632874.3 | UCM Automated | 0.15801
Worst profit (avail) | 340811.6384 | Competitor 2
Min TUR (avail) | 1.6667
Max TUR (avail) | 50
REQUIRED $/CAL TO HIT 3-YEAR PAYBACK
Machine | Machine Price | Current $/Cal | Global (PFA ≤ 2%) | Method 6 (Dobbert) | Specific Risk (PFA ≤ 2%) | Simple Acceptance | Capacity (lbf)
BCM | 25000 | 500 | 293.9626 | 303.6638 | 337.9071 | 292.2549 | 10000
PCM | 35000 | 500 | N/A | N/A | N/A | N/A | 2000
UCM Manual | 50000 | 500 | 313.5143 | 314.6211 | 329.3519 | 313.5143 | 1000000
UCM Automated | 100000 | 500 | 200.2295 | 200.5989 | 212.498 | 199.6552 | 1000000
Deadweight | 300000 | 500 | 290.5551 | 290.5551 | 291.0622 | 290.5551 | 1000000
Competitor 1 | 55000 | 500 | 331.764 | 374.1876 | 415 | 317.2549 | 1000000
Competitor 2 | 40000 | 500 | 336.2608 | 395.2597 | 412.5 | 315.9307 | 1000000
Note on A12: henry zumbrun: Simple Acceptence may not be applicable as in most cases it does not account for measurement uncertainty.
Note on A18: henry zumbrun: If you intend to use multiple load cells, check yes
Note on Q24: henry zumbrun: Global consumer's risk R_C (JCGM 106 eq.19), from Risk Calc row 165. Process-level risk across all items. For the per-result ILAC G8 specific risk see the PFA Specific (G8) column.
Note on AA24: henry zumbrun: ILAC G8:2009 §2.2 specific (pointwise) risk at the acceptance limit = 2.5% (pc=0.975) for the 0.98×U guard band. From Risk Calc row 58.

## Range Analysis
Range Analysis — Full Cost at High and Low Force Points
Enter your operating range, CMC at each end, swaps per machine. Risk via Sandia bivariate integral.
RANGE INPUTS |  |  |  | DERIVED
High Force (lbf) | 5000 |  |  | Tol @ High (lbf) | 5
Low Force (lbf) | 500 |  |  | Tol @ Low (lbf) | 0.5
Tolerance @ High (%) | 0.1% |  |  | σ_p @ High | 3.1285
Tolerance @ Low (%) | 0.1% |  |  | σ_p @ Low | 0.31285
EOPR | 89%
Calibrations per Year | 2000
Rework Cost per Event ($) | 116.6667
Swap Time per Change (min) | 30
Target TUR for Cell Sizing | 4 | Forward sizing: pick the TUR each reference cell must hold across the range. Required CMC = Tolerance (0.100 %) ÷ Target TUR. Lower TUR = fewer cells but more rework risk.
CMC OPTIMIZER
TARGET TUR REVERSE CALCULATOR
Desired TUR at Low Force | 4 | To achieve your target TUR at the low force point,
Required CMC (%) | 0.025% | you need a reference standard with this CMC or better.
Required U at Low (lbf) | 0.125
TUR at High Force (same CMC) | 4
What-If Tolerance (%) | 0.025% | Enter a second tolerance to compare (e.g., 0.025% for reference-grade)
MACHINE COMPARISON
Machine | Price ($) | Hrs/Cal | Cells | Swaps/Run | CMC @ High (%) | U @ High | TUR @ High | CMC @ Low (%) | U @ Low | TUR @ Low | PFR Glob @ Worst | PFR M6 @ Worst | PFR Spec @ Worst | RW Glob | RW M6 | RW Spec | Cal $/yr | Swap $/yr | Total Glob | Total M6 | Total Spec | What-If Tol (lbf) | TUR @ What-If Hi | TUR @ What-If Lo | RW M6 @ What-If | Total M6 @ What-If | CMC Model a | CMC Model b | Req CMC (%) | Swap Force (lbf) | Auto Cells | Auto Swaps
BCM | 25000 | 0.9 | 2 | 1 | 0.0005 | 2.5 | 2 | 0.06% | 0.3 | 1.6667 | 6.682% | 15% | 44.35% | 15591.7939 | 34994.2468 | 103480.7942 | 4000 | 200000 | 219591.7939 | 238994.2468 | 307480.7942 | 1.25 | 0.5 | 0.41667 | N/A — TUR<1 | Not Viable | 0.055556 | 0.00048889 | 0.025% | 5000 | 2 | 1
PCM | 35000 | 0.9 | 2 | 1 | 0.00033 | 1.65 | 3.0303 | 0.03% | 0.15 | 3.3333 | 3.062% | 4.472% | 18.63% | 7145.5244 | 10434.0816 | 43461.0209 | 4000 | 200000 | 211145.5244 | 214434.0816 | 247461.0209 | 1.25 | 0.75758 | 0.83333 | N/A — TUR<1 | Not Viable | -0.016667 | 0.00033333 | 0.025% | 5000 | 2 | 1
UCM Manual | 50000 | 1 | 2 | 1 | 0.0003 | 1.5 | 3.3333 | 0.03% | 0.15 | 3.3333 | 2.727% | 3.675% | 16.3% | 6361.9159 | 8575.4868 | 38037.0336 | 4000 | 200000 | 210361.9159 | 212575.4868 | 242037.0336 | 1.25 | 0.83333 | 0.83333 | N/A — TUR<1 | Not Viable | 0 | 0.0003 | 0.025% | 5000 | 2 | 1
UCM Automated | 100000 | 0.9 | 1 | 0 | 0.0002 | 1 | 5 | 0.025% | 0.125 | 4 | 2.197% | 2.513% | 12.71% | 5125.7 | 5864.4992 | 29662.6447 | 2000 | 0 | 7125.7 | 7864.4992 | 31662.6447 | 1.25 | 1.25 | 1 | Very High | See Calculator | 0.027778 | 0.00019444 | 0.025% | 500 | 1 | 0
Deadweight | 300000 | 1.2 | 0 | 0 | 2e-05 | 0.1 | 50 | 0.002% | 0.01 | 50 | 0.4758% | 0.4758% | 0.9104% | 1110.1212 | 1110.1212 | 2124.3245 | 0 | 0 | 1110.1212 | 1110.1212 | 2124.3245 | 1.25 | 12.5 | 12.5 | Low | See Calculator | 0 | 0 | 0 | 0 | 0 | 0
Competitor | 55000 | 1 | 2 | 1 | 0.0005 | 2.5 | 2 | 0.12% | 0.6 | 0.83333 | 17.65% | 54.02% | 89% | 41194.5924 | 126041.9068 | 207666.6667 | 4000 | 200000 | 245194.5924 | 330041.9068 | 411666.6667 | 1.25 | 0.5 | 0.20833 | N/A — TUR<1 | Not Viable | 0.38889 | 0.00042222 | 0.025% | 5000 | 2 | 1
Competitor 2 | 40000 | 1 | 2 | 1 | 0.0006 | 3 | 1.6667 | 0.15% | 0.75 | 0.66667 | 23.65% | 74.22% | 89% | 55188.3616 | 173185.9757 | 207666.6667 | 4000 | 200000 | 259188.3616 | 377185.9757 | 411666.6667 | 1.25 | 0.41667 | 0.16667 | N/A — TUR<1 | Not Viable | 0.5 | 0.0005 | 0.025% | 5000 | 2 | 1
Note: CMC@High (col F) and Tolerance@High (0.100%) are wired to Dashboard. CMC@Low (col I), Hrs/Cal (col C), and Price (col B) are independent — set per machine.
CHART DATA
Machine | Global |  | Method 6 | Specific Risk (PFA = 2.5 %) | Simple Accept
BCM | 219591.7939 |  | 238994.2468 | 307480.7942 | 216176.3971
PCM (N/A) | 211145.5244 |  | 214434.0816 | 247461.0209 | 0
UCM Manual | 210361.9159 |  | 212575.4868 | 242037.0336 | 210361.9159
UCM Automated | 7125.7 |  | 7864.4992 | 31662.6447 | 5977.1373
Deadweight | 1110.1212 |  | 1110.1212 | 2124.3245 | 1110.1212
Competitor | 245194.5924 |  | 330041.9068 | 411666.6667 | 216176.3971
Competitor 2 | 259188.3616 |  | 377185.9757 | 411666.6667 | 218528.1178
Notes:
PFR computed at both force points (5,000 and 500 lbf). Rework uses worst case (highest PFR). Blue cells are editable.
Swap cost = Swaps × (30 min / 60) × $200/hr × 2,000 cals/yr. Reset swaps to 0 if not switching standards.
CMC SENSITIVITY TABLE
CMC SENSITIVITY — Annual Cost vs. CMC at Low Force Point
Shows how total annual cost changes as CMC improves. Uses low force point (500 lbf, worst case). Cal cost = $2,000/cell.
CMC (%) | U (lbf) |  | TUR | PFR Global | PFR M6 | PFR Specific | Rework Global | Rework M6 | Rework Specific | Total Global* | Total M6* | Total Specific* | Savings vs BCM (M6) | Savings per 0.01% CMC
0.1% | 0.5 |  | 1 | 13.68% | 39.81% | 87.03% | 31918.4462 | 92883.6355 | 203066.5625 | 33918.4462 | 94883.6355 | 205066.5625 | 0 | —
0.08% | 0.4 |  | 1.25 | 9.963% | 26.36% | 66.21% | 23247.3222 | 61512.0252 | 154484.5257 | 25247.3222 | 63512.0252 | 156484.5257 | 0 | 15685.8052
0.06% | 0.3 |  | 1.6667 | 6.645% | 14.97% | 44.35% | 15504.3394 | 34936.8613 | 103472.1686 | 17504.3394 | 36936.8613 | 105472.1686 | 0 | 13287.5819
0.05% | 0.25 |  | 2 | 5.177% | 10.37% | 33.98% | 12079.9992 | 24195.8134 | 79279.5223 | 14079.9992 | 26195.8134 | 81279.5223 | 0 | 10741.048
0.04% | 0.2 |  | 2.5 | 3.851% | 6.583% | 24.52% | 8986.1548 | 15359.3299 | 57216.289 | 10986.1548 | 17359.3299 | 59216.289 | 0 | 8836.4835
0.035% | 0.175 |  | 2.8571 | 3.243% | 5% | 20.24% | 7565.9193 | 11666.4035 | 47225.9345 | 9565.9193 | 13666.4035 | 49225.9345 | 3692.9264 | 7385.8528
0.03% | 0.15 |  | 3.3333 | 2.67% | 3.623% | 16.3% | 6229.6033 | 8453.8607 | 38022.0803 | 8229.6033 | 10453.8607 | 40022.0803 | 6905.4692 | 6425.0856
0.025% | 0.125 |  | 4 | 2.132% | 2.451% | 12.71% | 4975.3149 | 5718.4109 | 29645.2282 | 6975.3149 | 7718.4109 | 31645.2282 | 9640.919 | 5470.8996
0.02% | 0.1 |  | 5 | 1.628% | 1.628% | 9.476% | 3799.4476 | 3799.4476 | 22111.5528 | 5799.4476 | 5799.4476 | 24111.5528 | 11559.8823 | 3837.9266
0.015% | 0.075 |  | 6.6667 | 1.155% | 1.155% | 6.606% | 2694.991 | 2694.991 | 15413.7526 | 4694.991 | 4694.991 | 17413.7526 | 12664.3388 | 2208.9131
0.01% | 0.05 |  | 10 | 0.7042% | 0.7042% | 4.082% | 1643.1837 | 1643.1837 | 9523.6006 | 3643.1837 | 3643.1837 | 11523.6006 | 13716.1462 | 2103.6148
0.005% | 0.025 |  | 20 | 0.2184% | 0.2184% | 1.939% | 509.5298 | 509.5298 | 4524.2667 | 2509.5298 | 2509.5298 | 6524.2667 | 14849.8001 | 2267.3077
0.003% | 0.015 |  | 33.3333 | 0.01703% | 0.01703% | 1.005% | 39.7411 | 39.7411 | 2345.8124 | 2039.7411 | 2039.7411 | 4345.8124 | 15319.5888 | 2348.9437
0.002% | 0.01 |  | 50 | 0% | 0% | 0.2018% | 0 | 0 | 470.8952 | 2000 | 2000 | 2470.8952 | 15359.3299 | 397.4105
* Total = Rework + $2,000 cal (1 cell). Swap cost excluded to isolate measurement capability impact.
Savings vs BCM shows cumulative annual savings from upgrading beyond 0.04% CMC. When Savings per 0.01% CMC flattens, further improvement has diminishing returns.
CELL SIZING CALCULATOR
How many reference cells does each machine need to cover the full range at the Target TUR? Based on CMC curve interpolation from CMC@High and CMC@Low.
Machine | CMC@Hi |  | CMC@Lo | Req CMC | Swap Force | Cells | Swaps | Cal $/yr | Swap $/yr | Total $/yr
BCM | 0.05% |  | 0.06% | 0.025% | N/V | 2 | 1 | 4000 | 200000 | 204000
PCM | 0.033% |  | 0.03% | 0.025% | N/V | 2 | 1 | 4000 | 200000 | 204000
UCM Manual | 0.03% |  | 0.03% | 0.025% | N/V | 2 | 1 | 4000 | 200000 | 204000
UCM Automated | 0.02% |  | 0.025% | 0.025% | 500 | 1 | 0 | 2000 | 0 | 2000
Deadweight | 0.002% |  | 0.002% | 0% | 0 | 0 | 0 | 0 | 0 | 0
Competitor | 0.05% |  | 0.12% | 0.025% | N/V | 2 | 1 | 4000 | 200000 | 204000
Competitor 2 | 0.06% |  | 0.15% | 0.025% | N/V | 2 | 1 | 4000 | 200000 | 204000
Target TUR (B13) controls how far down in a cell's range the CMC remains adequate. Lower TUR targets need fewer cells but accept more rework. The CMC model estimates CMC at any force from the two data points (CMC@High and CMC@Low) using a 1/F + offset curve. "N/V" means the machine cannot achieve the target TUR even at its best point.
When "Standard Switching Required?" = Yes on Dashboard, swaps are calculated. When No (matched capacity), cells are still needed but swaps = 0 — you buy matched reference cells for each sub-range.
EOPR SENSITIVITY — How Recommended CMC Changes with EOPR
Shows the optimal CMC at the low force point for different EOPR assumptions.
EOPR | σ_product |  | Req CMC for TUR 4:1 | M6 Rework @ TUR 3 | M6 Rework @ TUR 4 | M6 Rework @ TUR 5
70% | 0.48242 |  | 5e-05% | 12051.4009 | 6883.8403 | 4705.6922
75% | 0.43465 |  | 5e-05% | 12145.66 | 6880.7109 | 4682.0701
80% | 0.39015 |  | 5e-05% | 11965.359 | 6705.5765 | 4535.6766
85% | 0.34734 |  | 5e-05% | 11401.3492 | 6293.945 | 4222.1433
90% | 0.30398 |  | 5e-05% | 10248.3232 | 5527.4963 | 3661.3071
95% | 0.25511 |  | 5e-05% | 8000.9444 | 4125.8049 | 2667.501
At lower EOPR, more items are near the tolerance edge → PFR increases → rework goes up → the optimal CMC shifts lower (need better equipment).
REFERENCE CELL COVERAGE CALCULATOR
How many reference cells do you need to cover your range at your target TUR?
INPUTS
High Force (lbf) | 5000
Low Force (lbf) | 500
Target TUR | 4
Min Loading (%) | 5%
Device Tolerance (%) | 0.1%
RESULTS
Required CMC at Low Force (%) | 0.025%
Number of Cells Needed | 1
Swaps per Run | 0
Annual Swap Cost | 0
CELL CHAIN DETAIL
Cell # | Capacity (lbf) |  | Covers Down To | TUR at Bottom | Covers Low?
1 | 5000 |  | 500 | 4 | ✓ Yes
2
3
4
5
6
HOW THIS WORKS
Cell 1 starts at High Force and covers down to Capacity × Min Loading %.
If that doesn't reach Low Force, Cell 2 picks up where Cell 1 left off.
Each additional cell = one swap per calibration run.
At 5%: 10,000 → 500 → 25 = 2 cells, 1 swap. At 10%: 3 cells. At 20%: 4+ cells.
Better load cells (lower min loading %) = fewer cells = fewer swaps = lower cost.
EOPR SENSITIVITY — How Recommended CMC Changes with EOPR
Shows M6 rework at the low force point for different EOPR assumptions.
EOPR | σ_product |  | Req CMC for TUR 4:1 | M6 Rework @ TUR 3 | M6 Rework @ TUR 4 | M6 Rework @ TUR 5
70% | 0.48242 |  | 5e-05% | 12051.4009 | 6883.8403 | 4705.6922
75% | 0.43465 |  | 5e-05% | 12145.66 | 6880.7109 | 4682.0701
80% | 0.39015 |  | 5e-05% | 11965.359 | 6705.5765 | 4535.6766
85% | 0.34734 |  | 5e-05% | 11401.3492 | 6293.945 | 4222.1433
90% | 0.30398 |  | 5e-05% | 10248.3232 | 5527.4963 | 3661.3071
95% | 0.25511 |  | 5e-05% | 8000.9444 | 4125.8049 | 2667.501
At lower EOPR, more items are near the tolerance edge → PFR increases → you need better equipment to achieve the same cost.
Note on C24: henry zumbrun: Avg hours per calibration for BCM. Include setup, loading, data collection, documentation.
Note on C25: henry zumbrun: Avg hours per calibration for PCM.
Note on C26: henry zumbrun: Avg hours per calibration for UCM Manual.
Note on C27: henry zumbrun: Avg hours per calibration for UCM Automated.
Note on C28: henry zumbrun: Avg hours per calibration for Deadweight.
Note on C29: henry zumbrun: Avg hours per calibration for Competitor 1.
Note on C30: henry zumbrun: Avg hours per calibration for Competitor 2.

## Breakeven
Breakeven Analysis — Machine Investment Payback
Prices from Range Analysis. Savings = difference in total annual cost. Toggle swap cost on Dashboard.
MACHINE ANNUAL COST SUMMARY
Machine | Price | Total Glob/yr | Total M6/yr | Total Spec/yr
BCM | 25000 | 219591.7939 | 238994.2468 | 307480.7942
PCM | 35000 | 211145.5244 | 214434.0816 | 247461.0209
UCM Manual | 50000 | 210361.9159 | 212575.4868 | 242037.0336
UCM Automated | 100000 | 7125.7 | 7864.4992 | 31662.6447
Deadweight | 300000 | 1110.1212 | 1110.1212 | 2124.3245
Competitor | 55000 | 245194.5924 | 330041.9068 | 411666.6667
Competitor 2 | 40000 | 259188.3616 | 377185.9757 | 411666.6667
UPGRADE DELTAS (auto-calculated from prices above)
Path | Price Delta | M6 Savings/yr | Payback (M6, yrs)
Competitor 1 → BCM | -30000 | 91047.66 | -0.3295
Competitor 1 → UCM | -5000 | 117466.42 | -0.042565
Competitor 2 → UCM | 10000 | 164610.4889 | 0.060749
BCM → UCM Manual | 25000 | 26418.76 | 0.9463
BCM → UCM Auto | 75000 | 231129.7477 | 0.32449
UCM Manual → Deadweight | 250000 | 211465.3656 | 1.1822
PROFITABILITY MODEL
Price per Calibration ($) | 500
Target Payback (years) | 3
Calibrations per Year | 2000
Hourly Labor Rate ($/hr) | 200
PROFITABILITY PER MACHINE
Machine | Price | Revenue/yr | Cost Glob | Cost M6 | Cost Spec | Profit Glob | Profit M6 | Profit Spec | Payback Glob | Payback M6 | Payback Spec | Req Price (Glob) | Req Price (M6) | Hrs/Cal | Labor Cost/yr | Req Price (Spec) | Req $/Hr (M6) | Cost Simple | Profit Simple | Payback Simple
BCM | 25000 | 1000000 | 219591.7939 | 238994.2468 | 307480.7942 | 420408.2061 | 401005.7532 | 332519.2058 | 0.059466 | 0.062343 | 0.075184 | 293.9626 | 303.6638 | 0.9 | 360000 | 337.9071 | 337.4042 | 216176.3971 | 423823.6029 | 0.058987
PCM | 35000 | 1000000 | 211145.5244 | 214434.0816 | 247461.0209 | 428854.4756 | 425565.9184 | 392538.9791 | 0.081613 | 0.082243 | 0.089163 | 291.4061 | 293.0504 | 0.9 | 360000 | 309.5638 | 325.6115 | 211145.5244 | 428854.4756 | 0.081613
UCM Manual | 50000 | 1000000 | 210361.9159 | 212575.4868 | 242037.0336 | 389638.0841 | 387424.5132 | 357962.9664 | 0.12832 | 0.12906 | 0.13968 | 313.5143 | 314.6211 | 1 | 400000 | 329.3519 | 314.6211 | 210361.9159 | 389638.0841 | 0.12832
UCM Automated | 100000 | 1000000 | 7125.7 | 7864.4992 | 31662.6447 | 632874.3 | 632135.5008 | 608337.3553 | 0.15801 | 0.15819 | 0.16438 | 200.2295 | 200.5989 | 0.9 | 360000 | 212.498 | 222.8877 | 5977.1373 | 634022.8627 | 0.15772
Deadweight | 300000 | 1000000 | 1110.1212 | 1110.1212 | 2124.3245 | 518889.8788 | 518889.8788 | 517875.6755 | 0.57816 | 0.57816 | 0.57929 | 290.5551 | 290.5551 | 1.2 | 480000 | 291.0622 | 242.1292 | 1110.1212 | 518889.8788 | 0.57816
Competitor | 55000 | 1000000 | 245194.5924 | 330041.9068 | 411666.6667 | 354805.4076 | 269958.0932 | 188333.3333 | 0.15501 | 0.20374 | 0.29204 | 331.764 | 374.1876 | 1 | 400000 | 415 | 374.1876 | 216176.3971 | 383823.6029 | 0.14329
Competitor 2 | 40000 | 1000000 | 259188.3616 | 377185.9757 | 411666.6667 | 340811.6384 | 222814.0243 | 188333.3333 | 0.11737 | 0.17952 | 0.21239 | 336.2608 | 395.2597 | 1 | 400000 | 412.5 | 395.2597 | 218528.1178 | 381471.8822 | 0.10486
Hrs/Cal pulled from Range Analysis tab. Edit there to change.
REQUIRED CALIBRATIONS FOR TARGET PAYBACK
Machine | Price | Req Cals (Glob) | Req Cals (M6) | Req Cals (Spec)
BCM | 25000 | 118.932 | 124.6865 | 150.3673
PCM | 35000 | 163.2255 | 164.4869 | 178.3262
UCM Manual | 50000 | 256.6484 | 258.1148 | 279.3585
UCM Automated | 100000 | 316.0185 | 316.3879 | 328.7649
Deadweight | 300000 | 1156.3147 | 1156.3147 | 1158.5792
Competitor | 55000 | 310.0291 | 407.4707 | 584.0708
Competitor 2 | 40000 | 234.7338 | 359.0438 | 424.7788
CHART DATA — ANNUAL PROFIT
Machine | Profit (Global) | Profit (M6) | Profit (Specific) | Profit (Simple)
BCM | 420408.2061 | 401005.7532 | 332519.2058 | 423823.6029
PCM | 0 | 0 | 0 | 0
UCM Manual | 389638.0841 | 387424.5132 | 357962.9664 | 389638.0841
UCM Automated | 632874.3 | 632135.5008 | 608337.3553 | 634022.8627
Deadweight | 518889.8788 | 518889.8788 | 517875.6755 | 518889.8788
Competitor | 354805.4076 | 269958.0932 | 188333.3333 | 383823.6029
Competitor 2 | 340811.6384 | 222814.0243 | 188333.3333 | 381471.8822
Machine | Payback (Global) | Payback (M6) | Payback (Specific) | Payback (Simple)
BCM | 0.059466 | 0.062343 | 0.075184 | 0.058987
PCM | 0 | 0 | 0 | 0
UCM Manual | 0.12832 | 0.12906 | 0.13968 | 0.12832
UCM Automated | 0.15801 | 0.15819 | 0.16438 | 0.15772
Deadweight | 0.57816 | 0.57816 | 0.57929 | 0.57816
Competitor | 0.15501 | 0.20374 | 0.29204 | 0.14329
Competitor 2 | 0.11737 | 0.17952 | 0.21239 | 0.10486
PROFITABILITY vs TUR (Scatter Data)
Machine | TUR | Profit Global | TUR | Profit M6 | TUR | Profit Specific | TUR | Profit Simple
BCM | 2 | 420408.2061 | 2 | 401005.7532 | 2 | 332519.2058 | 2 | 423823.6029
PCM | 3.0303 | 428854.4756 | 3.0303 | 425565.9184 | 3.0303 | 392538.9791 | 3.0303 | 428854.4756
UCM Manual | 3.3333 | 389638.0841 | 3.3333 | 387424.5132 | 3.3333 | 357962.9664 | 3.3333 | 389638.0841
UCM Automated | 5 | 632874.3 | 5 | 632135.5008 | 5 | 608337.3553 | 5 | 634022.8627
Deadweight | 50 | 518889.8788 | 50 | 518889.8788 | 50 | 517875.6755 | 50 | 518889.8788
Competitor | 2 | 354805.4076 | 2 | 269958.0932 | 2 | 188333.3333 | 2 | 383823.6029
Competitor 2 | 1.6667 | 340811.6384 | 1.6667 | 222814.0243 | 1.6667 | 188333.3333 | 1.6667 | 381471.8822
$0 Line | 0 | 0
| 60 | 0
Each dot is a machine at its TUR. Points above the red $0 line are profitable; below are losses. The vertical spread at each TUR shows the cost of guard banding — Simple Accept (no guard band) is always the highest profit but carries the most customer risk.
Notes:
Revenue = Calibrations/yr × Price per Calibration.
Labor Cost = Hrs/Cal × Hourly Rate × Cals/yr. Edit Hrs/Cal per machine (green column) to reflect your actual calibration time.
Operating Cost = (Rework + Swap + Ref Cal) from Range Analysis. Does not include consumables or facility overhead.
Profit = Revenue − Labor − Operating Cost. Negative profit means the machine loses money at that price per cal.
Payback = Machine Price / Annual Profit. "Never" means the machine never pays for itself at current pricing.
Automation reduces Hrs/Cal directly, which is the single biggest lever on profitability (bigger than rework savings).

## Guard Band Analysis
Guard Band Tax Calculator
Shows tolerance consumed at various TUR levels
TUR | U (lbf) | GB Spec | Tol% Spec | Usable Spec | GB M6 | Tol% M6 | Usable M6 | GB Simple | Tol% Simple | Usable Simple
1.5 | 3.3333 | 3.2667 | 65.33% | 1.7333 | 1.2006 | 24.01% | 3.7994 | 0 | 0% | 5
2 | 2.5 | 2.45 | 49% | 2.55 | 0.70411 | 14.08% | 4.2959 | 0 | 0% | 5
2.5 | 2 | 1.96 | 39.2% | 3.04 | 0.42907 | 8.581% | 4.5709 | 0 | 0% | 5
3 | 1.6667 | 1.6333 | 32.67% | 3.3667 | 0.25886 | 5.177% | 4.7411 | 0 | 0% | 5
3.5 | 1.4286 | 1.4 | 28% | 3.6 | 0.14564 | 2.913% | 4.8544 | 0 | 0% | 5
4 | 1.25 | 1.225 | 24.5% | 3.775 | 0.066402 | 1.328% | 4.9336 | 0 | 0% | 5
4.6 | 1.087 | 1.0652 | 21.3% | 3.9348 | 0 | 0% | 5 | 0 | 0% | 5
5 | 1 | 0.98 | 19.6% | 4.02 | 0 | 0% | 5 | 0 | 0% | 5
10 | 0.5 | 0.49 | 9.8% | 4.51 | 0 | 0% | 5 | 0 | 0% | 5
24 | 0.20833 | 0.20417 | 4.083% | 4.7958 | 0 | 0% | 5 | 0 | 0% | 5

## Standard Switching
Standard Switching Cost Analysis
Risk from Sandia bivariate integral
INPUTS
Swap Time (min) | 30
Cal Cost/Cell ($/yr) | 2000
Downtime (weeks) | 2.5
Min Load% | #Cells | Swaps | SwapCost/Cal | AnnSwapCost | AnnCalCost | DT(wks) | TUR @0.1% | PFR Spec | RW Spec | PFR M6 | RW M6 | Total Spec | Total M6 | Ship Events | PFR Glob | RW Glob | Total Glob
5% | 1 | 0 | 0 | 0 | 2000 | 2.5 | 3 | 18.89% | 44074.4498 | 4.565% | 10651.333 | 46074.4498 | 12651.333 | 1 | 3.101% | 7234.5304 | 9234.5304
10% | 2 | 1 | 100 | 200000 | 4000 | 2.5 | 4 | 12.71% | 29653.2213 | 2.513% | 5864.4992 | 233653.2213 | 209864.4992 | 2 | 2.197% | 5125.7 | 209125.7
20% | 3 | 2 | 200 | 400000 | 6000 | 2.5 | 5 | 9.48% | 22121.1057 | 1.704% | 3977.1373 | 428121.1057 | 409977.1373 | 3 | 1.704% | 3977.1373 | 409977.1373
Deadweight | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Notes:
PFR via Sandia bivariate integral. M6 always ≤ Specific at same TUR.
Shipping Risk = 1 event/cell/yr in proper Morehouse case with custom-cut foam.
CHART DATA
Strategy | Swap Cost | Cal Cost | Rework (M6) | Total (M6)
5% (UPC) | 0 | 2000 | 10651.333 | 12651.333
10% (PC) | 200000 | 4000 | 5864.4992 | 209864.4992
20% (Commercial) | 400000 | 6000 | 3977.1373 | 409977.1373
Deadweight | 0 | 0 | 0 | 0

## Risk Calc
Sandia PSL Bivariate Normal Risk Calculator
Based on Sandia PSL PFA Risk Calculator (SAND2025-10736O). Uses ERF-based Drezner approximation.
SHARED INPUTS
Tolerance (± lbf) | 5
EOPR | 89%
σ_product | 3.1285
N (cals/yr) | 2000
Rework cost ($) | 116.6667
BVN Constants
cx | -1.095
cy | -0.75651
Machine |  | BCM | PCM | UCM Manual | UCM Automated | Deadweight | Interface Machine |  | 5% Load (TUR 3) | 10% Load (TUR 4) | 20% Load (TUR 5)
CMC (%) |  | 0.05% | 0.033% | 0.03% | 0.02% | 0.002% | 0.05% |  | 0 | 0 | 0
U (lbf) |  | 2.5 | 1.65 | 1.5 | 1 | 0.1 | 2.5 |  | 1.6667 | 1.25 | 1
σ_t |  | 1.25 | 0.825 | 0.75 | 0.5 | 0.05 | 1.25 |  | 0.83333 | 0.625 | 0.5
TUR |  | 2 | 3.0303 | 3.3333 | 5 | 50 | 2 |  | 3 | 4 | 5
σ_total |  | 3.369 | 3.2355 | 3.2172 | 3.1682 | 3.1289 | 3.369 |  | 3.2376 | 3.1904 | 3.1682
ρ |  | 0.92862 | 0.96694 | 0.97245 | 0.98747 | 0.99987 | 0.92862 |  | 0.96631 | 0.98062 | 0.98747
d = √(1-ρ²) |  | 0.37103 | 0.25499 | 0.23312 | 0.15782 | 0.01598 | 0.37103 |  | 0.25739 | 0.1959 | 0.15782
a_bvn = -ρ/d |  | -2.5028 | -3.7922 | -4.1714 | -6.2571 | -62.5707 | -2.5028 |  | -3.7542 | -5.0057 | -6.2571
GB Specific |  | 2.45 | 1.617 | 1.47 | 0.98 | 0.098 | 2.45 |  | 1.6333 | 1.225 | 0.98
Accept width (a_spec) |  | 2.55 | 3.383 | 3.53 | 4.02 | 4.902 | 2.55 |  | 3.3667 | 3.775 | 4.02
p1 = a/σ_t |  | 0.7569 | 1.0456 | 1.0972 | 1.2688 | 1.5667 | 0.7569 |  | 1.0399 | 1.1833 | 1.2688
p2 = -a/σ_t |  | -0.7569 | -1.0456 | -1.0972 | -1.2688 | -1.5667 | -0.7569 |  | -1.0399 | -1.1833 | -1.2688
q_neg = -t/σ_p |  | -1.5982 | -1.5982 | -1.5982 | -1.5982 | -1.5982 | -1.5982 |  | -1.5982 | -1.5982 | -1.5982
q_pos = t/σ_p |  | 1.5982 | 1.5982 | 1.5982 | 1.5982 | 1.5982 | 1.5982 |  | 1.5982 | 1.5982 | 1.5982
BVN1: b |  | 2.04 | 4.1006 | 4.7067 | 8.04 | 98.04 | 2.04 |  | 4.04 | 6.04 | 8.04
BVN1: aq+b |  | 6.04 | 10.1612 | 11.3733 | 18.04 | 198.04 | 6.04 |  | 10.04 | 14.04 | 18.04
BVN1: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN1: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN1: exp_arg1 |  | -0.22233 | -0.43983 | -0.48065 | -0.6186 | -0.85664 | -0.22233 |  | -0.43532 | -0.54948 | -0.6186
BVN1: exp_arg2 |  | 0.32814 | 0.094738 | 0.033952 | -0.21196 | -0.8054 | 0.32814 |  | 0.10112 | -0.080765 | -0.21196
BVN1: result
BVN2: b |  | -2.04 | -4.1006 | -4.7067 | -8.04 | -98.04 | -2.04 |  | -4.04 | -6.04 | -8.04
BVN2: aq+b |  | 1.96 | 1.96 | 1.96 | 1.96 | 1.96 | 1.96 |  | 1.96 | 1.96 | 1.96
BVN2: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN2: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN2: exp_arg1 |  | 0.32814 | 0.094738 | 0.033952 | -0.21196 | -0.8054 | 0.32814 |  | 0.10112 | -0.080765 | -0.21196
BVN2: exp_arg2 |  | -0.22233 | -0.43983 | -0.48065 | -0.6186 | -0.85664 | -0.22233 |  | -0.43532 | -0.54948 | -0.6186
BVN2: result
BVN3: b |  | -2.04 | -4.1006 | -4.7067 | -8.04 | -98.04 | -2.04 |  | -4.04 | -6.04 | -8.04
BVN3: aq+b |  | -6.04 | -10.1612 | -11.3733 | -18.04 | -198.04 | -6.04 |  | -10.04 | -14.04 | -18.04
BVN3: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN3: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN3: exp_arg1 |  | 0.32814 | 0.094738 | 0.033952 | -0.21196 | -0.8054 | 0.32814 |  | 0.10112 | -0.080765 | -0.21196
BVN3: exp_arg2 |  | -0.22233 | -0.43983 | -0.48065 | -0.6186 | -0.85664 | -0.22233 |  | -0.43532 | -0.54948 | -0.6186
BVN3: result
PFA Global, R_C (JCGM eq.19) |  | 0.06773% | 0.04798% | 0.04418% | 0.03074% | 0.003325% | 0.06773% |  | 0.0484% | 0.03761% | 0.03074%
PFR Global, R_P (JCGM eq.20) |  | 33.98% | 18.62% | 16.3% | 9.48% | 0.7225% | 33.98% |  | 18.89% | 12.71% | 9.48%
Specific pc / PFA at A_U (ILAC G8) |  | 2.5% | 2.5% | 2.5% | 2.5% | 2.5% | 2.5% |  | 2.5% | 2.5% | 2.5%
M6 Multiplier |  | 0.28165 | 0.15193 | 0.11918 | 0 | 0 | 0.28165 |  | 0.15532 | 0.053121 | 0
GB M6 |  | 0.70411 | 0.25069 | 0.17877 | 0 | 0 | 0.70411 |  | 0.25886 | 0.066402 | 0
Accept width (a_m6) |  | 4.2959 | 4.7493 | 4.8212 | 5 | 5 | 4.2959 |  | 4.7411 | 4.9336 | 5
p1_m6 |  | 1.2751 | 1.4679 | 1.4986 | 1.5782 | 1.598 | 1.2751 |  | 1.4644 | 1.5464 | 1.5782
p2_m6 |  | -1.2751 | -1.4679 | -1.4986 | -1.5782 | -1.598 | -1.2751 |  | -1.4644 | -1.5464 | -1.5782
BVN4: b |  | 3.4367 | 5.7567 | 6.4283 | 10 | 100 | 3.4367 |  | 5.6894 | 7.8938 | 10
BVN4: aq+b |  | 7.4367 | 11.8173 | 13.095 | 20 | 200 | 7.4367 |  | 11.6894 | 15.8938 | 20
BVN4: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN4: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN4: exp_arg1 |  | -0.91495 | -1.0676 | -1.0867 | -1.105 | -0.90671 | -0.91495 |  | -1.0653 | -1.111 | -1.105
BVN4: exp_arg2 |  | 0.012403 | -0.31714 | -0.3839 | -0.59921 | -0.85444 | 0.012403 |  | -0.30985 | -0.49844 | -0.59921
BVN4: result
BVN5: b |  | -3.4367 | -5.7567 | -6.4283 | -10 | -100 | -3.4367 |  | -5.6894 | -7.8938 | -10
BVN5: aq+b |  | 0.56329 | 0.30387 | 0.23836 | 0 | 0 | 0.56329 |  | 0.31064 | 0.10624 | 0
BVN5: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN5: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN5: exp_arg1 |  | 0.012403 | -0.31714 | -0.3839 | -0.59921 | -0.85444 | 0.012403 |  | -0.30985 | -0.49844 | -0.59921
BVN5: exp_arg2 |  | -0.91495 | -1.0676 | -1.0867 | -1.105 | -0.90671 | -0.91495 |  | -1.0653 | -1.111 | -1.105
BVN5: result
BVN6: b |  | -3.4367 | -5.7567 | -6.4283 | -10 | -100 | -3.4367 |  | -5.6894 | -7.8938 | -10
BVN6: aq+b |  | -7.4367 | -11.8173 | -13.095 | -20 | -200 | -7.4367 |  | -11.6894 | -15.8938 | -20
BVN6: z |  | 5.7389 | 11.879 | 14.1636 | 30.6181 | 2962.8083 | 5.7389 |  | 11.6625 | 19.9556 | 30.6181
BVN6: √z |  | 2.3956 | 3.4466 | 3.7635 | 5.5334 | 54.4317 | 2.3956 |  | 3.415 | 4.4672 | 5.5334
BVN6: exp_arg1 |  | 0.012403 | -0.31714 | -0.3839 | -0.59921 | -0.85444 | 0.012403 |  | -0.30985 | -0.49844 | -0.59921
BVN6: exp_arg2 |  | -0.91495 | -1.0676 | -1.0867 | -1.105 | -0.90671 | -0.91495 |  | -1.0653 | -1.111 | -1.105
BVN6: result
PFA (Method 6) |  | 1.135% | 1.199% | 1.212% | 1.155% | 4.028e-05% | 1.135% |  | 1.197% | 1.234% | 1.155%
PFR (Method 6) |  | 10.4% | 4.472% | 3.675% | 1.704% | 0.4758% | 10.4% |  | 4.565% | 2.513% | 1.704%
FR/yr (Specific) |  | 679.5786 | 372.4595 | 325.9628 | 189.6095 | 14.4493 | 679.5786 |  | 377.781 | 254.1705 | 189.6095
FR/yr (M6) |  | 207.9942 | 89.435 | 73.5042 | 34.0897 | 9.5153 | 207.9942 |  | 91.2971 | 50.2671 | 34.0897
Rework $/yr (Spec) |  | 79284.1758 | 43453.6065 | 38028.9887 | 22121.1057 | 1685.7521 | 79284.1758 |  | 44074.4498 | 29653.2213 | 22121.1057
Rework $/yr (M6) |  | 24265.9842 | 10434.0816 | 8575.4868 | 3977.1373 | 1110.1212 | 24265.9842 |  | 10651.333 | 5864.4992 | 3977.1373
SIMPLE ACCEPTANCE (Shared Risk, GB = 0)
PFR (Simple) |  | 5.218% | 3.062% | 2.727% | 1.704% | 0.4758% | 5.218%
FR/yr (Simple) |  | 104.3691 | 61.2474 | 54.5307 | 34.0897 | 9.5153 | 104.3691
Rework $/yr (Simple) |  | 12176.3971 | 7145.5244 | 6361.9159 | 3977.1373 | 1110.1212 | 12176.3971
NUMERICAL INTEGRATION (Trapezoidal Rule)
Integration points | 40 | [11 empty] | BCM Spec | BCM M6 | PCM Spec | PCM M6 | UCM Spec | UCM M6 | UCMAuto Spec | UCMAuto M6 | DW Spec | DW M6 | Interface Spec | Interface M6 | TUR3 Spec | TUR3 M6 | TUR4 Spec | TUR4 M6 | TUR5 Spec | TUR5 M6 | Global C | Global D | Global E | Global F | Global G | Global H | Global J | Global K | Global L
-5 | 0.25641 | [11 empty] | 0.00088885 | 0.010191 | 0.00088885 | 0.013534 | 0.00088885 | 0.014429 | 0.00088885 | 0.017779 | 0.00088885 | 0.017779 | 0.00088885 | 0.010191 | 0.00088885 | 0.013442 | 0.00088885 | 0.016274 | 0.00088885 | 0.017779 | 0.015642 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.015642 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779
-4.7436 | [12 empty] | 0.0016014 | 0.014548 | 0.0020018 | 0.020311 | 0.0021337 | 0.021864 | 0.0029863 | 0.028115 | 0.040367 | 0.040398 | 0.0016014 | 0.014548 | 0.001989 | 0.020151 | 0.0024482 | 0.025024 | 0.0029863 | 0.028115 | 0.021067 | 0.025128 | 0.025603 | 0.028115 | 0.040398 | 0.021067 | 0.02508 | 0.02663 | 0.028115 | 0.023482 | 0.025128 | 0.025603 | 0.028115 | 0.040398 | 0.023482
-4.4872 | [12 empty] | 0.0027628 | 0.020022 | 0.0041205 | 0.028478 | 0.0046016 | 0.030636 | 0.0079809 | 0.038636 | 0.04559 | 0.04559 | 0.0027628 | 0.020022 | 0.0040746 | 0.028253 | 0.0058013 | 0.034761 | 0.0079809 | 0.038636 | 0.027454 | 0.033412 | 0.034326 | 0.038636 | 0.04559 | 0.027454 | 0.033319 | 0.0362 | 0.038636 | 0.030052 | 0.033412 | 0.034326 | 0.038636 | 0.04559 | 0.030052
-4.2308 | [12 empty] | 0.0045674 | 0.026614 | 0.0077714 | 0.037571 | 0.0089463 | 0.040088 | 0.017206 | 0.047938 | 0.051104 | 0.051104 | 0.0045674 | 0.026614 | 0.0076598 | 0.0373 | 0.011904 | 0.044441 | 0.017206 | 0.047938 | 0.034687 | 0.042132 | 0.043309 | 0.047938 | 0.051104 | 0.034687 | 0.042009 | 0.045524 | 0.047938 | 0.03735 | 0.042132 | 0.043309 | 0.047938 | 0.051104 | 0.03735
-3.9744 | [12 empty] | 0.0072409 | 0.034227 | 0.013472 | 0.047014 | 0.015749 | 0.049539 | 0.030521 | 0.055758 | 0.056903 | 0.056903 | 0.0072409 | 0.034227 | 0.013254 | 0.046731 | 0.021331 | 0.053351 | 0.030521 | 0.055758 | 0.042583 | 0.05082 | 0.052025 | 0.055758 | 0.056903 | 0.042583 | 0.050689 | 0.054035 | 0.055758 | 0.045183 | 0.05082 | 0.052025 | 0.055758 | 0.056903 | 0.045183
-3.7179 | [12 empty] | 0.011017 | 0.042675 | 0.021547 | 0.056288 | 0.025241 | 0.058489 | 0.045761 | 0.06261 | 0.062935 | 0.062935 | 0.011017 | 0.042675 | 0.021189 | 0.056028 | 0.033756 | 0.061306 | 0.045761 | 0.06261 | 0.050918 | 0.059153 | 0.060186 | 0.06261 | 0.062935 | 0.050918 | 0.059035 | 0.061669 | 0.06261 | 0.053336 | 0.059153 | 0.060186 | 0.06261 | 0.062935 | 0.053336
-3.4615 | [12 empty] | 0.016105 | 0.051701 | 0.031949 | 0.065043 | 0.037085 | 0.066726 | 0.060013 | 0.069069 | 0.069141 | 0.069141 | 0.016105 | 0.051701 | 0.031437 | 0.064831 | 0.047846 | 0.068501 | 0.060013 | 0.069069 | 0.059449 | 0.06699 | 0.06775 | 0.069069 | 0.069141 | 0.059449 | 0.066898 | 0.068663 | 0.069069 | 0.06159 | 0.06699 | 0.06775 | 0.069069 | 0.069141 | 0.06159
-3.2051 | [12 empty] | 0.022643 | 0.061006 | 0.044165 | 0.07314 | 0.050367 | 0.074274 | 0.071558 | 0.075438 | 0.07545 | 0.07545 | 0.022643 | 0.061006 | 0.043523 | 0.072987 | 0.061798 | 0.075236 | 0.071558 | 0.075438 | 0.067946 | 0.074334 | 0.07482 | 0.075438 | 0.07545 | 0.067946 | 0.074271 | 0.075296 | 0.075438 | 0.069752 | 0.074334 | 0.07482 | 0.075438 | 0.07545 | 0.069752
-2.9487 | [12 empty] | 0.030658 | 0.070287 | 0.057305 | 0.080595 | 0.06386 | 0.081271 | 0.080469 | 0.081782 | 0.081784 | 0.081784 | 0.030658 | 0.070287 | 0.056595 | 0.080496 | 0.074172 | 0.081722 | 0.080469 | 0.081782 | 0.076211 | 0.081256 | 0.081529 | 0.081782 | 0.081784 | 0.076211 | 0.081218 | 0.081741 | 0.081782 | 0.077662 | 0.081256 | 0.081529 | 0.081782 | 0.081784 | 0.077662
-2.6923 | [12 empty] | 0.040036 | 0.07927 | 0.070335 | 0.087498 | 0.076431 | 0.087856 | 0.087707 | 0.088055 | 0.088055 | 0.088055 | 0.040036 | 0.07927 | 0.069635 | 0.087441 | 0.084391 | 0.088041 | 0.087707 | 0.088055 | 0.084089 | 0.087828 | 0.087963 | 0.088055 | 0.088055 | 0.084089 | 0.087808 | 0.088046 | 0.088055 | 0.085199 | 0.087828 | 0.087963 | 0.088055 | 0.088055 | 0.085199
-2.4359 | [12 empty] | 0.050508 | 0.087734 | 0.082356 | 0.093936 | 0.087364 | 0.094104 | 0.094101 | 0.094173 | 0.094173 | 0.094173 | 0.050508 | 0.087734 | 0.081741 | 0.093906 | 0.09266 | 0.09417 | 0.094101 | 0.094173 | 0.09147 | 0.094085 | 0.094144 | 0.094173 | 0.094173 | 0.09147 | 0.094075 | 0.094172 | 0.094173 | 0.092279 | 0.094085 | 0.094144 | 0.094173 | 0.094173 | 0.092279
-2.1795 | [12 empty] | 0.061673 | 0.095519 | 0.092808 | 0.09995 | 0.096453 | 0.10002 | 0.10003 | 0.10004 | 0.10004 | 0.10004 | 0.061673 | 0.095519 | 0.092326 | 0.099937 | 0.099508 | 0.10004 | 0.10003 | 0.10004 | 0.098278 | 0.10001 | 0.10003 | 0.10004 | 0.10004 | 0.098278 | 0.10001 | 0.10004 | 0.10004 | 0.09884 | 0.10001 | 0.10003 | 0.10004 | 0.10004 | 0.09884
-1.9231 | [12 empty] | 0.073033 | 0.10252 | 0.10151 | 0.10553 | 0.10387 | 0.10556 | 0.10556 | 0.10557 | 0.10557 | 0.10557 | 0.073033 | 0.10252 | 0.10117 | 0.10553 | 0.1054 | 0.10557 | 0.10556 | 0.10557 | 0.10446 | 0.10556 | 0.10556 | 0.10557 | 0.10557 | 0.10446 | 0.10555 | 0.10557 | 0.10557 | 0.10484 | 0.10556 | 0.10556 | 0.10557 | 0.10557 | 0.10484
-1.6667 | [12 empty] | 0.084064 | 0.10869 | 0.10857 | 0.11064 | 0.10993 | 0.11065 | 0.11065 | 0.11065 | 0.11065 | 0.11065 | 0.084064 | 0.10869 | 0.10836 | 0.11064 | 0.11061 | 0.11065 | 0.11065 | 0.11065 | 0.10999 | 0.11064 | 0.11065 | 0.11065 | 0.11065 | 0.10999 | 0.11064 | 0.11065 | 0.11065 | 0.11022 | 0.11064 | 0.11065 | 0.11065 | 0.11065 | 0.11022
-1.4103 | [12 empty] | 0.094266 | 0.11399 | 0.11423 | 0.1152 | 0.11493 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.094266 | 0.11399 | 0.11411 | 0.11519 | 0.11519 | 0.1152 | 0.1152 | 0.1152 | 0.11482 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.11482 | 0.1152 | 0.1152 | 0.1152 | 0.11496 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.11496
-1.1538 | [12 empty] | 0.10322 | 0.11842 | 0.11872 | 0.11913 | 0.11904 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.10322 | 0.11842 | 0.11866 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11893 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11893 | 0.11913 | 0.11913 | 0.11913 | 0.11901 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11901
-0.89744 | [12 empty] | 0.11063 | 0.12197 | 0.12222 | 0.12238 | 0.12235 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.11063 | 0.12197 | 0.12219 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12227 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12227 | 0.12238 | 0.12238 | 0.12238 | 0.12231 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12231
-0.64103 | [12 empty] | 0.11629 | 0.12465 | 0.12481 | 0.12487 | 0.12486 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.11629 | 0.12465 | 0.1248 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12481 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12481 | 0.12487 | 0.12487 | 0.12487 | 0.12484 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12484
-0.38462 | [12 empty] | 0.1201 | 0.12643 | 0.12654 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.1201 | 0.12643 | 0.12654 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12653 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12653 | 0.12656 | 0.12656 | 0.12656 | 0.12654 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12654
-0.12821 | [12 empty] | 0.12201 | 0.12733 | 0.1274 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12201 | 0.12733 | 0.1274 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12739 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12739 | 0.12741 | 0.12741 | 0.12741 | 0.1274 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.1274
0.12821 | [12 empty] | 0.12201 | 0.12733 | 0.1274 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12201 | 0.12733 | 0.1274 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12739 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12739 | 0.12741 | 0.12741 | 0.12741 | 0.1274 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.1274
0.38462 | [12 empty] | 0.1201 | 0.12643 | 0.12654 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.1201 | 0.12643 | 0.12654 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12653 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12653 | 0.12656 | 0.12656 | 0.12656 | 0.12654 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12654
0.64103 | [12 empty] | 0.11629 | 0.12465 | 0.12481 | 0.12487 | 0.12486 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.11629 | 0.12465 | 0.1248 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12481 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12481 | 0.12487 | 0.12487 | 0.12487 | 0.12484 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12484
0.89744 | [12 empty] | 0.11063 | 0.12197 | 0.12222 | 0.12238 | 0.12235 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.11063 | 0.12197 | 0.12219 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12227 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12227 | 0.12238 | 0.12238 | 0.12238 | 0.12231 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12231
1.1538 | [12 empty] | 0.10322 | 0.11842 | 0.11872 | 0.11913 | 0.11904 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.10322 | 0.11842 | 0.11866 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11893 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11893 | 0.11913 | 0.11913 | 0.11913 | 0.11901 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11901
1.4103 | [12 empty] | 0.094266 | 0.11399 | 0.11423 | 0.1152 | 0.11493 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.094266 | 0.11399 | 0.11411 | 0.11519 | 0.11519 | 0.1152 | 0.1152 | 0.1152 | 0.11482 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.11482 | 0.1152 | 0.1152 | 0.1152 | 0.11496 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.11496
1.6667 | [12 empty] | 0.084064 | 0.10869 | 0.10857 | 0.11064 | 0.10993 | 0.11065 | 0.11065 | 0.11065 | 0.11065 | 0.11065 | 0.084064 | 0.10869 | 0.10836 | 0.11064 | 0.11061 | 0.11065 | 0.11065 | 0.11065 | 0.10999 | 0.11064 | 0.11065 | 0.11065 | 0.11065 | 0.10999 | 0.11064 | 0.11065 | 0.11065 | 0.11022 | 0.11064 | 0.11065 | 0.11065 | 0.11065 | 0.11022
1.9231 | [12 empty] | 0.073033 | 0.10252 | 0.10151 | 0.10553 | 0.10387 | 0.10556 | 0.10556 | 0.10557 | 0.10557 | 0.10557 | 0.073033 | 0.10252 | 0.10117 | 0.10553 | 0.1054 | 0.10557 | 0.10556 | 0.10557 | 0.10446 | 0.10556 | 0.10556 | 0.10557 | 0.10557 | 0.10446 | 0.10555 | 0.10557 | 0.10557 | 0.10484 | 0.10556 | 0.10556 | 0.10557 | 0.10557 | 0.10484
2.1795 | [12 empty] | 0.061673 | 0.095519 | 0.092808 | 0.09995 | 0.096453 | 0.10002 | 0.10003 | 0.10004 | 0.10004 | 0.10004 | 0.061673 | 0.095519 | 0.092326 | 0.099937 | 0.099508 | 0.10004 | 0.10003 | 0.10004 | 0.098278 | 0.10001 | 0.10003 | 0.10004 | 0.10004 | 0.098278 | 0.10001 | 0.10004 | 0.10004 | 0.09884 | 0.10001 | 0.10003 | 0.10004 | 0.10004 | 0.09884
2.4359 | [12 empty] | 0.050508 | 0.087734 | 0.082356 | 0.093936 | 0.087364 | 0.094104 | 0.094101 | 0.094173 | 0.094173 | 0.094173 | 0.050508 | 0.087734 | 0.081741 | 0.093906 | 0.09266 | 0.09417 | 0.094101 | 0.094173 | 0.09147 | 0.094085 | 0.094144 | 0.094173 | 0.094173 | 0.09147 | 0.094075 | 0.094172 | 0.094173 | 0.092279 | 0.094085 | 0.094144 | 0.094173 | 0.094173 | 0.092279
2.6923 | [12 empty] | 0.040036 | 0.07927 | 0.070335 | 0.087498 | 0.076431 | 0.087856 | 0.087707 | 0.088055 | 0.088055 | 0.088055 | 0.040036 | 0.07927 | 0.069635 | 0.087441 | 0.084391 | 0.088041 | 0.087707 | 0.088055 | 0.084089 | 0.087828 | 0.087963 | 0.088055 | 0.088055 | 0.084089 | 0.087808 | 0.088046 | 0.088055 | 0.085199 | 0.087828 | 0.087963 | 0.088055 | 0.088055 | 0.085199
2.9487 | [12 empty] | 0.030658 | 0.070287 | 0.057305 | 0.080595 | 0.06386 | 0.081271 | 0.080469 | 0.081782 | 0.081784 | 0.081784 | 0.030658 | 0.070287 | 0.056595 | 0.080496 | 0.074172 | 0.081722 | 0.080469 | 0.081782 | 0.076211 | 0.081256 | 0.081529 | 0.081782 | 0.081784 | 0.076211 | 0.081218 | 0.081741 | 0.081782 | 0.077662 | 0.081256 | 0.081529 | 0.081782 | 0.081784 | 0.077662
3.2051 | [12 empty] | 0.022643 | 0.061006 | 0.044165 | 0.07314 | 0.050367 | 0.074274 | 0.071558 | 0.075438 | 0.07545 | 0.07545 | 0.022643 | 0.061006 | 0.043523 | 0.072987 | 0.061798 | 0.075236 | 0.071558 | 0.075438 | 0.067946 | 0.074334 | 0.07482 | 0.075438 | 0.07545 | 0.067946 | 0.074271 | 0.075296 | 0.075438 | 0.069752 | 0.074334 | 0.07482 | 0.075438 | 0.07545 | 0.069752
3.4615 | [12 empty] | 0.016105 | 0.051701 | 0.031949 | 0.065043 | 0.037085 | 0.066726 | 0.060013 | 0.069069 | 0.069141 | 0.069141 | 0.016105 | 0.051701 | 0.031437 | 0.064831 | 0.047846 | 0.068501 | 0.060013 | 0.069069 | 0.059449 | 0.06699 | 0.06775 | 0.069069 | 0.069141 | 0.059449 | 0.066898 | 0.068663 | 0.069069 | 0.06159 | 0.06699 | 0.06775 | 0.069069 | 0.069141 | 0.06159
3.7179 | [12 empty] | 0.011017 | 0.042675 | 0.021547 | 0.056288 | 0.025241 | 0.058489 | 0.045761 | 0.06261 | 0.062935 | 0.062935 | 0.011017 | 0.042675 | 0.021189 | 0.056028 | 0.033756 | 0.061306 | 0.045761 | 0.06261 | 0.050918 | 0.059153 | 0.060186 | 0.06261 | 0.062935 | 0.050918 | 0.059035 | 0.061669 | 0.06261 | 0.053336 | 0.059153 | 0.060186 | 0.06261 | 0.062935 | 0.053336
3.9744 | [12 empty] | 0.0072409 | 0.034227 | 0.013472 | 0.047014 | 0.015749 | 0.049539 | 0.030521 | 0.055758 | 0.056903 | 0.056903 | 0.0072409 | 0.034227 | 0.013254 | 0.046731 | 0.021331 | 0.053351 | 0.030521 | 0.055758 | 0.042583 | 0.05082 | 0.052025 | 0.055758 | 0.056903 | 0.042583 | 0.050689 | 0.054035 | 0.055758 | 0.045183 | 0.05082 | 0.052025 | 0.055758 | 0.056903 | 0.045183
4.2308 | [12 empty] | 0.0045674 | 0.026614 | 0.0077714 | 0.037571 | 0.0089463 | 0.040088 | 0.017206 | 0.047938 | 0.051104 | 0.051104 | 0.0045674 | 0.026614 | 0.0076598 | 0.0373 | 0.011904 | 0.044441 | 0.017206 | 0.047938 | 0.034687 | 0.042132 | 0.043309 | 0.047938 | 0.051104 | 0.034687 | 0.042009 | 0.045524 | 0.047938 | 0.03735 | 0.042132 | 0.043309 | 0.047938 | 0.051104 | 0.03735
4.4872 | [12 empty] | 0.0027628 | 0.020022 | 0.0041205 | 0.028478 | 0.0046016 | 0.030636 | 0.0079809 | 0.038636 | 0.04559 | 0.04559 | 0.0027628 | 0.020022 | 0.0040746 | 0.028253 | 0.0058013 | 0.034761 | 0.0079809 | 0.038636 | 0.027454 | 0.033412 | 0.034326 | 0.038636 | 0.04559 | 0.027454 | 0.033319 | 0.0362 | 0.038636 | 0.030052 | 0.033412 | 0.034326 | 0.038636 | 0.04559 | 0.030052
4.7436 | [12 empty] | 0.0016014 | 0.014548 | 0.0020018 | 0.020311 | 0.0021337 | 0.021864 | 0.0029863 | 0.028115 | 0.040367 | 0.040398 | 0.0016014 | 0.014548 | 0.001989 | 0.020151 | 0.0024482 | 0.025024 | 0.0029863 | 0.028115 | 0.021067 | 0.025128 | 0.025603 | 0.028115 | 0.040398 | 0.021067 | 0.02508 | 0.02663 | 0.028115 | 0.023482 | 0.025128 | 0.025603 | 0.028115 | 0.040398 | 0.023482
5 | [12 empty] | 0.00088885 | 0.010191 | 0.00088885 | 0.013534 | 0.00088885 | 0.014429 | 0.00088885 | 0.017779 | 0.00088885 | 0.017779 | 0.00088885 | 0.010191 | 0.00088885 | 0.013442 | 0.00088885 | 0.016274 | 0.00088885 | 0.017779 | 0.015642 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.015642 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779
[13 empty] | 0.55019 | 0.786 | 0.70374 | 0.84528 | 0.72698 | 0.85325 | 0.79515 | 0.87296 | 0.8809 | 0.88524 | 0.55019 | 0.786 | 0.70108 | 0.84435 | 0.76287 | 0.86487 | 0.79515 | 0.87296 | 0.82625 | 0.85938 | 0.86273 | 0.87296 | 0.88524 | 0.82625 | 0.85899 | 0.86803 | 0.87296 | 0.83782 | 0.85938 | 0.86273 | 0.87296 | 0.88524 | 0.83782
2%
0.41714
| 5.2086 | [11 empty] | 0.00053312 | 0.0074198 | 0.00042913 | 0.0092129 | 0.00040209 | 0.0096561 | 0.00027822 | 0.010789 | 1.3893e-11 | 4.828e-07 | 0.00053312 | 0.0074198 | 0.00043191 | 0.0091668 | 0.00034774 | 0.010524 | 0.00027822 | 0.010789
| 5.6257 | [11 empty] | 0.0001756 | 0.0036381 | 8.3032e-05 | 0.003647 | 6.5846e-05 | 0.003588 | 1.672e-05 | 0.0026683 | 2.2348e-49 | 7.9102e-38 | 0.0001756 | 0.0036381 | 8.4955e-05 | 0.0036517 | 3.8801e-05 | 0.0033942 | 1.672e-05 | 0.0026683
| 6.0428 | [11 empty] | 5.1351e-05 | 0.0016017 | 1.2477e-05 | 0.001154 | 7.9645e-06 | 0.0010203 | 5.1502e-07 | 0.00036533 | 3.0773e-117 | 1.303e-98 | 5.1351e-05 | 0.0016017 | 1.3039e-05 | 0.0011676 | 2.8137e-06 | 0.0007496 | 5.1502e-07 | 0.00036533
| 6.46 | [11 empty] | 1.3312e-05 | 0.00063078 | 1.4501e-06 | 0.00028832 | 7.0781e-07 | 0.00021849 | 8.0248e-09 | 2.6476e-05 | 2.8399e-215 | 1.4821e-189 | 1.3312e-05 | 0.00063078 | 1.5553e-06 | 0.00029609 | 1.3155e-07 | 0.0001104 | 8.0248e-09 | 2.6476e-05
| 6.8771 | [11 empty] | 3.0554e-06 | 0.00022156 | 1.2992e-07 | 5.6375e-05 | 4.6041e-08 | 3.4846e-05 | 6.2724e-11 | 9.8968e-07 | 0 | 0 | 3.0554e-06 | 0.00022156 | 1.4373e-07 | 5.9037e-05 | 3.9431e-09 | 1.0662e-05 | 6.2724e-11 | 9.8968e-07
| 7.2943 | [11 empty] | 6.2034e-07 | 6.9244e-05 | 8.9536e-09 | 8.5726e-06 | 2.1858e-09 | 4.1073e-06 | 2.4456e-13 | 1.8787e-08 | 0 | 0 | 6.2034e-07 | 6.9244e-05 | 1.0267e-08 | 9.1991e-06 | 7.5484e-11 | 6.6786e-07 | 2.4456e-13 | 1.8787e-08
| 7.7114 | [11 empty] | 1.1131e-07 | 1.9219e-05 | 4.7374e-10 | 1.0092e-06 | 7.5579e-11 | 3.5585e-07 | 4.7381e-16 | 1.7935e-10 | 0 | 0 | 1.1131e-07 | 1.9219e-05 | 5.6591e-10 | 1.1151e-06 | 9.2025e-13 | 2.6932e-08 | 4.7381e-16 | 1.7935e-10
| 8.1285 | [11 empty] | 1.7641e-08 | 4.7298e-06 | 1.9217e-11 | 9.1669e-08 | 1.9002e-12 | 2.2573e-08 | 4.5484e-19 | 8.5559e-13 | 0 | 0 | 1.7641e-08 | 4.7298e-06 | 2.4034e-11 | 1.0482e-07 | 7.1295e-15 | 6.9563e-10 | 4.5484e-19 | 8.5559e-13
| 8.5457 | [11 empty] | 2.468e-09 | 1.0309e-06 | 5.9698e-13 | 6.4081e-09 | 3.4692e-14 | 1.0453e-09 | 2.159e-22 | 2.0305e-15 | 0 | 0 | 2.468e-09 | 1.0309e-06 | 7.8563e-13 | 7.6202e-09 | 3.5044e-17 | 1.1467e-11 | 2.159e-22 | 2.0305e-15
| 8.9628 | [11 empty] | 3.0462e-10 | 1.9877e-07 | 1.4189e-14 | 3.4408e-10 | 4.5948e-16 | 3.5263e-11 | 5.0596e-26 | 2.3899e-18 | 0 | 0 | 3.0462e-10 | 1.9877e-07 | 1.9748e-14 | 4.2767e-10 | 1.0915e-19 | 1.2031e-13 | 5.0596e-26 | 2.3899e-18
| 9.3799 | [11 empty] | 3.316e-11 | 3.3878e-08 | 2.5786e-16 | 1.417e-11 | 4.4112e-18 | 8.6506e-13 | 5.8466e-30 | 1.3918e-21 | 0 | 0 | 3.316e-11 | 3.3878e-08 | 3.8146e-16 | 1.8501e-11 | 2.152e-22 | 8.0187e-16 | 5.8466e-30 | 1.3918e-21
| 9.7971 | [11 empty] | 3.1823e-12 | 5.1004e-09 | 3.5805e-18 | 4.47e-13 | 3.0676e-20 | 1.5412e-14 | 3.3282e-34 | 4.0036e-25 | 0 | 0 | 3.1823e-12 | 5.1004e-09 | 5.6585e-18 | 6.1622e-13 | 2.6837e-25 | 3.3897e-18 | 3.3282e-34 | 4.0036e-25
| 10.2142 | [11 empty] | 2.6917e-13 | 6.779e-10 | 3.797e-20 | 1.0792e-14 | 1.5444e-22 | 1.9922e-16 | 9.3258e-39 | 5.681e-29 | 0 | 0 | 2.6917e-13 | 6.779e-10 | 6.4431e-20 | 1.5787e-14 | 2.1155e-28 | 9.0777e-21 | 9.3258e-39 | 5.681e-29
| 10.6314 | 0.06773% | 0.04798% | 0.04418% | 0.03074% | 0.003325% | 0.06773% |  | 0.04435% | 0.03249% | 0.02465% |  | 2.006e-14 | 7.95e-11 | 3.0739e-22 | 1.9923e-16 | 5.6263e-25 | 1.8667e-18 | 1.2855e-43 | 3.9724e-33 | 0 | 0 | 2.006e-14 | 7.95e-11 | 5.629e-22 | 3.1085e-16 | 1.0535e-31 | 1.5386e-23 | 1.2855e-43 | 3.9724e-33
| 11.0485 | 1.135% | 1.199% | 1.212% | 1.155% | 4.028e-05% | 1.135% |  | 1.197% | 1.234% | 1.155% |  | 1.317e-15 | 8.223e-12 | 1.8989e-24 | 2.8109e-18 | 1.4826e-27 | 1.2671e-20 | 8.7122e-49 | 1.3676e-37 | 0 | 0 | 1.317e-15 | 8.223e-12 | 3.7719e-24 | 4.7013e-18 | 3.3131e-35 | 1.6492e-26 | 8.7122e-49 | 1.3676e-37
| 11.4656 | [11 empty] | 7.6151e-17 | 7.499e-13 | 8.9495e-27 | 3.0291e-20 | 2.8252e-30 | 6.2266e-23 | 2.9019e-54 | 2.3167e-42 | 0 | 0 | 7.6151e-17 | 7.499e-13 | 1.9379e-26 | 5.4585e-20 | 6.5769e-39 | 1.1172e-29 | 2.9019e-54 | 2.3167e-42
| 11.8828 | 2.389% | 1.77% | 1.64% | 1.155% | 4.028e-05% | 2.389% |  | 1.783% | 1.409% | 1.155% |  | 3.8774e-18 | 6.0277e-14 | 3.2168e-29 | 2.4921e-22 | 3.8916e-33 | 2.2143e-25 | 4.7486e-60 | 1.9298e-47 | 0 | 0 | 3.8774e-18 | 6.0277e-14 | 7.6326e-29 | 4.8633e-22 | 8.2387e-43 | 4.7811e-33 | 4.7486e-60 | 1.9298e-47
| 12.2999 | [11 empty] | 1.7382e-19 | 4.2692e-15 | 8.8165e-32 | 1.5648e-24 | 3.8743e-36 | 5.6962e-28 | 3.8164e-66 | 7.9018e-53 | 0 | 0 | 1.7382e-19 | 4.2692e-15 | 2.3038e-31 | 3.3237e-24 | 6.5108e-47 | 1.2919e-36 | 3.8164e-66 | 7.9018e-53
| 12.717 | [11 empty] | 6.8598e-21 | 2.6637e-16 | 1.8421e-34 | 7.4964e-27 | 2.7869e-39 | 1.0596e-30 | 1.506e-72 | 1.5897e-58 | 0 | 0 | 6.8598e-21 | 2.6637e-16 | 5.3284e-34 | 1.7418e-26 | 3.2452e-51 | 2.2034e-40 | 1.506e-72 | 1.5897e-58
| 1313% | [11 empty] | 2.3829e-22 | 1.4639e-17 | 2.9338e-37 | 2.7391e-29 | 1.4483e-42 | 1.4249e-33 | 2.9172e-79 | 1.5709e-64 | 0 | 0 | 2.3829e-22 | 1.4639e-17 | 9.4411e-37 | 6.9978e-29 | 1.02e-55 | 2.3713e-44 | 2.9172e-79 | 1.5709e-64
| 13.5513 | [11 empty] | 7.2851e-24 | 7.0843e-19 | 3.5608e-40 | 7.632e-32 | 5.4364e-46 | 1.3849e-36 | 2.7735e-86 | 7.6228e-71 | 0 | 0 | 7.2851e-24 | 7.0843e-19 | 1.2814e-39 | 2.1548e-31 | 2.0212e-60 | 1.6098e-48 | 2.7735e-86 | 7.6228e-71
| 13.9685 | [11 empty] | 1.96e-25 | 3.0186e-20 | 3.2933e-43 | 1.6212e-34 | 1.4738e-49 | 9.7257e-40 | 1.2939e-93 | 1.8159e-77 | 0 | 0 | 1.96e-25 | 3.0186e-20 | 1.3319e-42 | 5.0842e-34 | 2.5247e-65 | 6.8927e-53 | 1.2939e-93 | 1.8159e-77
| 14.3856 | [11 empty] | 4.6403e-27 | 1.1323e-21 | 2.3206e-46 | 2.6251e-37 | 2.8851e-53 | 4.9344e-43 | 2.9616e-101 | 2.1232e-84 | 0 | 0 | 4.6403e-27 | 1.1323e-21 | 1.0602e-45 | 9.1908e-37 | 1.9877e-70 | 1.8609e-57 | 2.9616e-101 | 2.1232e-84
| 14.8027 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 9.6663e-29 | 3.7387e-23 | 1.2458e-49 | 3.2395e-40 | 4.078e-57 | 1.8083e-46 | 3.3255e-109 | 1.2183e-91 | 0 | 0 | 9.6663e-29 | 3.7387e-23 | 6.4621e-49 | 1.2727e-39 | 9.8621e-76 | 3.1674e-62 | 3.3255e-109 | 1.2183e-91
| 15.2199 | 2.45 | 1.617 | 1.47 | 0.98 | 0.098 | 2.45 |  | 1.6333 | 1.225 | 0.98 |  | 1.7716e-30 | 1.0865e-24 | 5.0943e-53 | 3.0464e-43 | 4.1614e-61 | 4.7862e-50 | 1.8316e-117 | 3.4299e-99 | 0 | 0 | 1.7716e-30 | 1.0865e-24 | 3.0156e-52 | 1.3498e-42 | 3.0834e-81 | 3.3983e-67 | 1.8316e-117 | 3.4299e-99
| 15.637 | 1.225 | 0.8085 | 0.735 | 0.49 | 0.049 | 1.225 |  | 0.81667 | 0.6125 | 0.49 |  | 2.8564e-32 | 2.7786e-26 | 1.5867e-56 | 2.1828e-46 | 3.0656e-65 | 9.1478e-54 | 4.9476e-126 | 4.7372e-107 | 0 | 0 | 2.8564e-32 | 2.7786e-26 | 1.0774e-55 | 1.0963e-45 | 6.0739e-87 | 2.298e-72 | 4.9476e-126 | 4.7372e-107
| 16.0541 | 0.0056682 | 0.0040111 | 0.0036778 | 0.0024496 | 3.44e-09 | 0.0056682 |  | 0.0040472 | 0.0030888 | 0.0024496 |  | 4.0514e-34 | 6.2532e-28 | 3.7641e-60 | 1.1915e-49 | 1.6302e-69 | 1.2624e-57 | 6.5543e-135 | 3.2095e-115 | 0 | 0 | 4.0514e-34 | 6.2532e-28 | 2.9464e-59 | 6.8185e-49 | 7.5382e-93 | 9.7928e-78 | 6.5543e-135 | 3.2095e-115
| 16.4713 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 5.0548e-36 | 1.2382e-29 | 6.8004e-64 | 4.9547e-53 | 6.2568e-74 | 1.2578e-61 | 4.2577e-144 | 1.0665e-123 | 0 | 0 | 5.0548e-36 | 1.2382e-29 | 6.168e-63 | 3.2469e-52 | 5.8936e-99 | 2.6296e-83 | 4.2577e-144 | 1.0665e-123
| 16.8884 | 1.225 | 0.8085 | 0.735 | 0.49 | 0.049 | 1.225 |  | 0.81667 | 0.6125 | 0.49 |  | 5.5475e-38 | 2.1572e-31 | 9.3559e-68 | 1.5693e-56 | 1.7332e-78 | 9.0469e-66 | 1.3562e-153 | 1.738e-132 | 0 | 0 | 5.5475e-38 | 2.1572e-31 | 9.8828e-67 | 1.1837e-55 | 2.9025e-105 | 4.4488e-89 | 1.3562e-153 | 1.738e-132
| 17.3056 | 0.6125 | 0.40425 | 0.3675 | 0.245 | 0.0245 | 0.6125 |  | 0.40833 | 0.30625 | 0.245 |  | 5.3549e-40 | 3.3064e-33 | 9.8015e-72 | 3.7858e-60 | 3.4649e-83 | 4.6971e-70 | 2.1179e-163 | 1.389e-141 | 0 | 0 | 5.3549e-40 | 3.3064e-33 | 1.212e-70 | 3.3035e-59 | 9.0039e-112 | 4.7418e-95 | 2.1179e-163 | 1.389e-141
|  | 0.01266 | 0.009161 | 0.0084467 | 0.0057954 | 4.1787e-08 | 0.01266 |  | 0.0092384 | 0.0071779 | 0.0057954
|  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0
|  | 0.6125 | 0.40425 | 0.3675 | 0.245 | 0.0245 | 0.6125 |  | 0.40833 | 0.30625 | 0.245
|  | 0.30625 | 0.20212 | 0.18375 | 0.1225 | 0.01225 | 0.30625 |  | 0.20417 | 0.15313 | 0.1225
|  | 0.017741 | 0.012987 | 0.012007 | 0.0083505 | 1.3353e-07 | 0.017741 |  | 0.013093 | 0.01026 | 0.0083505
|  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0
|  | 0.30625 | 0.20212 | 0.18375 | 0.1225 | 0.01225 | 0.30625 |  | 0.20417 | 0.15313 | 0.1225
|  | 0.15313 | 0.10106 | 0.091875 | 0.06125 | 0.006125 | 0.15313 |  | 0.10208 | 0.076563 | 0.06125
|  | 0.020686 | 0.015233 | 0.014102 | 0.0098711 | 2.3358e-07 | 0.020686 |  | 0.015355 | 0.012083 | 0.0098711
|  | 0.15313 | 0 | 0 | 0 | 0 | 0.15313 |  | 0 | 0 | 0
|  | 0.30625 | 0.10106 | 0.091875 | 0.06125 | 0.006125 | 0.30625 |  | 0.10208 | 0.076563 | 0.06125
|  | 0.22969 | 0.050531 | 0.045937 | 0.030625 | 0.0030625 | 0.22969 |  | 0.051042 | 0.038281 | 0.030625
|  | 0.019181 | 0.016437 | 0.015228 | 0.010693 | 3.0728e-07 | 0.019181 |  | 0.016568 | 0.013064 | 0.010693
|  | 0.15313 | 0 | 0 | 0 | 0 | 0.15313 |  | 0 | 0 | 0
|  | 0.22969 | 0.050531 | 0.045937 | 0.030625 | 0.0030625 | 0.22969 |  | 0.051042 | 0.038281 | 0.030625
|  | 0.19141 | 0.025266 | 0.022969 | 0.015312 | 0.0015313 | 0.19141 |  | 0.025521 | 0.019141 | 0.015312
|  | 0.019926 | 0.01706 | 0.015809 | 0.011119 | 3.5197e-07 | 0.019926 |  | 0.017195 | 0.013572 | 0.011119
|  | 0.15313 | 0 | 0 | 0 | 0 | 0.15313 |  | 0 | 0 | 0
|  | 0.19141 | 0.025266 | 0.022969 | 0.015312 | 0.0015313 | 0.19141 |  | 0.025521 | 0.019141 | 0.015312
|  | 0.17227 | 0.012633 | 0.011484 | 0.0076562 | 0.00076563 | 0.17227 |  | 0.01276 | 0.0095703 | 0.0076562
|  | 0.020304 | 0.017376 | 0.016105 | 0.011335 | 3.7656e-07 | 0.020304 |  | 0.017513 | 0.013831 | 0.011335
|  | 0.17227 | 0 | 0 | 0 | 0 | 0.17227 |  | 0 | 0 | 0
|  | 0.19141 | 0.012633 | 0.011484 | 0.0076562 | 0.00076563 | 0.19141 |  | 0.01276 | 0.0095703 | 0.0076562
|  | 0.18184 | 0.0063164 | 0.0057422 | 0.0038281 | 0.00038281 | 0.18184 |  | 0.0063802 | 0.0047852 | 0.0038281
|  | 0.020114 | 0.017535 | 0.016254 | 0.011445 | 3.8946e-07 | 0.020114 |  | 0.017673 | 0.013961 | 0.011445
|  | 0.18184 | 0 | 0 | 0 | 0 | 0.18184 |  | 0 | 0 | 0
|  | 0.19141 | 0.0063164 | 0.0057422 | 0.0038281 | 0.00038281 | 0.19141 |  | 0.0063802 | 0.0047852 | 0.0038281
|  | 0.18662 | 0.0031582 | 0.0028711 | 0.0019141 | 0.00019141 | 0.18662 |  | 0.0031901 | 0.0023926 | 0.0019141
|  | 0.02002 | 0.017615 | 0.016329 | 0.0115 | 3.9607e-07 | 0.02002 |  | 0.017754 | 0.014026 | 0.0115
|  | 0.18662 | 0 | 0 | 0 | 0 | 0.18662 |  | 0 | 0 | 0
|  | 0.19141 | 0.0031582 | 0.0028711 | 0.0019141 | 0.00019141 | 0.19141 |  | 0.0031901 | 0.0023926 | 0.0019141
|  | 0.18901 | 0.0015791 | 0.0014355 | 0.00095703 | 9.5703e-05 | 0.18901 |  | 0.0015951 | 0.0011963 | 0.00095703
|  | 0.019973 | 0.017655 | 0.016366 | 0.011527 | 3.9941e-07 | 0.019973 |  | 0.017794 | 0.014059 | 0.011527
Solved GB (Global) |  | 0.18901 | 0 | 0 | 0 | 0 | 0.18901 |  | 0 | 0 | 0
Accept width (Global) |  | 4.811 | 5 | 5 | 5 | 5 | 4.811 |  | 5 | 5 | 5
PFA (Global) |  | 1.997% | 1.77% | 1.64% | 1.155% | 4.028e-05% | 1.997% |  | 1.783% | 1.409% | 1.155%
PFR (Global) |  | 6.375% | 3.062% | 2.727% | 1.704% | 0.4758% | 6.375% |  | 3.101% | 2.197% | 1.704%
FR/yr (Global) |  | 127.4998 | 61.2474 | 54.5307 | 34.0897 | 9.5153 | 127.4998 |  | 62.0103 | 43.9346 | 34.0897
Rework $/yr (Global) |  | 14874.9718 | 7145.5244 | 6361.9159 | 3977.1373 | 1110.1212 | 14874.9718 |  | 7234.5304 | 5125.7 | 3977.1373
RANGE ANALYSIS INTEGRATION
dx_hi | 0.25641
-5 |  | 0.00088885 | 0.00088885 | 0.00088885 | 0.00088885 | 0.00088885 | 0.00088885 | 0.00088881 | 0.010191 | 0.013534 | 0.014429 | 0.017779 | 0.017779 | 0.010191 | 0.008999 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779
-4.7436 |  | 0.0016014 | 0.0020018 | 0.0021337 | 0.0029863 | 0.040367 | 0.0016014 | 0.0014866 | 0.014548 | 0.020311 | 0.021864 | 0.028115 | 0.040398 | 0.014548 | 0.012551 | 0.023482 | 0.025128 | 0.025603 | 0.028115 | 0.040398 | 0.023482 | 0.02294
-4.4872 |  | 0.0027628 | 0.0041205 | 0.0046016 | 0.0079809 | 0.04559 | 0.0027628 | 0.0024077 | 0.020022 | 0.028478 | 0.030636 | 0.038636 | 0.04559 | 0.020022 | 0.017022 | 0.030052 | 0.033412 | 0.034326 | 0.038636 | 0.04559 | 0.030052 | 0.028894
-4.2308 |  | 0.0045674 | 0.0077714 | 0.0089463 | 0.017206 | 0.051104 | 0.0045674 | 0.0037771 | 0.026614 | 0.037571 | 0.040088 | 0.047938 | 0.051104 | 0.026614 | 0.022465 | 0.03735 | 0.042132 | 0.043309 | 0.047938 | 0.051104 | 0.03735 | 0.035567
-3.9744 |  | 0.0072409 | 0.013472 | 0.015749 | 0.030521 | 0.056903 | 0.0072409 | 0.0057419 | 0.034227 | 0.047014 | 0.049539 | 0.055758 | 0.056903 | 0.034227 | 0.028882 | 0.045183 | 0.05082 | 0.052025 | 0.055758 | 0.056903 | 0.045183 | 0.042844
-3.7179 |  | 0.011017 | 0.021547 | 0.025241 | 0.045761 | 0.062935 | 0.011017 | 0.008462 | 0.042675 | 0.056288 | 0.058489 | 0.06261 | 0.062935 | 0.042675 | 0.036207 | 0.053336 | 0.059153 | 0.060186 | 0.06261 | 0.062935 | 0.053336 | 0.050577
-3.4615 |  | 0.016105 | 0.031949 | 0.037085 | 0.060013 | 0.069141 | 0.016105 | 0.012096 | 0.051701 | 0.065043 | 0.066726 | 0.069069 | 0.069141 | 0.051701 | 0.044311 | 0.06159 | 0.06699 | 0.06775 | 0.069069 | 0.069141 | 0.06159 | 0.058595
-3.2051 |  | 0.022643 | 0.044165 | 0.050367 | 0.071558 | 0.07545 | 0.022643 | 0.016779 | 0.061006 | 0.07314 | 0.074274 | 0.075438 | 0.07545 | 0.061006 | 0.053008 | 0.069752 | 0.074334 | 0.07482 | 0.075438 | 0.07545 | 0.069752 | 0.066718
-2.9487 |  | 0.030658 | 0.057305 | 0.06386 | 0.080469 | 0.081784 | 0.030658 | 0.0226 | 0.070287 | 0.080595 | 0.081271 | 0.081782 | 0.081784 | 0.070287 | 0.062064 | 0.077662 | 0.081256 | 0.081529 | 0.081782 | 0.081784 | 0.077662 | 0.074772
-2.6923 |  | 0.040036 | 0.070335 | 0.076431 | 0.087707 | 0.088055 | 0.040036 | 0.029579 | 0.07927 | 0.087498 | 0.087856 | 0.088055 | 0.088055 | 0.07927 | 0.071223 | 0.085199 | 0.087828 | 0.087963 | 0.088055 | 0.088055 | 0.085199 | 0.082599
-2.4359 |  | 0.050508 | 0.082356 | 0.087364 | 0.094101 | 0.094173 | 0.050508 | 0.037641 | 0.087734 | 0.093936 | 0.094104 | 0.094173 | 0.094173 | 0.087734 | 0.080225 | 0.092279 | 0.094085 | 0.094144 | 0.094173 | 0.094173 | 0.092279 | 0.090059
-2.1795 |  | 0.061673 | 0.092808 | 0.096453 | 0.10003 | 0.10004 | 0.061673 | 0.04661 | 0.095519 | 0.09995 | 0.10002 | 0.10004 | 0.10004 | 0.095519 | 0.088827 | 0.09884 | 0.10001 | 0.10003 | 0.10004 | 0.10004 | 0.09884 | 0.097038
-1.9231 |  | 0.073033 | 0.10151 | 0.10387 | 0.10556 | 0.10557 | 0.073033 | 0.056204 | 0.10252 | 0.10553 | 0.10556 | 0.10557 | 0.10557 | 0.10252 | 0.09682 | 0.10484 | 0.10556 | 0.10556 | 0.10557 | 0.10557 | 0.10484 | 0.10344
-1.6667 |  | 0.084064 | 0.10857 | 0.10993 | 0.11065 | 0.11065 | 0.084064 | 0.06605 | 0.10869 | 0.11064 | 0.11065 | 0.11065 | 0.11065 | 0.10869 | 0.10403 | 0.11022 | 0.11064 | 0.11065 | 0.11065 | 0.11065 | 0.11022 | 0.10919
-1.4103 |  | 0.094266 | 0.11423 | 0.11493 | 0.1152 | 0.1152 | 0.094266 | 0.075709 | 0.11399 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.11399 | 0.11035 | 0.11496 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.11496 | 0.11424
-1.1538 |  | 0.10322 | 0.11872 | 0.11904 | 0.11913 | 0.11913 | 0.10322 | 0.084708 | 0.11842 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11842 | 0.11567 | 0.11901 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11901 | 0.11851
-0.89744 |  | 0.11063 | 0.12222 | 0.12235 | 0.12238 | 0.12238 | 0.11063 | 0.092581 | 0.12197 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12197 | 0.11996 | 0.12231 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12231 | 0.12199
-0.64103 |  | 0.11629 | 0.12481 | 0.12486 | 0.12487 | 0.12487 | 0.11629 | 0.098901 | 0.12465 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12465 | 0.12318 | 0.12484 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12484 | 0.12463
-0.38462 |  | 0.1201 | 0.12654 | 0.12656 | 0.12656 | 0.12656 | 0.1201 | 0.10332 | 0.12643 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12643 | 0.12534 | 0.12654 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12654 | 0.1264
-0.12821 |  | 0.12201 | 0.1274 | 0.12741 | 0.12741 | 0.12741 | 0.12201 | 0.1056 | 0.12733 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12733 | 0.12641 | 0.1274 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.1274 | 0.1273
0.12821 |  | 0.12201 | 0.1274 | 0.12741 | 0.12741 | 0.12741 | 0.12201 | 0.1056 | 0.12733 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.12733 | 0.12641 | 0.1274 | 0.12741 | 0.12741 | 0.12741 | 0.12741 | 0.1274 | 0.1273
0.38462 |  | 0.1201 | 0.12654 | 0.12656 | 0.12656 | 0.12656 | 0.1201 | 0.10332 | 0.12643 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12643 | 0.12534 | 0.12654 | 0.12656 | 0.12656 | 0.12656 | 0.12656 | 0.12654 | 0.1264
0.64103 |  | 0.11629 | 0.12481 | 0.12486 | 0.12487 | 0.12487 | 0.11629 | 0.098901 | 0.12465 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12465 | 0.12318 | 0.12484 | 0.12487 | 0.12487 | 0.12487 | 0.12487 | 0.12484 | 0.12463
0.89744 |  | 0.11063 | 0.12222 | 0.12235 | 0.12238 | 0.12238 | 0.11063 | 0.092581 | 0.12197 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12197 | 0.11996 | 0.12231 | 0.12238 | 0.12238 | 0.12238 | 0.12238 | 0.12231 | 0.12199
1.1538 |  | 0.10322 | 0.11872 | 0.11904 | 0.11913 | 0.11913 | 0.10322 | 0.084708 | 0.11842 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11842 | 0.11567 | 0.11901 | 0.11913 | 0.11913 | 0.11913 | 0.11913 | 0.11901 | 0.11851
1.4103 |  | 0.094266 | 0.11423 | 0.11493 | 0.1152 | 0.1152 | 0.094266 | 0.075709 | 0.11399 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.11399 | 0.11035 | 0.11496 | 0.1152 | 0.1152 | 0.1152 | 0.1152 | 0.11496 | 0.11424
1.6667 |  | 0.084064 | 0.10857 | 0.10993 | 0.11065 | 0.11065 | 0.084064 | 0.06605 | 0.10869 | 0.11064 | 0.11065 | 0.11065 | 0.11065 | 0.10869 | 0.10403 | 0.11022 | 0.11064 | 0.11065 | 0.11065 | 0.11065 | 0.11022 | 0.10919
1.9231 |  | 0.073033 | 0.10151 | 0.10387 | 0.10556 | 0.10557 | 0.073033 | 0.056204 | 0.10252 | 0.10553 | 0.10556 | 0.10557 | 0.10557 | 0.10252 | 0.09682 | 0.10484 | 0.10556 | 0.10556 | 0.10557 | 0.10557 | 0.10484 | 0.10344
2.1795 |  | 0.061673 | 0.092808 | 0.096453 | 0.10003 | 0.10004 | 0.061673 | 0.04661 | 0.095519 | 0.09995 | 0.10002 | 0.10004 | 0.10004 | 0.095519 | 0.088827 | 0.09884 | 0.10001 | 0.10003 | 0.10004 | 0.10004 | 0.09884 | 0.097038
2.4359 |  | 0.050508 | 0.082356 | 0.087364 | 0.094101 | 0.094173 | 0.050508 | 0.037641 | 0.087734 | 0.093936 | 0.094104 | 0.094173 | 0.094173 | 0.087734 | 0.080225 | 0.092279 | 0.094085 | 0.094144 | 0.094173 | 0.094173 | 0.092279 | 0.090059
2.6923 |  | 0.040036 | 0.070335 | 0.076431 | 0.087707 | 0.088055 | 0.040036 | 0.029579 | 0.07927 | 0.087498 | 0.087856 | 0.088055 | 0.088055 | 0.07927 | 0.071223 | 0.085199 | 0.087828 | 0.087963 | 0.088055 | 0.088055 | 0.085199 | 0.082599
2.9487 |  | 0.030658 | 0.057305 | 0.06386 | 0.080469 | 0.081784 | 0.030658 | 0.0226 | 0.070287 | 0.080595 | 0.081271 | 0.081782 | 0.081784 | 0.070287 | 0.062064 | 0.077662 | 0.081256 | 0.081529 | 0.081782 | 0.081784 | 0.077662 | 0.074772
3.2051 |  | 0.022643 | 0.044165 | 0.050367 | 0.071558 | 0.07545 | 0.022643 | 0.016779 | 0.061006 | 0.07314 | 0.074274 | 0.075438 | 0.07545 | 0.061006 | 0.053008 | 0.069752 | 0.074334 | 0.07482 | 0.075438 | 0.07545 | 0.069752 | 0.066718
3.4615 |  | 0.016105 | 0.031949 | 0.037085 | 0.060013 | 0.069141 | 0.016105 | 0.012096 | 0.051701 | 0.065043 | 0.066726 | 0.069069 | 0.069141 | 0.051701 | 0.044311 | 0.06159 | 0.06699 | 0.06775 | 0.069069 | 0.069141 | 0.06159 | 0.058595
3.7179 |  | 0.011017 | 0.021547 | 0.025241 | 0.045761 | 0.062935 | 0.011017 | 0.008462 | 0.042675 | 0.056288 | 0.058489 | 0.06261 | 0.062935 | 0.042675 | 0.036207 | 0.053336 | 0.059153 | 0.060186 | 0.06261 | 0.062935 | 0.053336 | 0.050577
3.9744 |  | 0.0072409 | 0.013472 | 0.015749 | 0.030521 | 0.056903 | 0.0072409 | 0.0057419 | 0.034227 | 0.047014 | 0.049539 | 0.055758 | 0.056903 | 0.034227 | 0.028882 | 0.045183 | 0.05082 | 0.052025 | 0.055758 | 0.056903 | 0.045183 | 0.042844
4.2308 |  | 0.0045674 | 0.0077714 | 0.0089463 | 0.017206 | 0.051104 | 0.0045674 | 0.0037771 | 0.026614 | 0.037571 | 0.040088 | 0.047938 | 0.051104 | 0.026614 | 0.022465 | 0.03735 | 0.042132 | 0.043309 | 0.047938 | 0.051104 | 0.03735 | 0.035567
4.4872 |  | 0.0027628 | 0.0041205 | 0.0046016 | 0.0079809 | 0.04559 | 0.0027628 | 0.0024077 | 0.020022 | 0.028478 | 0.030636 | 0.038636 | 0.04559 | 0.020022 | 0.017022 | 0.030052 | 0.033412 | 0.034326 | 0.038636 | 0.04559 | 0.030052 | 0.028894
4.7436 |  | 0.0016014 | 0.0020018 | 0.0021337 | 0.0029863 | 0.040367 | 0.0016014 | 0.0014866 | 0.014548 | 0.020311 | 0.021864 | 0.028115 | 0.040398 | 0.014548 | 0.012551 | 0.023482 | 0.025128 | 0.025603 | 0.028115 | 0.040398 | 0.023482 | 0.02294
5 |  | 0.00088885 | 0.00088885 | 0.00088885 | 0.00088885 | 0.00088885 | 0.00088885 | 0.00088881 | 0.010191 | 0.013534 | 0.014429 | 0.017779 | 0.017779 | 0.010191 | 0.008999 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779 | 0.017779
|  | 0.55019 | 0.70374 | 0.72698 | 0.79515 | 0.8809 | 0.55019 | 0.44651 | 0.786 | 0.84528 | 0.85325 | 0.87296 | 0.88524 | 0.786 | 0.74002 | 0.83782 | 0.85938 | 0.86273 | 0.87296 | 0.88524 | 0.83782 | 0.82318
|  | 33.98% | 18.63% | 16.3% | 9.485% | 0.9104% | 33.98% | 44.35% | 10.4% | 4.472% | 3.675% | 1.704% | 0.4758% | 10.4% | 15% | 5.218% | 3.062% | 2.727% | 1.704% | 0.4758% | 5.218% | 6.682%
dx_lo | 0.025641
-0.5 |  | 0.0088881 | 0.0088885 | 0.0088885 | 0.0088885 | 0.0088885 | -0.021273 | -0.076408 | 0.08999 | 0.14429 | 0.14429 | 0.16274 | 0.17779 | 0.053639 | 0.029678 | 0.17779 | 0.17779 | 0.17779 | 0.17779 | 0.17779 | 0.17763 | 0.17642
-0.47436 |  | 0.014866 | 0.021337 | 0.021337 | 0.024482 | 0.40367 | -0.027659 | -0.093872 | 0.12551 | 0.21864 | 0.21864 | 0.25024 | 0.40398 | 0.06873 | 0.036782 | 0.2294 | 0.25603 | 0.25603 | 0.2663 | 0.40398 | 0.21551 | 0.21111
-0.44872 |  | 0.024077 | 0.046016 | 0.046016 | 0.058013 | 0.4559 | -0.03547 | -0.11408 | 0.17022 | 0.30636 | 0.30636 | 0.34761 | 0.4559 | 0.086917 | 0.045074 | 0.28894 | 0.34326 | 0.34326 | 0.362 | 0.4559 | 0.25853 | 0.25014
-0.42308 |  | 0.037771 | 0.089463 | 0.089463 | 0.11904 | 0.51104 | -0.044862 | -0.13714 | 0.22465 | 0.40088 | 0.40088 | 0.44441 | 0.51104 | 0.10848 | 0.054615 | 0.35567 | 0.43309 | 0.43309 | 0.45524 | 0.51104 | 0.3067 | 0.29352
-0.39744 |  | 0.057419 | 0.15749 | 0.15749 | 0.21331 | 0.56903 | -0.055962 | -0.16308 | 0.28882 | 0.49539 | 0.49539 | 0.53351 | 0.56903 | 0.13363 | 0.065433 | 0.42844 | 0.52025 | 0.52025 | 0.54035 | 0.56903 | 0.35985 | 0.34108
-0.37179 |  | 0.08462 | 0.25241 | 0.25241 | 0.33756 | 0.62935 | -0.06885 | -0.19184 | 0.36207 | 0.58489 | 0.58489 | 0.61306 | 0.62935 | 0.16246 | 0.077513 | 0.50577 | 0.60186 | 0.60186 | 0.61669 | 0.62935 | 0.41764 | 0.39255
-0.34615 |  | 0.12096 | 0.37085 | 0.37085 | 0.47846 | 0.69141 | -0.083543 | -0.22323 | 0.44311 | 0.66726 | 0.66726 | 0.68501 | 0.69141 | 0.19494 | 0.090791 | 0.58595 | 0.6775 | 0.6775 | 0.68663 | 0.69141 | 0.47954 | 0.44746
-0.32051 |  | 0.16779 | 0.50367 | 0.50367 | 0.61798 | 0.7545 | -0.099981 | -0.25695 | 0.53008 | 0.74274 | 0.74274 | 0.75236 | 0.7545 | 0.23088 | 0.10515 | 0.66718 | 0.7482 | 0.7482 | 0.75296 | 0.7545 | 0.54479 | 0.50519
-0.29487 |  | 0.226 | 0.6386 | 0.6386 | 0.74172 | 0.81784 | -0.11801 | -0.29258 | 0.62064 | 0.81271 | 0.81271 | 0.81722 | 0.81784 | 0.26988 | 0.12041 | 0.74772 | 0.81529 | 0.81529 | 0.81741 | 0.81784 | 0.61248 | 0.56496
-0.26923 |  | 0.29579 | 0.76431 | 0.76431 | 0.84391 | 0.88055 | -0.13738 | -0.32956 | 0.71223 | 0.87856 | 0.87856 | 0.88041 | 0.88055 | 0.31139 | 0.13634 | 0.82599 | 0.87963 | 0.87963 | 0.88046 | 0.88055 | 0.6815 | 0.62584
-0.24359 |  | 0.37641 | 0.87364 | 0.87364 | 0.9266 | 0.94173 | -0.15773 | -0.36721 | 0.80225 | 0.94104 | 0.94104 | 0.9417 | 0.94173 | 0.35463 | 0.15264 | 0.90059 | 0.94144 | 0.94144 | 0.94172 | 0.94173 | 0.75061 | 0.68676
-0.21795 |  | 0.4661 | 0.96453 | 0.96453 | 0.99508 | 1.0004 | -0.17862 | -0.40475 | 0.88827 | 1.0002 | 1.0002 | 1.0004 | 1.0004 | 0.39863 | 0.16897 | 0.97038 | 1.0003 | 1.0003 | 1.0004 | 1.0004 | 0.81843 | 0.74656
-0.19231 |  | 0.56204 | 1.0387 | 1.0387 | 1.054 | 1.0557 | -0.19949 | -0.44134 | 0.9682 | 1.0556 | 1.0556 | 1.0557 | 1.0557 | 0.4423 | 0.18495 | 1.0344 | 1.0556 | 1.0556 | 1.0557 | 1.0557 | 0.88354 | 0.80399
-0.16667 |  | 0.6605 | 1.0993 | 1.0993 | 1.1061 | 1.1065 | -0.21975 | -0.47605 | 1.0403 | 1.1065 | 1.1065 | 1.1065 | 1.1065 | 0.48441 | 0.20017 | 1.0919 | 1.1065 | 1.1065 | 1.1065 | 1.1065 | 0.9445 | 0.85779
-0.14103 |  | 0.75709 | 1.1493 | 1.1493 | 1.1519 | 1.152 | -0.23874 | -0.50796 | 1.1035 | 1.152 | 1.152 | 1.152 | 1.152 | 0.52367 | 0.2142 | 1.1424 | 1.152 | 1.152 | 1.152 | 1.152 | 0.99987 | 0.90672
-0.11538 |  | 0.84708 | 1.1904 | 1.1904 | 1.1913 | 1.1913 | -0.25582 | -0.53618 | 1.1567 | 1.1913 | 1.1913 | 1.1913 | 1.1913 | 0.55881 | 0.22665 | 1.1851 | 1.1913 | 1.1913 | 1.1913 | 1.1913 | 1.0483 | 0.94958
-0.089744 |  | 0.92581 | 1.2235 | 1.2235 | 1.2238 | 1.2238 | -0.27035 | -0.55988 | 1.1996 | 1.2238 | 1.2238 | 1.2238 | 1.2238 | 0.5886 | 0.23712 | 1.2199 | 1.2238 | 1.2238 | 1.2238 | 1.2238 | 1.0887 | 0.98529
-0.064103 |  | 0.98901 | 1.2486 | 1.2486 | 1.2487 | 1.2487 | -0.28179 | -0.57834 | 1.2318 | 1.2487 | 1.2487 | 1.2487 | 1.2487 | 0.61197 | 0.24529 | 1.2463 | 1.2487 | 1.2487 | 1.2487 | 1.2487 | 1.1199 | 1.0129
-0.038462 |  | 1.0332 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | -0.28969 | -0.59098 | 1.2534 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | 0.62807 | 0.2509 | 1.264 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | 1.1412 | 1.0318
-0.012821 |  | 1.056 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | -0.29372 | -0.5974 | 1.2641 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | 0.63628 | 0.25375 | 1.273 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | 1.1519 | 1.0414
0.012821 |  | 1.056 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | -0.29372 | -0.5974 | 1.2641 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | 0.63628 | 0.25375 | 1.273 | 1.2741 | 1.2741 | 1.2741 | 1.2741 | 1.1519 | 1.0414
0.038462 |  | 1.0332 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | -0.28969 | -0.59098 | 1.2534 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | 0.62807 | 0.2509 | 1.264 | 1.2656 | 1.2656 | 1.2656 | 1.2656 | 1.1412 | 1.0318
0.064103 |  | 0.98901 | 1.2486 | 1.2486 | 1.2487 | 1.2487 | -0.28179 | -0.57834 | 1.2318 | 1.2487 | 1.2487 | 1.2487 | 1.2487 | 0.61197 | 0.24529 | 1.2463 | 1.2487 | 1.2487 | 1.2487 | 1.2487 | 1.1199 | 1.0129
0.089744 |  | 0.92581 | 1.2235 | 1.2235 | 1.2238 | 1.2238 | -0.27035 | -0.55988 | 1.1996 | 1.2238 | 1.2238 | 1.2238 | 1.2238 | 0.5886 | 0.23712 | 1.2199 | 1.2238 | 1.2238 | 1.2238 | 1.2238 | 1.0887 | 0.98529
0.11538 |  | 0.84708 | 1.1904 | 1.1904 | 1.1913 | 1.1913 | -0.25582 | -0.53618 | 1.1567 | 1.1913 | 1.1913 | 1.1913 | 1.1913 | 0.55881 | 0.22665 | 1.1851 | 1.1913 | 1.1913 | 1.1913 | 1.1913 | 1.0483 | 0.94958
0.14103 |  | 0.75709 | 1.1493 | 1.1493 | 1.1519 | 1.152 | -0.23874 | -0.50796 | 1.1035 | 1.152 | 1.152 | 1.152 | 1.152 | 0.52367 | 0.2142 | 1.1424 | 1.152 | 1.152 | 1.152 | 1.152 | 0.99987 | 0.90672
0.16667 |  | 0.6605 | 1.0993 | 1.0993 | 1.1061 | 1.1065 | -0.21975 | -0.47605 | 1.0403 | 1.1065 | 1.1065 | 1.1065 | 1.1065 | 0.48441 | 0.20017 | 1.0919 | 1.1065 | 1.1065 | 1.1065 | 1.1065 | 0.9445 | 0.85779
0.19231 |  | 0.56204 | 1.0387 | 1.0387 | 1.054 | 1.0557 | -0.19949 | -0.44134 | 0.9682 | 1.0556 | 1.0556 | 1.0557 | 1.0557 | 0.4423 | 0.18495 | 1.0344 | 1.0556 | 1.0556 | 1.0557 | 1.0557 | 0.88354 | 0.80399
0.21795 |  | 0.4661 | 0.96453 | 0.96453 | 0.99508 | 1.0004 | -0.17862 | -0.40475 | 0.88827 | 1.0002 | 1.0002 | 1.0004 | 1.0004 | 0.39863 | 0.16897 | 0.97038 | 1.0003 | 1.0003 | 1.0004 | 1.0004 | 0.81843 | 0.74656
0.24359 |  | 0.37641 | 0.87364 | 0.87364 | 0.9266 | 0.94173 | -0.15773 | -0.36721 | 0.80225 | 0.94104 | 0.94104 | 0.9417 | 0.94173 | 0.35463 | 0.15264 | 0.90059 | 0.94144 | 0.94144 | 0.94172 | 0.94173 | 0.75061 | 0.68676
0.26923 |  | 0.29579 | 0.76431 | 0.76431 | 0.84391 | 0.88055 | -0.13738 | -0.32956 | 0.71223 | 0.87856 | 0.87856 | 0.88041 | 0.88055 | 0.31139 | 0.13634 | 0.82599 | 0.87963 | 0.87963 | 0.88046 | 0.88055 | 0.6815 | 0.62584
0.29487 |  | 0.226 | 0.6386 | 0.6386 | 0.74172 | 0.81784 | -0.11801 | -0.29258 | 0.62064 | 0.81271 | 0.81271 | 0.81722 | 0.81784 | 0.26988 | 0.12041 | 0.74772 | 0.81529 | 0.81529 | 0.81741 | 0.81784 | 0.61248 | 0.56496
0.32051 |  | 0.16779 | 0.50367 | 0.50367 | 0.61798 | 0.7545 | -0.099981 | -0.25695 | 0.53008 | 0.74274 | 0.74274 | 0.75236 | 0.7545 | 0.23088 | 0.10515 | 0.66718 | 0.7482 | 0.7482 | 0.75296 | 0.7545 | 0.54479 | 0.50519
0.34615 |  | 0.12096 | 0.37085 | 0.37085 | 0.47846 | 0.69141 | -0.083543 | -0.22323 | 0.44311 | 0.66726 | 0.66726 | 0.68501 | 0.69141 | 0.19494 | 0.090791 | 0.58595 | 0.6775 | 0.6775 | 0.68663 | 0.69141 | 0.47954 | 0.44746
0.37179 |  | 0.08462 | 0.25241 | 0.25241 | 0.33756 | 0.62935 | -0.06885 | -0.19184 | 0.36207 | 0.58489 | 0.58489 | 0.61306 | 0.62935 | 0.16246 | 0.077513 | 0.50577 | 0.60186 | 0.60186 | 0.61669 | 0.62935 | 0.41764 | 0.39255
0.39744 |  | 0.057419 | 0.15749 | 0.15749 | 0.21331 | 0.56903 | -0.055962 | -0.16308 | 0.28882 | 0.49539 | 0.49539 | 0.53351 | 0.56903 | 0.13363 | 0.065433 | 0.42844 | 0.52025 | 0.52025 | 0.54035 | 0.56903 | 0.35985 | 0.34108
0.42308 |  | 0.037771 | 0.089463 | 0.089463 | 0.11904 | 0.51104 | -0.044862 | -0.13714 | 0.22465 | 0.40088 | 0.40088 | 0.44441 | 0.51104 | 0.10848 | 0.054615 | 0.35567 | 0.43309 | 0.43309 | 0.45524 | 0.51104 | 0.3067 | 0.29352
0.44872 |  | 0.024077 | 0.046016 | 0.046016 | 0.058013 | 0.4559 | -0.03547 | -0.11408 | 0.17022 | 0.30636 | 0.30636 | 0.34761 | 0.4559 | 0.086917 | 0.045074 | 0.28894 | 0.34326 | 0.34326 | 0.362 | 0.4559 | 0.25853 | 0.25014
0.47436 |  | 0.014866 | 0.021337 | 0.021337 | 0.024482 | 0.40367 | -0.027659 | -0.093872 | 0.12551 | 0.21864 | 0.21864 | 0.25024 | 0.40398 | 0.06873 | 0.036782 | 0.2294 | 0.25603 | 0.25603 | 0.2663 | 0.40398 | 0.21551 | 0.21111
0.5 |  | 0.0088881 | 0.0088885 | 0.0088885 | 0.0088885 | 0.0088885 | -0.021273 | -0.076408 | 0.08999 | 0.14429 | 0.14429 | 0.16274 | 0.17779 | 0.053639 | 0.029678 | 0.17779 | 0.17779 | 0.17779 | 0.17779 | 0.17779 | 0.17763 | 0.17642
|  | 0.44651 | 0.72698 | 0.72698 | 0.76287 | 0.8809 | -0.15734 | -0.35388 | 0.74002 | 0.85325 | 0.85325 | 0.86487 | 0.88524 | 0.34982 | 0.14777 | 0.82318 | 0.86273 | 0.86273 | 0.86803 | 0.88524 | 0.71345 | 0.65348
|  | 44.35% | 16.3% | 16.3% | 12.71% | 0.9104% | 104.7% | 124.4% | 15% | 3.675% | 3.675% | 2.513% | 0.4758% | 54.02% | 74.22% | 6.682% | 2.727% | 2.727% | 2.197% | 0.4758% | 17.65% | 23.65%
dx_pfa_hi | 0.43152
5.2158
5.6473
6.0788
6.5103
6.9418
7.3734
7.8049
8.2364
8.6679
9.0995
9.531
9.9625
10.394
10.8255
11.2571
11.6886
12.1201
12.5516
12.9832
13.4147
13.8462
14.2777
14.7092
15.1408
15.5723
16.0038
16.4353
16.8668
17.2984
17.7299
dx_pfa_lo | 0.043152
0.52158
0.56473
0.60788
0.65103
0.69418
0.73734
0.78049
0.82364
0.86679
0.90995
0.9531
0.99625
1.0394
1.0826
1.1257
1.1689
1.212
1.2552
1.2983
1.3415
1.3846
1.4278
1.4709
1.5141
1.5572
1.6004
1.6435
1.6867
1.7298
1.773
CMC OPTIMIZER — Integration x-grid (midpoints at LOW force)
dx | 0.033333
-0.48333
-0.45
-0.41667
-0.38333
-0.35
-0.31667
-0.28333
-0.25
-0.21667
-0.18333
-0.15
-0.11667
-0.083333
-0.05
-0.016667
0.016667
0.05
0.083333
0.11667
0.15
0.18333
0.21667
0.25
0.28333
0.31667
0.35
0.38333
0.41667
0.45
0.48333
Note on A56: henry zumbrun: GLOBAL consumer's risk per JCGM 106 eq.(19): probability a random item is non-conforming AND accepted, integrated over the production PDF. NOT the same as ILAC G8 specific risk (see row 58). Value also computed at C165:H165 which feeds the Dashboard.
Note on C56: JCGM 106 eq.19 global consumer's risk R_C. v17.6: integrated via sigma_t-substitution eta=Tol+sigma_t*u, 240-pt midpoint over u in [0,16], doubled for the symmetric two-sided tolerance. Resolves the measurement kernel even when sigma_t << sigma_product (high TUR), fixing the prior 30-node sigma_product-grid that under-counted tail risk at TUR>=12. Matches scipy to <0.2% at all TUR.
Note on A58: henry zumbrun: ILAC G8:2009 §2.2 specific (pointwise) risk. With the 0.98×U guard band the measured value at the acceptance limit A_U sits 1.96σ_t inside the tolerance, giving specific conformance probability pc = Φ(1.96) = 0.975 and specific PFA = 1 - 0.975 = 2.5%. This is the per-result risk a lab reports, distinct from the global process risk in row 56.
Note on C58: henry zumbrun: ILAC G8 specific PFA at acceptance limit = 1 - Φ((T-A_U)/σ_t). With 0.98×U GB this = 1-Φ(1.96) = 2.5%.
Note on C89: henry zumbrun: JCGM 106 eq. A.15 — PFA at Method 6 (empirical TUR-based) guard band. Numerical integration over tail of g0*P(accept¦eta), doubled for symmetric two-sided tolerance. Previous closed-form was wrong by 5-90x.

## Validation
Formula Validation & Proof Report
Morehouse True Cost Calculator v17.7
Tol=5.0 lbf ¦ EOPR=89% ¦ N=2,000 ¦ Force=5,000 lbf ¦ Rate=$200/hr ¦ 2026-06-03
Reference: JCGM 106:2012 (ISO/IEC Guide 98-4) · UKAS LAB 48 Ed.4 (2022) · scipy.integrate.quad · Sandia PSL (SAND2025-10736O)
TABLE 1: DASHBOARD PFA vs SCIPY — Specific (per-cert) · Global (process) · PFR (cost ref)
Machine | CMC | TUR | Excel Spec PFA | Scipy Spec PFA | Δ | ✓ | Excel Glob PFA | Scipy Glob PFA | Δ | ✓ | Excel PFR | Scipy PFR | Δ | ✓
BCM | 0.05% | 2 | 2.5% | 2.5% | 1.049e-05% | PASS | 0.06773% | 0.0679% | 0.0001716% | PASS | 33.98% | 33.98% | 3.25e-05% | PASS
PCM | 0.033% | 3.0303 | 2.5% | 2.5% | 1.049e-05% | PASS | 0.04798% | 0.0481% | 0.0001159% | PASS | 18.62% | 18.62% | 7.423e-05% | PASS
UCM Manual | 0.03% | 3.3333 | 2.5% | 2.5% | 1.049e-05% | PASS | 0.04418% | 0.0442% | 1.915e-05% | PASS | 16.3% | 16.3% | 3.8e-05% | PASS
UCM Auto | 0.02% | 5 | 2.5% | 2.5% | 1.049e-05% | PASS | 0.03074% | 0.0308% | 5.636e-05% | PASS | 9.48% | 9.48% | 7.389e-05% | PASS
Deadweight | 0.002% | 50 | 2.5% | 2.5% | 1.049e-05% | PASS | 0.003325% | 0.0033% | 2.521e-05% | PASS | 0.7225% | 0.7225% | 3.482e-05% | PASS
Competitor 1 | 0.05% | 2 | 2.5% | 2.5% | 1.049e-05% | PASS | 0.06773% | 0.0679% | 0.0001716% | PASS | 33.98% | 33.98% | 3.25e-05% | PASS
Competitor 2 | 0.06% | 1.6667 | 2.5% | 2.5% | 1.049e-05% | PASS | 0.07804% | 0.0782% | 0.0001614% | PASS | 44.35% | 44.35% | 6.168e-05% | PASS
TABLE 2: REWORK = PFR × N × $Cost (7 checks)
Machine | PFR (M6) | × Cals | × $Cost | Expected | Excel | Δ | Status
BCM | 7.061% | 1000 | 100 | 7061.174 | 7061.174 | 0 | PASS
PCM | 4.84% | 1000 | 100 | 4839.5276 | 4839.5276 | 0 | PASS
UCM Manual | 4.003% | 1000 | 100 | 4002.552 | 4002.552 | 0 | PASS
UCM Auto | 1.893% | 1000 | 100 | 1892.8336 | 1892.8336 | 0 | PASS
Deadweight | 0.541% | 1000 | 100 | 541.0064 | 541.0064 | 1.1369e-13 | PASS
Competitor 1 | 10.87% | 1000 | 100 | 10868.7975 | 10868.7975 | 0 | PASS
Competitor 2 | 15.12% | 1000 | 100 | 15121.1237 | 15121.1237 | 0 | PASS
TABLE 3: GUARD BANDS — Deterministic (20 checks)
TUR | U | Excel GB_Spec | Calc GB_Spec |  | Excel GB_M6 | Calc GB_M6 |  | Consumed
1.5 | 6.6667 | 6.5333 | 6.5333 | PASS | 2.4012 | 2.4012 | PASS | 65.33%
2 | 5 | 4.9 | 4.9 | PASS | 1.4082 | 1.4082 | PASS | 49%
2.5 | 4 | 3.92 | 3.92 | PASS | 0.85814 | 0.85814 | PASS | 39.2%
3 | 3.3333 | 3.2667 | 3.2667 | PASS | 0.51773 | 0.51773 | PASS | 32.67%
3.5 | 2.8571 | 2.8 | 2.8 | PASS | 0.29128 | 0.29128 | PASS | 28%
4 | 2.5 | 2.45 | 2.45 | PASS | 0.1328 | 0.1328 | PASS | 24.5%
4.6 | 2.1739 | 2.1304 | 2.1304 | PASS | 0 | 0 | PASS | 21.3%
5 | 2 | 1.96 | 1.96 | PASS | 0 | 0 | PASS | 19.6%
10 | 1 | 0.98 | 0.98 | PASS | 0 | 0 | PASS | 9.8%
24 | 0.41667 | 0.40833 | 0.40833 | PASS | 0 | 0 | PASS | 4.083%
TABLE 4: ORDERING — Global ≤ M6 ≤ Specific (7 checks)
Machine | RW Global | RW M6 | RW Specific | G≤M6 | M6≤S | Status
BCM | 5215.792 | 7061.174 | 25100.7166 | ✓ | ✓ | PASS
PCM | 3544.7645 | 4839.5276 | 19392.2608 | ✓ | ✓ | PASS
UCM Manual | 2980.6053 | 4002.552 | 17106.6567 | ✓ | ✓ | PASS
UCM Auto | 1892.8336 | 1892.8336 | 10238.5617 | ✓ | ✓ | PASS
Deadweight | 541.0064 | 541.0064 | 1039.1573 | ✓ | ✓ | PASS
Competitor 1 | 8085.3146 | 10868.7975 | 34053.0739 | ✓ | ✓ | PASS
Competitor 2 | 6465.3029 | 15121.1237 | 43680.5922 | ✓ | ✓ | PASS
TABLE 5: CROSS-SHEET REFERENCES (15 checks)
Reference | Description | Value | Expected | Status
Risk Calc B5 | Tol | 10 | 10 | PASS
Risk Calc B6 | EOPR | 0.85 | 0.85 | PASS
Risk Calc B7 | σ_p | 6.9467 | 6.9467 | PASS
Risk Calc B8 | N | 1000 | 1000 | PASS
Risk Calc B9 | RW$ | 100 | 100 | PASS
RA B9 | EOPR | 0.85 | 0.85 | PASS
RA B10 | N | 1000 | 1000 | PASS
RA B11 | RW$ | 100 | 100 | PASS
RA E5 | TolHi | 10 | 10 | PASS
RA E6 | TolLo | 0.025 | 0.025 | PASS
BRK B27 | N | 1000 | 1000 | PASS
BRK B28 | Rate | 200 | 200 | PASS
Dash B9 | RW$ | 100 | 100 | PASS
Dash B16 | Tol | 10 | 10 | PASS
BRK O31 | Hrs/Cal link | 0.9 | 0.9 | PASS
TABLE 6: BREAKEVEN — Rev($150/cal) − Labor($200/hr) − OpCost = Profit
Machine | Price | Hrs/Cal | Labor/yr | OpCost | Revenue | Profit M6 | Check | Payback | Req$/Cal | Req$/Hr | Status
BCM | 35000 | 1.5 | 300000 | 42867.7719 | 150000 | -192867.7719 | -192867.7719 | Never | 354.5344 | 236.3563 | PASS
PCM | 25000 | 1.5 | 300000 | 30263.5085 | 150000 | -180263.5085 | -180263.5085 | Never | 338.5968 | 225.7312 | PASS
UCM Manual | 50000 | 1 | 200000 | 12868.7975 | 150000 | -62868.7975 | -62868.7975 | Never | 229.5355 | 229.5355 | PASS
UCM Automated | 100000 | 0.5 | 100000 | 7436.1928 | 150000 | 42563.8072 | 42563.8072 | 2.3yr | 140.7695 | 281.5391 | PASS
Deadweight | 300000 | 1 | 200000 | 541.0064 | 150000 | -50541.0064 | -50541.0064 | Never | 300.541 | 300.541 | PASS
Competitor | 30000 | 1.5 | 300000 | 56122.264 | 150000 | -206122.264 | -206122.264 | Never | 366.1223 | 244.0815 | PASS
Competitor 2 | 40000 | 1.5 | 300000 | 75059.1029 | 150000 | -225059.1029 | -225059.1029 | Never | 388.3924 | 258.9283 | PASS
TABLE 7: STANDARD SWITCHING — Swap+Cal+RW=Total (3 checks)
MinLoad% | Swap$/yr | Cal$/yr | RW M6 | Expected | Excel | Δ | Status
5% | 0 | 2000 | 4936.835 | 6936.835 | 6936.835 | 0 | PASS
10% | 0 | 4000 | 2765.5605 | 6765.5605 | 6765.5605 | 0 | PASS
20% | 0 | 6000 | 1892.8336 | 7892.8336 | 7892.8336 | 0 | PASS
TABLE 8: CELL CHAIN — 10,000→25 lbf at 5% Loading
Check | Excel | Expected | Status
Cells needed | 2 | 2 | PASS
Swaps/run | 1 | 1 | PASS
Cell 1 | 10,000 lbf | → 500 lbf | No — need next cell | INFO
Cell 2 | 500 lbf | → 25 lbf | ✓ Yes | PASS
VALIDATION SUMMARY
Result | ✅ ZERO ERRORS — ALL 5,741 CHECKS PASSED
Total formulas | 5,588
Validation checks | 5741
Errors | 0
Warnings | 0
JCGM 106 conformance | Global PFA via eq.19/20 integral; resistor benchmark RC=0.99% / RP=6.9% vs JCGM 1% / 7%
LAB 48 conformance | C₉₅≥1.5 all machines; specific PFA=2.5% at Aᵤ (App C kᵥ=1.96)
Guard band convention | Engine w=0.98×U (fixed 2.5% specific risk) — permitted App C alternative to default w=U₉₅
Global PFA accuracy | < 0.2% vs scipy across all TUR (σₜ-resolved integration, N=240)
Global PFA at high TUR | RESOLVED v17.6: σₜ-substitution resolves the measurement kernel; Deadweight TUR 50 now 0.12% vs scipy (was 0.54%)
Reference | JCGM 106:2012 · UKAS LAB 48 Ed.4 (2022) · scipy · Sandia PSL (SAND2025-10736O)
INPUT CELL LEGEND
Editable input | Sage shading + green bold + thin green border
Calculated | No shading — formula cell, do not edit
TABLE 9: JCGM 106:2012 CONFORMANCE — Global risk integrals (eq.19/20, A.15–A.17)
Check | Clause | Engine | JCGM / scipy | Δ | Status
σ_product from EOPR | A.4.1 / eq.7 | 3.1285 | 3.1285 | 0 | PASS
Cm = T/2U identity | eq.12 | 2 | 2 | 0 | PASS
Resistor RC (benchmark) | 9.5.3 ex. | 0.99% | 1% | 0.01% | PASS
Resistor RP (benchmark) | 9.5.3 ex. | 6.9% | 7% | 0.1% | PASS
Global PFA ≤ 2% (per machine, eq.19)
| BCM | 0.06773% | 2% | 1.932% | PASS
| PCM | 0.04798% | 2% | 1.952% | PASS
| UCM Manual | 0.04418% | 2% | 1.956% | PASS
| UCM Auto | 0.03074% | 2% | 1.969% | PASS
| Deadweight | 0.003325% | 2% | 1.997% | PASS
TABLE 10: UKAS LAB 48 Ed.4 CONFORMANCE — Decision rule & guard band
Check | Clause | Engine | LAB 48 req | Status
C₉₅ = (Tᵤ−Tʟ)/(2·U₉₅) identity | p.7 | 2 | 2 | PASS
C₉₅ ≥ 1.5 (validates w=U₉₅ rule) | p.9
BCM |  | 2 | 1.5 | PASS
PCM |  | 3.0303 | 1.5 | PASS
UCM Manual |  | 3.3333 | 1.5 | PASS
UCM Auto |  | 5 | 1.5 | PASS
Deadweight |  | 50 | 1.5 | PASS
Competitor 2 (worst) |  | 1.6667 | 1.5 | PASS
Specific PFA at Aᵤ (guarded accept) | B.7 / App C |  | PFAₘₐₓ
BCM |  | 2.5% | 2.5% | PASS
PCM |  | 2.5% | 2.5% | PASS
UCM Manual |  | 2.5% | 2.5% | PASS
UCM Auto |  | 2.5% | 2.5% | PASS
Deadweight |  | 2.5% | 2.5% | PASS
App C two-sided check (PFA at Aᵤ ≈ PFAₘₐₓ) | App C step 7 | 2.5% | 2.5% | PASS
Guard band convention | p.9 / 8.3.2.3 | w = 0.98×U (kᵥ≈1.96) | w = U₉₅ (k=2) default | NOTE
TABLE 11: INDEPENDENT SELF-CHECK — recomputes engine risks via a DIFFERENT method (Simpson's rule). Edit any input, re-read: all stay PASS if engine is sound.
Machine | Risk | Engine | Self-check (Simpson) | Δ | ✓
BCM | Global PFA | 0.06773% | 0.06785% | 0.0001233% | PASS
PCM | Global PFA | 0.04798% | 0.04806% | 7.533e-05% | PASS
UCM Manual | Global PFA | 0.04418% | 0.04425% | 6.752e-05% | PASS
UCM Auto | Global PFA | 0.03074% | 0.03079% | 4.287e-05% | PASS
Deadweight | Global PFA | 0.003325% | 0.003329% | 3.903e-06% | PASS
BCM | Specific PFA | 2.5% | 2.5% | 0% | PASS
PCM | Specific PFA | 2.5% | 2.5% | 0% | PASS
UCM Manual | Specific PFA | 2.5% | 2.5% | 0% | PASS
UCM Auto | Specific PFA | 2.5% | 2.5% | 0% | PASS
Deadweight | Specific PFA | 2.5% | 2.5% | 0% | PASS
BCM | PFR | 33.98% | 33.98% | 5.804e-05% | PASS
PCM | PFR | 18.62% | 18.62% | 5.494e-05% | PASS
UCM Manual | PFR | 16.3% | 16.3% | 5.402e-05% | PASS
UCM Auto | PFR | 9.48% | 9.48% | 4.899e-05% | PASS
Deadweight | PFR | 0.7225% | 0.7226% | 0.0001065% | PASS
SELF-CHECK RESULT | ✅ ALL 15 SELF-CHECKS PASS — engine matches independent method | 0.0001233% | max Δ
Note on E7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Scipy columns are a static benchmark regenerated for the current scenario (Tol=5 lbf, Force=5000, EOPR=89%) via scipy.integrate.quad. Excel columns are live engine links (Risk Calc rows 56/57/58); Δ recomputes automatically. Re-run scipy if Dashboard inputs change materially.
Note on C112: henry zumbrun: Engine quadrature replicated on JCGM 106 §9.5.3 resistor case (TU=1500.2, TL=1499.8, AU=1500.18, u0=0.12, um=0.04): engine R_C=0.99%, scipy=0.99%. Source: JCGM 106:2012 §9.5.3, verified via scipy.integrate.quad.
Note on C113: henry zumbrun: Engine quadrature replicated on JCGM 106 §9.5.3 resistor case: engine R_P=6.92%, scipy=6.90%, JCGM published 7%. Source: JCGM 106:2012 §9.5.3, verified via scipy.integrate.quad.
Note on E137: henry zumbrun: LAB 48 App C, steps 6-7: for a two-sided spec, recompute PFA with the measured value sitting at the acceptance limit A_U. Because C95>=1.5 for every machine, only one tolerance limit significantly overlaps, so the single-sided guard-band factor (kw=1.96 at PFAmax=2.5%) remains valid and PFA stays ~2.5%. Verified via scipy two-sided integral.
Note on E138: henry zumbrun: DEVIATION (conservative, conformant): LAB 48 p.9 cites a DEFAULT guard band w=U95 (k=2), which yields ~2.3-2.5% PFA. The engine instead sets w=0.98*U, i.e. the acceptance limit sits exactly 1.96-sigma_t inside the tolerance, pinning the specific PFA to exactly 2.5% (kw=NORM.S.INV(1-0.025)=1.96 per LAB 48 App C eq.C.1/C.3). This is a permitted alternative guard-band factor under App C, not the k=2 default. Both are valid LAB 48 decision rules; the engine's choice targets a fixed 2.5% specific risk rather than a fixed k=2 band.